## Notes
In 2018, COCIR and MedTech Europe commissioned the consultancy Rina Technology Ltd. to perform an analysis of all CMR 1A and 1B and human health-relevant endocrine disrupting (ED) substances, identifying the ones that may be used in medical devices, to help companies comply with their obligations under Section 10.4. of MDR Annex I, Chapter II. In 2019 and 2020 this analysis has reviewed and updated by BOMcheck. Thanks to this analysis, companies now have access to a reduced list of screened substances (out of a total of >1000) that they may consider as a first priority when determining those medical devices in scope of Section 10.4., containing substances that trigger justification and labelling obligations. Also the "COCIR flowchart" and the "MedTech Europe (MTE) flowchart" to guide companies towards compliance are being made available to the members of both associations through this document. This spreadsheet does not relieve companies from any requirement to assess the full list of substances that may be relevant for MDR compliance.
Last updated: January 2024
With the kind support of BOMcheck, the list has been updated to:
1. Reflect the screening of 1 newly classified CMR 1B substance (from the 20th ATP to CLP) which was carried out by BOMcheck
2. Align it with the IEC 62474 / BOMcheck screening of the REACH Candidate List substances
3. Remove 4 REACH Candidate List substances from the 17 January REACH Candidate List update in the 'All CMR' section which are not classed as CMR or ED substances
These and further updates (e.g. those based on information received from members) are detailed in the 'All CMR' section of this spreadsheet.
Terms of use: FOR INTERNAL USE BY MEDTECH EUROPE AND COCIR MEMBERS ONLY
The data in this spreadsheet including the list of substances and their uses belong to BOMcheck, MedTech Europe and COCIR. Members of MedTech Europe and COCIR are permitted to use this at no charge for the purposes of MDR Annex I Section 10.4. compliance but are not permitted to share this with third parties or use for its own internal development.
Notes
Substance names, CAS and EC numbers are those listed in the ECHA Classification and Labelling Inventory. Only substances with harmonised CMR 1A and 1B classifications have been considered apart from those proposed for reclassification.
The substances 1 - 99 in the CMR list tab may occur in medical devices of MedTech Europe and COCIR members and be relevant to Annex I Section 10.4. The other substances are, according to RINA's assessment, less likely to be present in medical devices. Note that some lines include more than one substance, e.g. lead compounds, or more than one material, e.g. lead which is used as lead metal and in various alloys.
The 'All CMR' list contains all CMR 1A and 1B substances with harmonised classifications. Note that only CMR of category 1A and 1B are relevant to the MDR (CMR 2 category included for information only in brackets).
All human health-relevant endocrine disruptors (as identified in the REACH Candidate list) are also CMR 1A or 1B. However some of the environmental endocrine disruptors are not CMRs. These can be found at https://echa.europa.eu/candidate-list-table (filter list for environmental endocrine disruptors). Currently, no endocrine disrupting substances have been identified according to the criteria established under the Biocidal Products Regulation.
Legend
1 = May occur at > 0.1% in materials of medical device in scope of Annex I Section 10.4. Likelihood is variable with some being much less likely than others. Some are restricted or regulated in some applications. Strictest concentration limit assumed.
2 = May be used in medical devices, but either not in patient contact materials or gas, liquid or solid transfer materials, or only at < 0.1% (N if used and present in batteries or in electronic components).
3 = Much less likely to occur in patient contact materials or gas, liquid or solid transfer materials.
4 = Uses uncertain or unknown, however very unlikely to be used in medical devices.
N = Not used in medical device materials or will never occur in patient contact materials or gas, liquid or solid transfer materials.

## MTE Flowchart
COCIR members may also make use of this flowchart provided they keep it for internal use and do not distribute it to third parties.

## CMR & ED list (hidden)
| EC No. | CAS No. | Name all | Uses | Hazard classification | Type | Applicable EU legislation | UPDATE
1 | 204-327-1 | 119-47-1 | 6,6'-di-tert-butyl-2,2'-methylenedi-p-cresol [DBMC] | Assessed as an SVHC as part of September 2021 screening. This substance can be used as an antioxidant additive in the manufacture of polymers, where it is added after the polymerisation process and so may be present in the supplied article in concentrations >0.1% by weight of the article. Polymers containing this substance as an antioxidant may be used in medical device materials and patient contact materials. https://www.solvay.com/en/chemical-categories/polymer-additives-uv-light-stabilizers-antioxidants-and-antistatic-agents/antioxidants-sterile-medical-applications-ppe | Repr. 1B | 1 |  | NOV 2021
2 | 235-541-3 | 12267-73-1 | tetraboron disodium heptaoxide, hydrate | Flame retardant used in paper, card, paint, wood and also added to cellulose insulation | Repr. 1B | 1 | REACH SVHC | NOV 2021
3 | 237-560-2 | 13840-56-7 | orthoboric acid, sodium salt | Flame retardant used in paper, card, paint, wood and also added to cellulose insulation | Repr. 1B | 1 | REACH SVHC | NOV 2021
4 | 215-540-4 | 1303-96-4, 1330-43-4, 12179-04-3 | disodium tetraborate decahydrate: borax decahydrate disodium tetraborate pentahydrate | Flame retardant used in paper, card, paint, wood and also added to cellulose insulation. The substances can also be used in contact lens storage solutions and multi-purpose disinfecting solutions. | Repr. 1B | 1 | REACH SVHC | NOV 2021
5 | 201-025-1 | 77-40-7 | 4,4'-(1-methylpropylidene)bisphenol | BPB is structurally similar to BPA and so may be used to make polymers such as polycarbonate and epoxy resins, but usually present at <0.1%. Similarly, may also be used as alternative to BPA as an uncommon additive in PVC at about 0.1% | ED Health ED Envi | 1 | REACH SVHC | NOV 2021
6 | 222-883-3 293-901-5 | 3648-18-8 91648-39-4 | dioctyltin dilaurate; [1] stannane, dioctyl-, bis(coco acyloxy) derivs. [2] | Dioctyltin (DOT) compounds are added to PVC plastic (at a typical addition rate of between 0.001 and 1% w/w) to make it more stable to heating. DOT compounds prevent the plastic from degrading when it is heated to form different products (for example when PVC is injection moulded it may be heated to between 170°C and 210°C). DOT also protects the PVC plastic against degradation from sunlight. DOT compounds are generally used in colourless and/or transparent plastic articles. DOT compounds are used in rigid extrusions (e.g. trims, conduits), rigid mouldings (e.g. bottles) and medical applications (e.g. valves, containers, flexible tubes, blood bags, surgical gloves). | Repr. 1B | 1 |  | OCT 2020
7 | 248-523-5 | 27554-26-3 | diisooctyl phthalate | Plasticizer in plastic materials such as PVC and other plastic polymers. In flexible PVC the typical phthalate content can range from 20 to 50% w/w. | Repr. 1B | 1 |  | OCT 2020
8 | 236-671-3 | 13463-41-7 | pyrithione zinc; (T-4)-bis[1-(hydroxy-.kappa.O)pyridine-2(1H)-thionato-.kappa.S]zinc | This substance is primarily used for its antibacterial properties as a biocide in cosmetic and personal care products. It can also be used to produce antimicrobial textiles and as a biocide in polyurethane foam, PVC and film preservatives. Zinc pyrithione can be added to PVC in a DIDP paste which contains 5% zinc pyrithione. The DIDP paste is added to flexible and plastisol PVC applications at 1% to 4% by weight of the total PVC formulation to provide protection against attack from a wide spectrum of fungi and bacteria. Therefore, the zinc pyrithione could be present up to 0.2% in the supplied PVC article - https://www.valtris.com/intercide-zinc-pyrithione-paste/ Powedered ZNP can also be added to Styrene-ethylene/butylene-styrene (SEBS) based thermoplastic elastomers (TPE) at 1.5% w/w, which are used in a wide range of products with elastic properties, such as remote control keyboards, cell phone covers, tooth brush cables, toys and others, to provide antimicrobial properties - https://www.scielo.br/pdf/mr/v20n5/1516-1439-mr-1980-5373-MR-2017-0137.pdf | Repr. 1B | 1 |  | OCT 2020
9 | 222-583-2 | 3542-36-7 | dichlorodioctylstannane | Dioctyltin (DOT) compounds are added to PVC plastic (at a typical addition rate of between 0.001 and 1% w/w) to make it more stable to heating. DOT compounds prevent the plastic from degrading when it is heated to form different products (for example when PVC is injection moulded it may be heated to between 170°C and 210°C). DOT also protects the PVC plastic against degradation from sunlight. DOT compounds are generally used in colourless and/or transparent plastic articles. DOT compounds are used in rigid extrusions (e.g. trims, conduits), rigid mouldings (e.g. bottles) and medical applications (e.g. valves, containers, flexible tubes, blood bags, surgical gloves). | Repr. 1B | 1 |  | 2020-11-01
10 | 245-152-0 | 22673-19-4 | dibutylbis(pentane-2,4- dionato-O,O’)tin | Dibutylbis(pentane-2,4-dionato-O,O')tin has similar properties and uses to dibutyltin dichloride which is already listed in the Candidate list as likely to be present in supplied articles at concentrations > 0.1% w/w. Butylin compounds are used as biocides and as stabilisers in plastics. Dibutylbis(pentane-2,4-dionato-O,O')tin has widespread uses as a catalyst and in the manufacturing of adhesives, sealants, coatings, dyes, polymer preparations, resins and rubber. It can be found in supplied articles in the following applications and concentration ranges: • Adhesive/sealant: 0.1-10% • Hardener/coating: 5-10% | Repr 1B | 1 |  | APRIL 2020
11 | 308-208-6 | 97925-95-6 | ethanol, 2,2’-iminobis-, N- (C13-15-branched and linear alkyl) derivs. | Ethanol, 2,2'-iminobis-, N-(C13-15-branched and linear alkyl) derivatives are added during the manufacture of plastics as an antistatic agent to reduce or eliminate a static build up on the surface of the plastic. Anti-static additives work by lowering the resistivity of the plastic so that charges are mobile and therefore will not cause static issues such as dust attraction or electrostatic discharge. Typical addition rates for the substance are 1-4% in ABS plastic and 0.1-0.4% w/w in polyethylene. | Reprotox 1B | 1 |  | APRIL 2020
12 | 231-158-0 | 7440-48-4 | cobalt | Alloys (steels, hard coatings, etc.) | Carcinogen 1B Reprotox 1B | 1 |  | APRIL 2020
13 | 276-090-2 | 71850-09-4 | Diisohexyl phthalate | Used as plasticizer in plastic materials such as PVC and other plastic polymers. In flexible PVC the typical phthalate content can range from 20 to 50% w/w. | Reprotox 1B | 1 | REACH SVHC | APRIL 2020 (included in REACH Candidate List: JAN 2020)
14 | 215-125-8 | 1303-86-2 | diboron trioxide: boric oxide | Diboron trioxide, also known as boric oxide, and is the anhydride form of boric acid. Assessment by BOMcheck and IEC 62474 identifies that Borates are used in wooden veneer sheets and pressed wooden panels, as a constituent within the starch adhesive, as a flame retardant and also as stabilizer in aminoplastic resins. Borates are used as a biocide in professional and industrial wood preservation. Diboron trioxide is used as a flame retardant in wood, paper, cotton and other plant-derived materials. | Reprotox 1B | 1 | REACH SVHC | APRIL 2019 (included in REACH Candidate List: June 2012)
15 |  |  | Benzidine based azo dyes: 4,4'-diarylazobiphenyl dyes, with the exception of those specified elsewhere in this Annex | Dyes of various colours, although some are banned by REACH in fabrics with skin contact | Carcinogen 1B | 1 | REACH Annex XVII | FEB 2018
16 | 401-500-5 | 118658-99-4 | (Methylenebis(4,1-phenylenazo(1-(3-(dimethylamino)propyl)-1,2-dihydro-6-hydroxy-4-methyl-2-oxopyridine-5,3-diyl)))-1,1'-dipyridinium dichloride dihydrochloride | Azo dye used in textiles, paper | Carcinogen 1B | 1 |  | FEB 2018
17 |  |  | [Calcium-aluminium-silicate fibres with random orientation with the following representative composition (% given by weight): SiO2 50,0-56,0 %, Al2O3 13,0-16,0 %, B2O3 5,8-10,0 %, Na2O < 0,6 %, K2O < 0,4 %, CaO 15,0-24,0 %, MgO < 5,5 %, Fe2O3 < 0,5 %, F2 < 1,0 %. Process: typically produced by flame attenuation and rotary process. (Additional individual elements may be present at low levels; the process list does not preclude innovation).] | Used as thermal insulation, also is a type of e-glass used mainly as fibre reinforcement of plastics | Carcinogen 1B | 1 | REACH SVHC | FEB 2018
18 | 271-093-5 | 68515-50-4 | 1,2-benzenedicarboxylic acid, dihexyl ester, branched and linear | Plasticiser | Reproductive toxin 1B | 1 | REACH SVHC. Phthalate | FEB 2018
19 | 284-032-2 | 84777-06-0 | 1,2-benzenedicarboxylic acid, dipentylester, branched and linear | Plasticiser | Reproductive toxin 1B | 1 | REACH SVHC. Phthalate | FEB 2018
20 | 276-158-1 | 71888-89-6 | 1,2-benzenedicarboxylic acid: di-C<sub>6-8</sub>-branched alkylesters, C<sub>7</sub>-rich | Plasticiser | Reproductive toxin 1B | 1 | REACH SVHC. Phthalate | FEB 2018
21 | 271-084-6 | 68515-42-4 | 1,2-benzenedicarboxylic acid: di-C<sub>7-11</sub>-branched and linear alkylesters | Plasticiser | Reproductive toxin 1B | 1 | REACH SVHC. Phthalate | FEB 2018
22 | 219-603-7 | 2475-45-8 | 1,4,5,8-tetraaminoanthraquinone: C.I. Disperse Blue 1 | Blue dye used in plastics, fabrics, etc. | Carcinogen 1B | 1 |  | FEB 2018
23 | 202-918-9 | 101-14-4 | 2,2'-dichloro-4,4'-methylenedianiline: 4,4'-methylene bis(2-chloroaniline): salts of 2,2'-dichloro-4,4'-methylenedianiline: salts of 4,4'-methylenebis(2-chloroaniline) | "MOCA", used mainly to make polyurethanes (some used also for epoxy resins), but residues may exceed 0.1% acccording to ECHA | Carcinogen 1B | 1 | REACH SVHC | FEB 2018
24 |  | 72319-19-8 | 2,7-naphthalenedisulfonic acid, nickel(II) salt | May be used as a green textile dye but uncertain and unlikely | Carcinogen 1A, Reproductive toxin 1B (Mutagen 2) | 1 |  | FEB 2018
25 | 420-580-2 | 151798-26-4 | 2-[2-hydroxy-3-(2-chlorophenyl)carbamoyl-1-naphthylazo]-7-[2-hydroxy-3-(3-methylphenyl)carbamoyl-1-naphthylazo]fluoren-9-one | Pigment (used in some photocopiers) | Reproductive toxin 1B | 1 |  | FEB 2018
26 | 239-622-4 | 15571-58-1 | 2-ethylhexyl 10-ethyl-4,4-dioctyl-7-oxo-8-oxa-3,5-dithia-4-stannatetradecanoate (DOTE) | PVC stabiliser. Mainly used in rigid PVC including transparent PVC, but may also occur in flexible PVC | Reproductive toxin 1B | 1 | REACH SVHC | FEB 2018
27 | 400-600-6 | 71868-10-5 | 2-methyl-1-(4-methylthiophenyl)-2-morpholinopropan-1-one | Photoinitiator resulting in inclusion into a matrix, including application in coatings, adhesives and inks. | Reproductive toxin 1B | 1 | REACH SVHC | JAN 2020
28 | 209-321-2 | 569-61-9 | 4,4'-(4-iminocyclohexa-2,5-dienylidenemethylene)dianiline hydrochloride: C.I. Basic Red 9 | Red dye | Carcinogen 1B | 1 |  | FEB 2018
29 | 202-027-5 | 90-94-8 | 4,4'-bis(dimethylamino)benzophenone: Michler's ketone | Impurity in plastics and textiles from various dues and pigments | Carcinogen 1B (Mutagen 2) | 1 | REACH SVHC | FEB 2018
30 | 202-696-3 | 98-73-7 | 4-<i>tert</i>-butylbenzoic acid, 4-t-butylbenzoic acid | PVC stabiliser and resin modifier (alkyds) | Reproductive toxin 1B | 1 |  | FEB 2018
31 | 200-453-6 | 60-09-3 | 4-aminoazobenzene: 4-phenylazoaniline | Also called aniline yellow. Used as a yellow pigments and in inks including inks for inkjet printersmicroscopy. Used in insecticides, lacquers, varnishes, waxes, oil stains, and styrene resins.Also used for vital staining, in pyrotechnics for yellow colored smokes. It is also an intermediate in synthesis of other dyes | Carcinogen 1B | 1 |  | FEB 2018
32 | 400-340-3 | 85136-74-9 | 6-hydroxy-1-(3-isopropoxypropyl)-4-methyl-2-oxo-5-[4-(phenylazo)phenylazo]-1,2-dihydro-3-pyridinecarbonitrile | Also called Disperse Orange 149. Orange pigment. | Carcinogen 1B | 1 |  | FEB 2018
33 | 232-143-1 | 7789-09-5 | ammonium dichromate | Mainly used as a process chemical, but may rarely be used as a mordant for dying fabrics and textiles (may possibly exceed 0.1%) | Carcinogen 1B, Mutagen 1B, Reproductive toxin 1B | 1 | REACH SVHC | FEB 2018
34 | 221-470-5 | 3108-42-7 | ammonium nonadecafluorodecanoate | Water-proofing textiles | Reproductive toxin 1B (carcinogen 2) | 1 | REACH SVHC | FEB 2018
35 | 223-320-4 | 3825-26-1 | ammonium pentadeca- fluorooctanoate | used as a processing aid for manufacture of fluoropolymers such as PVDF. Also used as an additive in lubricating oils, in cleaning agents, for waterproofing and textile coatings and in liquid crystal displays. Occurs in water and stain resistant textiles. Will not exceed 0.1% in a material, but could as an ultra-thin waterproof coating | Reproductive toxin 1B (carcinogen 2) | 1 |  | FEB 2018
36 | 249-415-0 | 29081-56-9 | ammonium perfluorooctane sulfonate: ammonium heptadecafluorooctanesulfonate | used as a processing aid for manufacture of fluoropolymers such as PVDF. Also used as an additive in lubricating oils, in cleaning agents, for waterproofing and textile coatings and in liquid crystal displays. Occurs in water and stain resistant textiles. Will not exceed 0.1% in a material, but could as an ultra-thin waterproof coating | Reproductive toxin 1B (carcinogen 2) | 1 |  | FEB 2018
37 | 201-622-7 | 85-68-7 | Benzyl butyl phthalate, BBP | Plasticiser | Reproductive toxin 1B, Endocrine Disruptor | 1 | REACH SVHC. Phthalate | FEB 2018
38 | 231-150-7 | 7440-41-7 | beryllium metal (and alloys) | Copper beryllium and nickel beryllium alloys are the main alloy used for its high electrical and/or thermal conductivity, high strength and hardness, non-magnetic properties, and good corrosion and fatigue resistance, but is unlikely in patient contact materials. Nickel-beryllium alloy is used as springs in thermostats, bellows, diaphragms, burn-in connectors and sockets. Beryllium metal is also used as X-ray tube windows. Beryllium aluminium lightweight high strength alloys are used in hard disc drives, aerospace, space, etc. One unusual repoprted use is in loudspeakers. | Carcinogen 1B | 1 |  | FEB 2018
39 | 204-211-0 | 117-81-7 | bis(2-ethylhexyl) phthalate: di-(2-ethylhexyl) phthalate: DEHP | Plasticiser in PVC, rubber, sealants, paint, ink. Etc. | Reproductive toxin 1B, Endocrine Disruptor | 1 | REACH SVHC. Phthalate, environmental endocrine disruptor | FEB 2018
40 | 204-212-6 | 117-82-8 | bis(2-methoxyethyl) phthalate | SVHC: Phthalate plasticiser in polymeric materials and paints, lacquers and varnishes, including printing inks | Reproductive toxin 1B | 1 | REACH SVHC. Phthalate | FEB 2018
41 | 201-245-8 | 80-05-7 | bisphenol A: 4,4’-isopropylidenediphenol | Used to make polymers such as polycarbonate and epoxy resins, but usually present at <0.1%. Uncommon additive in PVC at about 0.1% | Reproductive toxin 1B, Endocrine Disruptor | 1 | REACH SVHC, environmental endocrine disruptor | FEB 2018
42 | 233-139-2 | 10043-35-3 | boric acid | Flame retardant in paper and wood and also used paints, inks and adhesives. Used as a preservative in wood preservative, preservatives for textile, paper, leather, rubber and polymerised materials. Also used in lubricants and other materials.The substances can also be used in contact lens storage solutions and multi-purpose disinfecting solutions. | Reproductive toxin 1B | 1 | REACH SVHC | FEB 2018
43 | 234-343-4 | 11113-50-1 | boric acid | Flame retardant in paper and wood and also used paints, inks and adhesives. Used as a preservative in wood preservative, preservatives for textile, paper, leather, rubber and polymerised materials. Also used in lubricants and other materials.The substances can also be used in contact lens storage solutions and multi-purpose disinfecting solutions. | Reproductive toxin 1B | 1 |  | FEB 2018
44 | 208-953-6 | 548-62-9 | C.I. Basic Violet 3: 4-[4,4'-bis(dimethylamino) benzhydrylidene]cyclohexa-2,5-dien-1-ylidene]dimethylammonium chloride | Violet pigment | Carcinogen 1B | 1 | REACH SVHC | FEB 2018
45 | 231-152-8 | 7440-43-9 | cadmium metal | Banned by RoHS, so should be absent in electrical equipment. Also restricted by REACH but with exemptions. Used as cadmium plating fasteners, connectors and as a component of very flexible alloys and low melting point alloys. | Carcinogen 1B (mutagen 2, repro tox 2) | 1 | RoHS and REACH restricted | FEB 2018
46 | 215-147-8 | 1306-23-6 | cadmium sulphide | Banned by RoHS, so should be absent in electrical equipment. Also restricted by REACH but with exemptions. Used as a yellow pigment in plastics, ceramics and in inks. | Carcinogen 1B (mutagen 2, repro tox 2) | 1 | RoHS and REACH restricted | FEB 2018
47 | 237-366-8 | 13765-19-0 | calcium chromate | Yellow pigment, corrosion inhibitor, electroplating chemical, photochemical processing and industrial waste treatment. | Carcinogen 1B | 1 | RoHS and REACH restricted | FEB 2018
48 |  |  | Chromium (VI) compounds, with the exception of barium chromate and of compounds specified elsewhere in this Annex | Mainly yellow pigments, corrosion inhibitors | Carcinogen 1B | 1 | RoHS and REACH restricted | FEB 2018
49 | 200-755-8 | 71-48-7, 6147-53-1 | cobalt di(acetate) | Paint and varnish dryer, anodised aluminium sealing. | Carcinogen 1B, Reproductive toxin 1B (Mutagen 2) | 1 |  | FEB 2018
50 | 268-169-5 | 68016-03-5 | cobalt dimolybdenum nickel octaoxide | Yellow pigment | Carcinogen 1A | 1 |  | FEB 2018
51 | 269-051-6 | 68186-89-0 | cobalt nickel gray periclase: C.I. Pigment Black 25: C.I. 77332 | Black pigment | Carcinogen 1A | 1 |  | FEB 2018
52 | 201-557-4 | 84-74-2, 93952-11-5 | dibutyl phthalate: DBP | Plasticiser | Reproductive toxin 1B, Endocrine Disruptor | 1 | REACH SVHC. Phthalate | FEB 2018
53 | 211-670-0 | 683-18-1 | dibutyltin dichloride: (DBTC) | PVC stabiliser and polyurethane foam catalyst . Has also been an additive for rubbers. | Reproductive toxin 1B (Mutagen 2) | 1 |  | FEB 2018
54 | 201-039-8 | 77-58-7 | dibutyltin dilaurate: dibutyl[bis(dodecanoyloxy)]stannane | Dibutyltin dilaurate is used as an accelerator for solvent based, chemical cross-linking, two-component coatings like polyurethane systems. PVC stabiliser | Reproductive toxin 1B (Mutagen 2) | 1 |  | FEB 2018
55 | 201-545-9 | 84-61-7 | dicyclohexyl phthalate | Plasticiser | Reproductive toxin 1B, Endocrine Disruptor | 1 | REACH SVHC. Phthalate | JUNE 2018 (update REACH Candidate list entry)
56 | 201-559-5 | 84-75-3 | dihexyl phthalate | Plasticiser | Reproductive toxin 1B | 1 | REACH SVHC. Phthalate | FEB 2018
57 | 201-553-2 | 84-69-5 | diisobutyl phthalate | Plasticiser | Reproductive toxin 1B, Endocrine Disruptor | 1 | REACH SVHC. Phthalate | FEB 2018
58 | 210-088-4 | 605-50-5 | diisopentylphthalate | Plasticiser | Reproductive toxin 1B | 1 | REACH SVHC. Phthalate | FEB 2018
59 | 234-494-6 | 12007-01-1 | dinickel boride | Process catalyst and in hard coatings | Carcinogen 1A | 1 |  | FEB 2018
60 | 234-828-0 | 12035-64-2 | dinickel phosphide | Semiconductor used in laser diodes. Occurs in electroless nickel plating (if phosphorous reducing agent is used) | Carcinogen 1A | 1 |  | FEB 2018
61 | 215-217-8 | 1314-06-3 | dinickel trioxide | Uncommonly used as a pigment (can be black, green, blue, yellow, brown) in ceramics and glass. Also used as a process chemical | Carcinogen 1A | 1 |  | FEB 2018
62 | 205-017-9 | 131-18-0 | di-n-pentyl phthalate | Plasticiser | Reproductive toxin 1B | 1 | REACH SVHC. Phthalate | FEB 2018
63 | 251-087-9 | 32536-52-0 | diphenylether; octabromo derivate | Octabromodiphenyl ether (Octa-BDE) is a flame retardant banned by RoHS, POPs and by REACH | Reproductive toxin 1B | 1 |  | FEB 2018
64 | 240-221-1 | 16071-86-6 | disodium {}{5-[(4'-((2,6-hydroxy-3-((2-hydroxy-5-sulphophenyl)azo)phenyl)azo)(1,1'-biphenyl)-4-yl)azo]salicylato(4-)}}cuprate(2-): CI Direct Brown 95 | Brown pigment | Carcinogen 1B | 1 |  | FEB 2018
65 | 209-358-4 | 573-58-0 | disodium 3,3'-[[1,1'-biphenyl]-4,4'-diylbis(azo)]bis(4-aminonaphthalene-1-sulphonate): C.I. Direct Red 28: <i>o</i>-dianisidine based azo dyes: 4,4'-diarylazo-3,3'-dimethoxybiphenyl dyes with the exception of those mentioned elsewhere in this Annex: <i>o</i>-tolidine based dyes: 4,4'-diarylazo-3,3'-dimethylbiphenyl dyes, with the exception of those mentioned elsewhere in this Annex | Red pigment (also called Congo Red) | Carcinogen 1B (Repro tox 2) | 1 | REACH SVHC | FEB 2018
66 | 217-710-3 | 1937-37-7 | disodium 4-amino-3-[[4'-[(2,4-diaminophenyl)azo][1,1'-biphenyl]-4-yl]azo]-5-hydroxy-6-(phenylazo)naphtalene-2,7-disulphonate: C.I. Direct Black 38 | Black pigment | Carcinogen 1B (Repro tox 2) | 1 | REACH SVHC | FEB 2018
67 | 234-541-0 | 12008-41-2, 12280-03-4 | disodium octaborate anhydrous: disodium octaborate tetrahydrate | Flame retardant used in paper, card, paint, wood and also added to cellulose insulation | Reproductive toxin 1B | 1 | REACH SVHC | JUNE 2018 (included in REACH Candidate list)
68 | 215-540-4 | 1303-96-4, 1330-43-4, 12179-04-3 | disodium tetraborate decahydrate: borax decahydrate | Flame retardant used in paper, card, paint, wood and also added to cellulose insulation | Reproductive toxin 1B | 1 | REACH SVHC | FEB 2018
69 | 231-846-0 | 7758-97-6 | lead chromate | Yellow pigment (Not RoHS compliant) | Carcinogen 1B, Reproductive toxin 1A | 1 | REACH SVHC | FEB 2018
70 | 235-759-9 | 12656-85-8 | lead chromate molybdate sulfate red: C.I. Pigment Red 104: [This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77605.] | Red pigment (Not RoHS compliant) | Carcinogen 1B, Reproductive toxin 1A | 1 | REACH SVHC | FEB 2018
71 |  |  | lead compounds with the exception of those specified elsewhere in this Annex | Various uses, but most have no patient contact. Includes lead zirconium niobate (PZT) used in ultrasound transducers and piezo devices, lead titanate (used in some electronic components), lead dioxide (used corrosion resistant paint and in lead-acid batteries), as PVC stabilisers (see REACH SVHC list of lead compounds), leaded glass (e.g. optical glass), white pigments (hydoxide and carbonate are banned by REACH), other ceramics that contain lead as a binder (may be used for electronic components, sensors, electodes). | Reproductive toxin 1A | 1 | RoHS, but may be exempt, some are REACH SVHCs | FEB 2018
72 | 231-100-4 | 7439-92-1 | lead powder and metal in other forms | Lead metal is used for radiation shielding, counterweights, X-ray test objects, in power semiconductors and in alloys (e.g. brass and bronze aluminium, steel) | Reproductive toxin 1A | 1 | RoHS, but may be exempt REACH SVHC | JUNE 2018 (included in REACH Candidate list) | See also COCIR study on uses of metallic lead in medical devices (also available to MedTech Europe members)
73 | 215-693-7 | 1344-37-2 | lead sulfochromate yellow: C.I. Pigment Yellow 34: [This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77603.] | Yellow pigment (Not RoHS compliant) | Carcinogen 1B, Reproductive toxin 1A | 1 | REACH SVHC | FEB 2018
74 | 210-894-6 | 625-45-6 | methoxyacetic acid | Solvent for resins and paints. Also used as a plasticiser, in colouring agents and for textile treatment | Reproductive toxin 1B | 1 |  | FEB 2018
75 | 258-051-1 | 52625-25-9 | nickel 3,5-bis(<i>tert</i>-butyl)-4-hydroxybenzoate (1:2) | Rubber adhesives | Carcinogen 1A, Reproductive toxin 1B (Mutagen 2) | 1 |  | FEB 2018
76 | 239-086-1 | 14998-37-9 | nickel acetate | Light stabilisers in PP and PE - impart a green colour. Dye mordant for textiles, and as a sealer for anodized aluminium. | Carcinogen 1A, Reproductive toxin 1B (Mutagen 2) | 1 |  | FEB 2018
77 | 271-853-6 | 68610-24-2 | nickel barium titanium primrose priderite: C.I. Pigment Yellow 157: C.I. 77900 | Yellow pigment | Carcinogen 1A | 1 |  | FEB 2018
78 | 235-723-2 | 12619-90-8 | nickel boride | Used as a process chemical but also used as a hard coating (to replace chrome) | Carcinogen 1A | 1 |  | FEB 2018
79 | 234-493-0 | 12007-00-0 | nickel boride (NiB) | Used as a process chemical but also used as a hard coating (to replace chrome) | Carcinogen 1A | 1 |  | FEB 2018
80 | 222-068-2 | 3333-67-3 | nickel carbonate: basic nickel carbonate: carbonic acid, nickel (2+) salt | Process chemical, industrial catalyst, blue/green pigment mainly used in ceramics | Carcinogen 1A, Reproductive toxin 1B (Mutagen 2) | 1 |  | FEB 2018
81 | 206-761-7 | 373-02-4, 6018-89-9 | nickel di(acetate) | Anodised aluminium sealing. Up to 1 - 2% in thin anodised coating | Carcinogen 1A, Reproductive toxin 1B (Mutagen 2) | 1 |  | FEB 2018
82 | 234-823-3 | 12035-36-8 | nickel dioxide | Pigment (green, browns, blues, grays and yellow) used mainly in glass, ceramics, porcelain enamels and in catalysts. Used to make ferrites (but not present as NiO2). | Carcinogen 1A | 1 |  | FEB 2018
83 | 215-215-7 | 1313-99-1 | nickel monoxide | Pigment (green, browns, blues, grays and yellow) used mainly in glass, ceramics, porcelain enamels and in catalysts. Used to make ferrites (but not present as NiO). | Carcinogen 1A | 1 |  | FEB 2018
84 | 234-323-5 | 11099-02-8 | nickel oxide | Pigment (green, browns, blues, grays and yellow) used mainly in glass, ceramics, porcelain enamels and in catalysts. Used to make ferrites (but not present as NiO). | Carcinogen 1A | 1 |  | FEB 2018
85 | 235-260-6 | 12142-88-0 | nickel telluride | Used in electronic industries for special resistors and capacitors | Carcinogen 1A | 1 |  | FEB 2018
86 | 235-752-0 | 12653-76-8 | nickel titanium oxide | Yellow pigment, gas sensors, semiconductor, high density data storage, solid fuel cells and as a catalyst | Carcinogen 1A | 1 |  | FEB 2018
87 | 234-825-4 | 12035-39-1 | nickel titanium trioxide | Yellow pigment, gas sensors, semiconductor, high density data storage, solid fuel cells and as a catalyst | Carcinogen 1A | 1 |  | FEB 2018
88 | 238-032-4 | 14177-51-6 | nickel tungsten tetraoxide | Thin film coatings | Carcinogen 1A | 1 |  | FEB 2018
89 | 274-755-1 | 70692-93-2 | nickel zirkonium trioxide | Pigment/ceramic | Carcinogen 1A | 1 |  | FEB 2018
90 | 206-400-3 | 335-76-2 | nonadecafluorodecanoic acid | Main uses is for waterproofing, may have been used to make plasticisers (very unlikely), lubricants, wetting agents, surfactants and corrosion inhibitors | Reproductive toxin 1B (mutagen 2) | 1 | REACH SVHC | FEB 2018
91 | 933-378-9 | 776297-69-9 | n-pentyl-isopentylphthalate | Plasticiser | Reproductive toxin 1B | 1 |  | FEB 2018
92 | 271-112-7 | 68515-84-4 | olivine, nickel green | Green pigment | Carcinogen 1A | 1 |  | FEB 2018
93 | 237-560-2 | 13840-56-7 | orthoboric acid, sodium salt | Sodium borate, borax, preservative and flame retardant. | Reproductive toxin 1B | 1 |  | FEB 2018
94 | 206-801-3 | 375-95-1 | perfluorononan-1-oic acid | Also called PFNA, used as a processing aid for manufacture of fluoropolymers such as PVDF. Also used as an additive in lubricating oils, in cleaning agents, for waterproofing and textile coatings and in liquid crystal displays. Occurs in water and stain resistant textiles. Will not exceed 0.1% in a material, but could as an ultra-thin waterproof coating | Reproductive toxin 1B (carcinogen 2) | 1 | REACH SVHC | FEB 2018
95 |  | 4149-60-4 | perfluorononan-1-oic acid ammonium salts | Used as a processing aid for manufacture of fluoropolymers such as PVDF. Also used as an additive in lubricating oils, in cleaning agents, for waterproofing and textile coatings and in liquid crystal displays. Occurs in water and stain resistant textiles. Will not exceed 0.1% in a material, but could as an ultra-thin waterproof coating | Reproductive toxin 1B (carcinogen 2) | 1 | REACH SVHC | FEB 2018
96 |  | 21049-39-8 | perfluorononan-1-oic acid sodium salts | Used as a processing aid for manufacture of fluoropolymers such as PVDF. Also used as an additive in lubricating oils, in cleaning agents, for waterproofing and textile coatings and in liquid crystal displays. Occurs in water and stain resistant textiles. Will not exceed 0.1% in a material, but could as an ultra-thin waterproof coating | Reproductive toxin 1B (carcinogen 2) | 1 | REACH SVHC | FEB 2018
97 | 217-179-8 | 1763-23-1 | perfluorooctane sulfonic acid: heptadecafluorooctane-1-sulfonic acid | PFOS - banned | Reproductive toxin 1B (carcinogen 2) | 1 | EU POPs Regulation | FEB 2018
98 | 206-397-9 | 335-67-1 | perfluorooctanoic acid | PFOA - water and stain resistant fabrics - applied as very thin coatings | Reproductive toxin 1B (carcinogen 2) | 1 | REACH SVHC | FEB 2018
99 | 232-373-2 | 8009-03-8 | Petrolatum: Petrolatum: [A complex combination of hydrocarbons obtained as a semi-solid from dewaxing paraffinic residual oil. It consists predominantly of saturated crystalline and liquid hydrocarbons having carbon numbers predominantly greater than C<sub>25.</sub>] | Also called petroleum jelly, mainly has medical uses, but is used as a lubricant. Petrolatums with different CAS numbers not likely to be used | Carcinogen 1B | 1 |  | FEB 2018
100 | 204-557-2 | 122-60-1 | phenyl glycidyl ether: 2,3-epoxypropyl phenyl ether: 1,2-epoxy-3-phenoxypropane | Process chemical for epoxy resins but also reported used as a stabiliser for various polymers (vinyl, epoxy, rubber), likely to be uncommon | Carcinogen 1B (mutagen 2) | 1 |  | FEB 2018
101 | 220-527-1 | 2795-39-3 | potassium perfluorooctanesulfonate: potassium heptadecafluorooctane-1-sulfonate | A PFOS salt, used as a fluorosurfactant in stain and water-proof coatings for fabrics and textiles. Also added to waxes, polishes, paints, varnishes, and cleaning products | Reproductive toxin 1B (carcinogen 2) | 1 |  | FEB 2018
102 | 253-461-7 | 37321-15-6 | silicic acid, nickel salt | Green pigment | Carcinogen 1A | 1 |  | FEB 2018
103 |  | 3830-45-3 | sodium nonadecafluorodecanoate | Main uses is for waterproofing, may have been used to make plasticisers (very unlikely), lubricants, wetting agents, surfactants and corrosion inhibitors. Unlikely to exceed 0.1% | Reproductive toxin 1B (carcinogen 2) | 1 | REACH SVHC | FEB 2018
104 | 232-142-6 | 7789-06-2 | strontium chromate | Paint primers (yellow) | Carcinogen 1B | 1 | REACH SVHC | FEB 2018
105 | 235-541-3 | 12267-73-1 | tetraboron disodium heptaoxide, hydrate | Also known as borax, has many uses including in household cleaning products, in steel welding fluxes, for silver soldering, as an anti-fungal agent (e.g. in wood), as a fire retardant for cellulose, paper and wood, in welding fluxes, in brake fluids, metal working fluids, in antifreeze as a corrosion inhibitor and to make other boron compounds. It is also used in starch adhesives used in corrugated cardboard packaging | Reproductive toxin 1B | 1 | REACH SVHC | FEB 2018
106 | 220-012-1 | 2602-46-2 | tetrasodium 3,3'-[[1,1'-biphenyl]-4,4'-diylbis(azo)]bis[5-amino-4-hydroxynaphthalene-2,7-disulphonate]: C.I. Direct Blue 6 | Blue pigment | Carcinogen 1B (repro tox 2) | 1 |  | FEB 2018
107 | 231-205-5 | 7446-27-7 | trilead bis(orthophosphate) | PVC (and rarely other polymer) stabiliser, used to make special glasses (but not present as trilead bis(orthophosphate)), some medicinal uses. | Reproductive toxin 1A | 1 | RoHS | FEB 2018
108 | 234-495-1 | 12007-02-2 | trinickel boride | Used as a process chemical but also used as a hard coating (to replace chrome) | Carcinogen 1A | 1 |  | FEB 2018
109 | 204-118-5 | 115-96-8 | tris(2-chloroethyl)phosphate | SVHC: Flame retardant and additive plasticizer used in various polymers including rigid polyurethane and flexible polyurethane foam, textile back coatings, PVC (uncommon), cellulose esters and polyester resin | Reproductive toxin 1B (carcinogen 2) | 1 | REACH SVHC | FEB 2018
110 | 413-590-3 | 164058-22-4 | trisodium [4'-(8-acetylamino-3,6-disulfonato-2-naphthylazo)-4''-(6-benzoylamino-3-sulfonato-2-naphthylazo)-biphenyl-1,3',3'',1'''-tetraolato-<i>O</i>,<i>O</i>',<i>O</i>'',<i>O</i>''']copper(II) | Believed to be used as a pigment (colour unknown) | Carcinogen 1B | 1 |  | FEB 2018
111 | 246-677-8 | 25155-23-1 | trixylyl phosphate | Fire retardant for many types of plastics, fire retardant in hydraulic fluids (e.g. for industrial generators). | Reproductive toxin 1B | 1 | REACH SVHC | FEB 2018
112 | 213-103-2 | 924-42-5 | N-(hydroxymethyl)acrylamide; methylolacrylamide; [NMA] | N-methylol acrylamide is a crosslinkable monomer widely used as in the manufacture of polymer products (adhesives, binders, surface coatings, resins and thermoplastics). N-methylolacrylamide is a bi-functional monomer with reactive vinyl and hydroxyethyl groups. Thermoplastic polymers can be formed by copolymerization of N-methylol-acrylamide with a variety of vinyl monomers by emulsion, solution and suspension techniques. The resulting products, which have pendant hydroxyethyl groups, can undergo cross-linking under moderate conditions, permitting conversion of thermoplastic backbone polymers to thermoset materials at the point of use in the absence of an external cross-linking agent. Conversely, the hydroxyethyl group can be reacted with a substrate like cellulose and subsequently cross-linked by free-radical polymerization. N-methylol acrylamide is also used as an intermediate in the production of other substances. N-methylol acrylamide is reacted in all of these uses and will not be present in supplied articles > 0.1% w/w | Carc. 1B Muta. 1B | 2 |  | OCT 2020
113 | 213-934-0 | 1067-53-4 | tris(2-methoxyethoxy)vinylsilane; 6-(2-methoxyethoxy)-6-vinyl-2,5,7,10-tetraoxa-6-silaundecane | tris(2-methoxyethoxy)vinylsilane, an organosilane, is used as a coupling agent and adhesion promoter in filled polymer systems to improve mechanical and electrical properties, particularly after wet conditioning. It can also be used as an intermediate and laboratory chemical. tris(2-methoxyethoxy)vinylsilane can be used as a co-monomer for the preparation of different polymers such as polyethylene or acrylics. A major application of polymers containing this substance is in wires, cables and pipes. However, tris(2-methoxyethoxy)vinylsilane will be reacted in all of the above applications and so is not likely to be found > 0.1% w/w in the supplied article. Coating used for outdoor decking - 0.3% w/w https://www.tor-coatings.com/wp-content/uploads/2020/01/RC5490001-Elastadeck-Walkway-Gritcoat-Base-gb.pdf | Repr. 1B | 2 |  | OCT 2020
114 | 205-886-4 | 191-30-0 | dibenzo[def,p]chrysene dibenzo[a,l]pyrene | dibenzo[def,p]chrysene,dibenzo[a,l]pyrene is a Polycyclic Aromatic Hydrocarbon (PAH) which can be present as an impurity in plastic and rubber parts, because they are present as impurities in some of the raw materials used in the production of these parts, in particular in extender oils and in carbon black. PAHs are not added intentionally to the articles and do not perform any specific function as constituents of the plastic or rubber parts. Extender oils can be used as plasticizers in rubber and plastics or in potting material used in electronic ballasts. PAHs may be present as contaminant in these oils. Black pigments and dyes used in rubber and plastic (e.g. Carbon Black) may contain PAH contaminated soot. In all of these applications the concentration of PAH in the supplied articles is typically below 0.02% and will always be considerably less than 0.1%. | Carc. 1B | 2 |  | OCT 2020
115 | 221-967-7 | 3296-90-0 | 2,2-bis(bromomethyl)propane-1,3-diol | 2,2-bis(bromomethyl)propane-1,3-diol, also known as BMP is a reactive brominated flame retardant that is used in the manufacture of unsaturated polyester resins for moulded products and rigid polyurethane foam with a typical addition rate of 1-5%. (https://www.industrialchemicals.gov.au/sites/default/files/1%2C3-Propanediol%2C%202%2C2-bis%28bromomethyl%29-_Human%20health%20tier%20II%20assessment.pdf and https://www.ncbi.nlm.nih.gov/books/NBK390857/). It can also be used as an intermediate in the production of other flame retardants. Reactive flame retardants are added during the polymerisation process and become an integral part of the polymer. The result is a modified polymer with flame retardant properties and different molecular structure compared to the original polymer molecule. As a result BMP is not present as a substance in the manufactured polymer > 0.1% w/w. BMP is also used as a reactive flame retardant in waterproofing coatings such as GACOFLEX® U-64 LIQUID POLYURETHANE COATING which is a two-part liquid polyurethane coating which contains 1-5% BMP (https://gaco.com/wp-content/uploads/U6402P-PEWTER-URE-CTG-POLY-SIDE-GHS04212015.pdf) and FIRESTONE® REGULAR URETHANE BLACK BASECOAT which also contains 1-5% BMP (https://41z32e1503dj3gyqe3151089-wpengine.netdna-ssl.com/wp-content/uploads/UB6421PC-%E2%80%93-REGULAR-URETHANE-BLACK-BASECOAT-VOC-COMPLIANT-POLYOL-COMPONENT-A-GHS07162018.pdf). As mentioned above, reactive flame retardants become an integral part of the polymer and therefore are not present in the coated material > 0.1% w/w. | Carc. 1B Muta. 1B | 2 |  | OCT 2020
116 | 202-987-5 | 101-90-6 | m-bis(2,3-epoxypropoxy)benzene; resorcinol diglycidyl ether | resorcinol diglycidyl ether is a reactive diluent and modifier in epoxy resins (used in coatings, potting and tooling compounds, adhesives and composites) and is also used as a curing agent in polysulfide rubber. resorcinol diglycidyl ether is reacted or cured in all of these applications and so will not be present >0.1% w/w in a supplied article. | Carc. 1B | 2 |  | NOV 2020
117 | 202-496-6 | 96-29-7 | butanone oxime; ethyl methyl ketoxime; ethyl methyl ketone oxime | butanone oxime, known as MEKO in the paint industry, is widely used as an anti-skinning agent to prevent oxidative drying in the container of alkyd (oil-based) paints, primers, varnishes and other coatings that can be used in consumer or industrial applications. Concentrations of butanone oxime in paints is up to 1% w/w and 0.1-0.2% w/w in colorants or pigments used in printing inks. Other reported uses of butanone oxime include as an ingredient in pesticide products and wood preservatives, as a chemical intermediate and as a corrosion inhibitor in industrial boilers and water treatement systems. None of the above applications will result in the substance being present >0.1% w/w in a supplied article. When used in paints, the substance will evapourate during the drying process. butanone oxime can also be used in silicone sealants particularly room temperture vulcanisation (RTV) silicone. Example addition rates of butnanone oxime in silicone sealants include 0.5-3% w/w. However, in this application butanone oxime evaporates during the curng process and so will not be present in a supplied article >0.1% w/w. Canada Gov Study - https://www.ec.gc.ca/ese-ees/default.asp?lang=En&n=32AD1FD8-1 Primer 0.3% w/w - https://www.crowntrade.co.uk/wp-content/uploads/2014/12/Crown-Trade-Red-Oxide-Primer-CLP-SDS-180417.pdf Gloss paint 0.3-1% w/w - https://www.hammerite.co.uk/files/2017/02/HM_GB_EN_RADIATOR_PAINT_GLOSS.pdf Coating <1% w/w - https://marinecoatings.brand.akzonobel.com/m/6e00da8714a4f5f4/original/CLT229_0010eu_UA_EN_20170602_1.pdf Silicate sealant 0.5-1% w/w - https://www.techsil.co.uk/media/pdf/SDS/TBSI19058-sds.pdf Silicone sealant 1-3 % w/w - https://docs.rs-online.com/d07b/0900766b8134b277.pdf Study of emission of butanone oxime from sealants - https://www.sciencedirect.com/science/article/pii/S0026265X2031136X | Carc. 1B | 2 |  | NOV 2020
118 | 201-279-3 | 80-43-3 | bis(α,α-dimethylbenzyl) peroxide | bis(α,α-dimethylbenzyl) peroxide also known as Dicumyl Peroxide is an organic peroxide used as a cross-linking agent in the maufacture of polymers and elastomers. Typical applications of Dicumyl Peroxide as cross linking agent include the vulcinaization of natural and synthetic rubbers and in the production of polyethylene and polymers used to produce hose, wires, rubber seals etc. In all of these applications Dicumyl Peroxide will be reacted during the curing process. Dicumyl Preoxide is also used as a flame retardant synergist in expanded polystyrene (EPS) where it is added in small quantities alongside halogonated flame retartardants to act as catalyst for the flame retardant action in case of heating. As it is an additive it may be present in supplied articles. However, EPS will not be used in medical devices in patient contact materials or gas, liquid or solid transfer materials. | Repr. 1B | 2 |  | NOV 2020
119 | 205-877-5 | 189-55-9 | benzo[rst]pentaphene | benzo[rst]pentaphene is a Polycyclic Aromatic Hydrocarbon (PAH) which can be present as an impurity in plastic and rubber parts, because they are present as impurities in some of the raw materials used in the production of these parts, in particular in extender oils and in carbon black. PAHs are not added intentionally to the articles and do not perform any specific function as constituents of the plastic or rubber parts. Extender oils can be used as plasticizers in rubber and plastics or in potting material used in electronic ballasts. PAHs may be present as contaminant in these oils. Black pigments and dyes used in rubber and plastic (e.g. Carbon Black) may contain PAH contaminated soot. In all of these applications the concentration of PAH in the supplied articles is typically below 0.02% and will always be considerably less than 0.1%. | Carcinogen 1B | 2 |  | APRIL 2020
120 | 205-878-0 | 189-64-0 | dibenzo[b,def]chrysene; dibenzo[a,h]pyrene | dibenzo[b,def]chrysene is a Polycyclic Aromatic Hydrocarbon (PAH) which can be present as an impurity in plastic and rubber parts, because they are present as impurities in some of the raw materials used in the production of these parts, in particular in extender oils and in carbon black. PAHs are not added intentionally to the articles and do not perform any specific function as constituents of the plastic or rubber parts. Extender oils can be used as plasticizers in rubber and plastics or in potting material used in electronic ballasts. PAHs may be present as contaminant in these oils. Black pigments and dyes used in rubber and plastic (e.g. Carbon Black) may contain PAH contaminated soot. In all of these applications the concentration of PAH in the supplied articles is typically below 0.02% and will always be considerably less than 0.1%. | Carcinogen 1B | 2 |  | APRIL 2020
121 | 203-772-9 | 110-49-6 | 2-methoxyethyl acetate: methylglycol acetate | Not registered under REACH; Solvent and intermediate; may be used in medical devices but only at < 0.1% This substance may be used as a reaction solvent for gums, resins, waxes, oils, greases, lacquers, paints, inks, adhesives and sealants. The typical concentration of EGMEA in these applications is 1 – 5%. The concentration of EGMEA is always less than 10% of the printing inks, sealants, adhesives, etc which in turn will represent considerably less than 1% of the weight of the supplied article. Therefore, the presence of EGMEA in printing inks, sealants, adhesives, etc will not lead to a concentration of more than 0.1% w/w of EGMEA in the supplied article. | Repr 1B | 2 | REACH SVHC | JUL 2019
122 | 202-506-9 | 96-45-7 | ethylene thiourea: imidazolidine-2-thione: 2-imidazoline-2-thiol | Assessment by BOMcheck and IEC 62474 identifies that Imidazolidine-2-thione is used as a catalyst in some acrylic adhesive glues which may be used in adhesive tapes (for example, double sided adhesive tapes which may be used to hold the back-light in place in mobile phones). The adhesive glues are generally not cured at high temperatures. As a result, imidazolidine-2-thione may be detectable as a substance > 0.1% w/w in acrylic adhesive glue. This substance is used as an accelerator in rubber vulcanising as the process with sulphur is extremely slow and impractical. Publications indicate that ethylene thiourea behaves as a catalyst and so, although it does take part in and is consumed by the chemical reactions, some is reformed and so this substance can remain (at >0.1%) in vulcanised rubber. Therefore, this substance may be present in devices in concentration > 0,1%. | Reproductive toxin 1B | 2 | REACH SVHC | APRIL 2019 (included in REACH Candidate List: December 2013)Assessment updated Septembe 2020
123 | 200-028-5 | 50-32-8, 63466-71-7 | benzo[a]pyrene: benzo[def]chrysene | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1% | Carcinogen 1B, Reproductive toxin 1B, Mutagen 1B | 2 | REACH SVHC | APRIL 2019 (included in REACH Candidate List: June 2016)
124 | 200-280-6 | 56-55-3, 1718-53-2 | benz[a]anthracene | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1% | Carcinogen 1B | 2 | REACH SVHC | APRIL 2019 (included in REACH Candidate List: January 2018)
125 | 205-892-7 | 192-97-2 | benzo[e]pyrene | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1% | Carcinogen 1B | 2 | Restricted under entry 50 of Annex XVII of the REACH Regulation in rubber or plastic components that come into direct, prolonged or repetitive skin or oral cavity contact | APRIL 2019
126 | 205-910-3 | 205-82-3 | benzo[j]fluoranthene | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1% | Carcinogen 1B | 2 | Restricted under entry 50 of Annex XVII of the REACH Regulation in rubber or plastic components that come into direct, prolonged or repetitive skin or oral cavity contact | APRIL 2019
127 | 205-911-9 | 205-99-2 | benz[e]acephenanthrylene | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1% | Carcinogen 1B | 2 | Restricted under entry 50 of Annex XVII of the REACH Regulation in rubber or plastic components that come into direct, prolonged or repetitive skin or oral cavity contact | APRIL 2019
128 | 200-181-8 | 53-70-3 | dibenz[a,h]anthracene | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1% | Carcinogen 1B | 2 | Restricted under entry 50 of Annex XVII of the REACH Regulation in rubber or plastic components that come into direct, prolonged or repetitive skin or oral cavity contact | APRIL 2019
129 | 205-916-6 | 207-08-9 | Benzo[k]fluoranthene | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1% | Carcinogen 1B | 2 | REACH SVHC | APRIL 2019 (included in REACH Candidate list: January 2019)
130 | 221-221-0 | 3033-77-0 | 2,3-epoxypropyltrimethylammonium chloride ...%: glycidyl trimethylammonium chloride ...% | Cationisation of starches which are added in paper to give paper better surface quality. Used reactively so that only up to 25mg/kg remain, so <0.1% | Carcinogen 1B (mutagen 2, repro tox 2) | 2 |  | FEB 2018
131 | 203-924-4 | 111-96-6 | bis(2-methoxyethyl) ether | Solvent for battery electrolytes (lithium ion and other). May also be used as a component of inks, but will not be present in printed material. | Reproductive toxin 1B | 2 | REACH SVHC | FEB 2018
132 | 215-146-2 | 1306-19-0 | cadmium oxide (non-pyrophoric) | Power contacts (RoHS exempt), but no patient contact | Carcinogen 1B (mutagen 2, repro tox 2) | 2 | REACH SVHC | FEB 2018
133 | 205-923-4 | 218-01-9, 1719-03-5 | chrysene | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1% | Carcinogen 1B (mutagen 2) | 2 | REACH SVHC | FEB 2018
134 | 261-346-8 | 58591-45-0 | cobalt nickel dioxide | Reported to be used in supercapacitors | Carcinogen 1A | 2 |  | FEB 2018
135 | 620-395-9 | 12737-30-3 | cobalt nickel oxide | Reported to be used in supercapacitors | Carcinogen 1A | 2 |  | FEB 2018
136 | 215-114-8 | 1303-00-0 | gallium arsenide | Semiconductor used inside ICs and LEDs. Also used in solar cells Always encapsulated so that physical contact should not be possible | Carcinogen 1B, Reproductive toxin 1B | 2 |  | FEB 2018
137 | 244-959-5 | 22398-80-7 | indium phosphide | Semiconductor used in laser diodes and high frequency electronics - no patient contact | Carcinogen 1B (Reproductive toxin 2) | 2 |  | FEB 2018
138 | 215-630-3 | 1335-32-6 | lead acetate, basic | MRI/CT contrast agent, but should be no patient exposure, unless leaks occur. | Reproductive toxin 1A (carcinogen 2) | 2 |  | FEB 2018
139 | 206-104-4 | 301-04-2, 6080-56-4 | lead di(acetate) | MRI/CT contrast agent, but should be no patient exposure, unless leaks occur. | Reproductive toxin 1A | 2 | REACH SVHC | FEB 2018
140 | 620-400-4 | 12031-65-1 | lithium nickel dioxide | Lithium ion battery anode | Carcinogen 1A | 2 |  | FEB 2018
141 | 231-106-7 | 7439-97-6 | mercury | Measurement instruments such as sphygmomanometers, although most uses are now banned | Reproductive toxin 1B | 2 |  | FEB 2018
142 | 220-250-6 | 2687-91-4 | N-ethyl-2-pyrrolidone: 1-ethylpyrrolidin-2-one | Solvent in photoresist stripper formulations. Ethylpyrrolidone can be used in electrolyte formulations for electrolytic capacitors (according to Eastman who are a manufacturer), ink-jet printing ink solvent, liquid crystal cells, solvent and in pharmaceuticals. | Reproductive toxin 1B | 2 |  | FEB 2018
143 | 224-699-9 | 4454-16-4 | nickel bis(2-ethylhexanoate) | Catalyst used to make rubber but likely to be <0.1% | Carcinogen 1A | 2 |  | FEB 2018
144 | 235-379-3 | 12201-89-7 | nickel disilicide | Semiconductor memory and other devices | Carcinogen 1A | 2 |  | FEB 2018
145 | 257-970-5 | 52502-12-2 | nickel divanadium hexaoxide | Electrochromic material. Also used as photocell electrode layer | Carcinogen 1A | 2 |  | FEB 2018
146 | 248-585-3 | 27637-46-3 | nickel isooctanoate | Catalyst used to make rubber but likely to be <0.1% | Carcinogen 1A | 2 |  | FEB 2018
147 | 234-824-9 | 12035-38-0 | nickel tin trioxide: nickel stannate | Additive in ceramic capacitors, may be used for electroplating | Carcinogen 1A | 2 |  | FEB 2018
148 | 287-468-1 | 85508-43-6 | nickel(II) isodecanoate | Catalyst used to make rubber but likely to be <0.1% | Carcinogen 1A | 2 |  | FEB 2018
149 | 249-555-2 | 29317-63-3 | nickel(II) isooctanoate | Catalyst used to make rubber but likely to be <0.1% | Carcinogen 1A | 2 |  | FEB 2018
150 | 287-469-7 | 85508-44-7 | nickel(II) neodecanoate | Catalyst used to make rubber but likely to be <0.1% | Carcinogen 1A | 2 |  | FEB 2018
151 | 300-094-6 | 93920-10-6 | nickel(II) neononanoate | Catalyst used to make rubber but likely to be <0.1% | Carcinogen 1A | 2 |  | FEB 2018
152 | 300-093-0 | 93920-09-3 | nickel(II) neoundecanoate | Catalyst used to make rubber but likely to be <0.1% | Carcinogen 1A | 2 |  | FEB 2018
153 | 225-656-7 | 4995-91-9 | nickel(II) octanoate | Catalyst used to make rubber but likely to be <0.1% | Carcinogen 1A | 2 |  | FEB 2018
1 | 211-076-1 | 629-14-1 | 1,2-diethoxyethane | Ethylene glycol diethyl ether is a important solvent, used for ink, paint and coating industry. Mainly used in polymer, electrochemistry, boracium chemistry, resin, nitro cellulose, it is also used in surface treatment, halogen analysis in gasoline, acetic acid recycle in diluted acetic acid, depaint agent, thinner, flushing agent, stabilizer, anti oxidant , thickener of lube. Boiling point 120°C so it is not likely to be present in polymers or coatings made using this solvent. May be an additive in hydraulic and brake fluids. | Reproductive toxin 1A | 3 | REACH SVHC | FEB 2018
2 | 202-491-9 | 96-23-1 | 1,3-dichloro-2-propanol | Process chemical, but can occur as an impurity breakdown product in paper and paints https://monographs.iarc.fr/ENG/Monographs/vol101/mono101-011.pdf | Carcinogen 1B | 3 |  | FEB 2018
3 | 201-549-0 | 84-65-1 | anthraquinone | Mainly as an intermediate to make dyes and pharmaceutocals. Uncommonly used as an organic scintillator in radiation detectors. Patient contact is very unlikely | Carcinogen 1B | 3 |  | FEB 2018
4 | 215-133-1 | 1304-56-9 | beryllium oxide | No patient contact. Beryllium oxide is used as a Beryllium oxide ceramics having excellent electrical insulation properties and a high thermal conductivity and is used in some power semiconductors. One report lists uses as guide bores of laser scalpels and parts of MRI and CAT scanners. | Carcinogen 1B | 3 |  | FEB 2018
5 | 231-589-4 | 7646-79-9, 7791-13-1 | cobalt dichloride | Moisture indicator cards, uncommon ink drier | Carcinogen 1B, Reproductive toxin 1B (mutagen 2) | 3 | REACH SVHC | FEB 2018
6 | 200-001-8 | 50-00-0 | formaldehyde …% | Water soluble gas used as a process chemical. Can be generated by degradation of some types of polymers. | Carcinogen 1B, (mutagen 2) | 3 |  | FEB 2018
7 | 278-011-7 | 74869-21-9 | Lubricating greases: Grease: [A complex combination of hydrocarbons having carbon numbers predominantly in the range of C<sub>12</sub> through C<sub>50.</sub> May contain organic salts of alkali metals, alkaline earth metals, and/or aluminium compounds.] | May be used in medical devices although unlikely to be patient contact | Carcinogen 1B | 3 |  | FEB 2018
8 | 297-474-6 | 93572-43-1 | Lubricating oils (petroleum), base oils, paraffinic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by refining of crude oil. It consists predominantly of aromatics, naphthenics and paraffinics and produces a finished oil with a viscosity of 120 SUS at 100 °F (23cSt at 40 °C).] | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
9 | 309-874-0 | 101316-69-2 | Lubricating oils (petroleum), C <sub>>25</sub>, solvent-extd., deasphalted, dewaxed, hydrogenated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by solvent extraction and hydrogenation of vacuum distillation residues. It consists predominantly of hydrocarbons having carbon numbers predominantly greater than C<sub>25</sub> and produces a finished oil with a viscosity in the order of 32cSt to 37cSt at 100 °C (212 °F).] | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
10 | 276-737-9 | 72623-86-0 | Lubricating oils (petroleum), C<sub>15-30</sub>, hydrotreated neutral oil-based: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treating light vacuum gas oil and heavy vacuum gas oil with hydrogen in the presence of a catalyst in a two stage process with dewaxing being carried out between the two stages. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil having a viscosity of approximately 15cSt at 40 °C. It contains a relatively large proportion of saturated hydrocabons.] | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
11 | 309-875-6 | 101316-70-5 | Lubricating oils (petroleum), C<sub>17-32</sub>, solvent-extd., dewaxed, hydrogenated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by solvent extraction and hydrogenation of atmospheric distillation residues. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>17</sub> through C<sub>32</sub> and produced a finished oil with a viscosity in the order of 17cSt to 23cSt at 40 °C (104 °F.] | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
12 | 295-423-2 | 92045-42-6 | Lubricating oils (petroleum), C<sub>17-35</sub>, solvent-extd., dewaxed, hydrotreated: Baseoil - unspecified | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
13 | 307-034-8 | 97488-95-4 | Lubricating oils (petroleum), C<sub>18-27</sub>, hydrocracked solvent-dewaxed: Baseoil - unspecified | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
14 | 305-594-8 | 94733-15-0 | Lubricating oils (petroleum), C<sub>18-40</sub>, solvent-dewaxed hydrocracked distillate-based: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by solvent deparaffination of the distillation residue from hydrocracked petroleum. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>18</sub> through C<sub>40</sub> and boiling in the range of approximately 370 °C to 550 °C (698 °F to 1022 °F).] | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
15 | 305-595-3 | 94733-16-1 | Lubricating oils (petroleum), C<sub>18-40</sub>, solvent-dewaxed hydrogenated raffinate-based: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by solvent deparaffination of the hydrogenated raffinate obtained by solvent extraction of a hydrotreated petroleum distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>18</sub> through C<sub>40</sub> and boiling in the range of approximately 370 °C to 550 °C (698 °F to 1022 °F).] | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
16 | 309-876-1 | 101316-71-6 | Lubricating oils (petroleum), C<sub>20-35</sub>, solvent-extd., dewaxed, hydrogenated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by solvent extraction and hydrogenation of atmospheric distillation residues. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>35</sub> and produces a finished oil with a viscosity in the order of 37cSt to 44cSt at 40 °C (104 °F).] | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
17 | 276-736-3 | 72623-85-9 | Lubricating oils (petroleum), C<sub>20-50</sub>, hydrotreated neutral oil-based, high-viscosity: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treating light vacuum gas oil, heavy vacuum gas oil, and solvent deasphalted residual oil with hydrogen in the presence of a catalyst in a two stage process with dewaxing being carried out between the two stages. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil having a viscosity of approximately 112cSt at 40 °C. It contains a relatively large proportion of saturated hydrocarbons.] | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
18 | 276-738-4 | 72623-87-1 | Lubricating oils (petroleum), C<sub>20-50</sub>, hydrotreated neutral oil-based: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treating light vacuum gas oil, heavy vacuum gas oil and solvent deasphalted residual oil with hydrogen in the presence of a catalyst in a two stage process with dewaxing being carried out between the two stages. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of approximately 32cSt at 40 °C. It contains a relatively large proportion of saturated hydrocarbons.] | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
19 | 309-877-7 | 101316-72-7 | Lubricating oils (petroleum), C<sub>24-50</sub>, solvent-extd., dewaxed, hydrogenated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by solvent extraction and hydrogenation of atmospheric distillation residues. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>24</sub> through C<sub>50</sub> and produces a finished oil with a viscosity in the order of 16cSt to 75cSt at 40 °C (104 °F).] | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
20 | 295-424-8 | 92045-43-7 | Lubricating oils (petroleum), hydrocracked nonarom. solvent-deparaffined: Baseoil - unspecified | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
21 | 278-012-2 | 74869-22-0 | Lubricating oils: Baseoil - unspecified: [A complex combination of hydrocarbons obtained from solvent extraction and dewaxing processes. It consists predominantly of saturated hydrocarbons having carbon numbers in the range C<sub>15</sub> through C<sub>50</sub>.] | Lubricants | Carcinogen 1B | 3 |  | FEB 2018
22 | 231-906-6 | 7778-50-9 | potassium dichromate | Process chemical and used for sealing anodised aluminium (to prevent dye leakage). Would not comply with RoHS and requires REACH authorisation if used in the EU. | Carcinogen 1B, Reproductive toxin 1B, Mutagen 1B | 3 | REACH SVHC | FEB 2018
23 | 234-190-3 | 7789-12-0, 10588-01-9 | sodium dichromate | Process chemical and aluminium anodised coating sealant (to prevent dye leakage) | Carcinogen 1B, Reproductive toxin 1B, Mutagen 1B | 3 | REACH SVHC | FEB 2018
24 | 402-060-7 | 108225-03-2 | (6-(4-hydroxy-3-(2-methoxyphenylazo)-2-sulfonato-7-naphthylamino)-1,3,5-triazin-2,4-diyl)bis[(amino-1-methylethyl)ammonium] formate | Registered by one company (Archroma Germany GmbH) who are manufacturers of colours for textiles, paper, coatings (paint), adhesives and sealants | Carcinogen 1B | 3 |  | FEB 2018
1 | 279-452-8 | 80387-97-9 | 2-ethylhexyl[[[3,5-bis(1,1-dimethylethyl)-4-hydroxyphenyl]methyl]thio]acetate | Unknown, not registered in the EU. Likely to be used as a process or laboratory chemical | Reproductive toxin 1B | 4 |  | FEB 2018
2 | 412-790-8 |  | reaction mass of: <i>N</i>-[3-hydroxy-2-(2-methylacryloylaminomethoxy)propoxymethyl]-2-methylacrylamide: <i>N</i>-[2,3-bis-(2-methylacryloylaminomethoxy)propoxymethyl]-2-methylacrylamide: methacrylamide: 2-methyl-<i>N</i>-(2-methylacryloylaminomethoxymethyl)-acrylamide: <i>N</i>-(2,3-dihydroxypropoxymethyl)-2-methylacrylamide | Uncertain but one reference indicates that it is used in cosmetics, so not likely to be applicable toi medical devices | Carcinogen 1B (mutagen 2) | 4 |  | FEB 2018
3 | 421-550-1 |  | reaction mass of: 1,3,5-tris(3-aminomethylphenyl)-1,3,5-(1<i>H</i>,3<i>H</i>,5<i>H</i>)-triazine-2,4,6-trione: reaction mass of oligomers of 3,5-bis(3-aminomethylphenyl)-1-poly[3,5-bis(3-aminomethylphenyl)-2,4,6-trioxo-1,3,5-(1<i>H</i>,3<i>H</i>,5<i>H</i>)-triazin-1-yl]-1,3,5-(1<i>H</i>,3<i>H</i>,5<i>H</i>)-triazine-2,4,6-trione | Uncertain, registration dossier for 1,3,5-tris(oxiranylmethyl)-1,3,5-triazine-2,4,6(1H,3H,5H)-trione indicates that this is a process chemical only, used to make powder coatings, solder resist inks and special electronic tapes. Reacts in use so would not occur in products. | Carcinogen 1B, Reproductive toxin 1B | 4 |  | FEB 2018
4 | 403-250-2 |  | reaction mass of: 4-[[bis-(4-fluorophenyl)methylsilyl]methyl]-4<i>H</i>-1,2,4-triazole: 1-[[bis-(4-fluorophenyl)methylsilyl]methyl]-1<i>H</i>-1,2,4-triazole | Uncertain, but may be used as a fungicide | Reproductive toxin 1B (carcinogen 2) | 4 |  | FEB 2018
5 | 435-960-3 |  | reaction mass of: dimethyl (2-(hydroxymethylcarbamoyl)ethyl)phosphonate: diethyl (2-(hydroxymethylcarbamoyl)ethyl)phosphonate: methyl ethyl (2-(hydroxymethylcarbamoyl)ethyl)phosphonate | Unknown, but indication that this may be used to give special effects to textiles | Carcinogen 1B, Mutagen 1B | 4 |  | FEB 2018
6 | 402-660-9 |  | reaction mass of: disodium 4-(3-ethoxycarbonyl-4-(5-(3-ethoxycarbonyl-5-hydroxy-1-(4-sulfonatophenyl)pyrazol-4-yl)penta-2,4-dienylidene)-4,5-dihydro-5-oxopyrazol-1-yl)benzenesulfonate: trisodium 4-(3-ethoxycarbonyl-4-(5-(3-ethoxycarbonyl-5-oxido-1-(4-sulfonatophenyl)pyrazol-4-yl)penta-2,4-dienylidene)-4,5-dihydro-5-oxopyrazol-1-yl)benzenesulfonate | Unknown. Registered for REACH by one legal entity (not identified) as a new substance, but uses not specified. Banned in cosmetics, which indicates an old use. | Reproductive toxin 1B | 4 |  | FEB 2018
| 206-019-2 | 288-32-4 | imidazole | Corrosion inhibitor (in water) and process chemical used to make imidazole derivatives Removed from the list in APRIL 2020. Imidazole is used as a polymerisation crosslinking accelerator and hardener for epoxy resin adhesives. As a highly reactive crosslinking accelerator and hardener, the 2-methylimidazole becomes fully reacted and is not present > 0.1% w/w in the cured epoxy. | Reprotox 1B | 1 N |  | APRIL 2020
Substances that are listed in the ECHA CLH Registry of Intentions that are likely to become CMR 1A or 1B or Endocrine Disruptors in the near future (or proposed hazard class not stated)
|  | 7440-48-4 | Cobalt metal | Alloys (steels, hard coatings, etc.) | Proposed as Carcinogen 1B & Repro tox 1B | 1
|  | 3648-18-8 | Dioctyltin dilaurate | Polyurethane polymerisation catalyst | Proposed as Repro tox 1B | 1
|  | 27554-26-3 | Di-isoctyl phthalate | Plasticiser used in flexible polymers, inks, paints, adhesives, etc. | Proposed as Repro tox 1B | 1
|  | 409-21-2 | Silicon carbide fibres (diameter <3 µm, length > 5 µm and aspect ratio ≥ 3:1) | Filler in ceramics and polymers | Proposed to be Carcinogen 1A | 1
|  | Not specified | Crystalline silica (tridymite, cristobalite) | Proposed by France but without specifying a proposed hazard classification. Classified in some countries (USA) as a carcinogen | Not specified | 1
|  | 13701-59-2 | Barium borate | Added to paints, coatings, adhesives, plastics, and paper products as a UV stabiliser and pigment. Also used for non-linera optics applications such as in lasers. Proposed to be a Repro Tox 1B. | Proposed to be a Repro Tox 1B. | 1
|  | 108-78-1 | Melamine | Fire-retardant used in paints, plastics, and paper, however melamine derivatives are normally used, not pure melamine, although this is possible. Proposed for carcinogeneity | Proposed to be Carcinogen 1B | 3
|  | 5888-33-5 | Exo-1,7,7-trimethylbicyclo[2.2.1]hept-2-yl acrylate | Component of light-cure adhesiv &, inks but may not occur in cured material - proposed classification not published by ECHA, but is unlikely to be a CMR. | Not specified | 3
|  | 101-90-6 | m-bis(2,3-epoxypropoxy)benzene, Resorcinol diglycidyl ether | Epoxy resin and rubber curing agent. Will not occur in cured resin or rubber. Proposed to be: Carcingen 1B | Proposed to b: Carcingen 1B | N
|  | 106-87-6 | 7-oxa-3-oxiranylbicyclo[4.1.0]heptane, Vinylcyclohexene dioxide | Epoxy curing agent, will not occur in cured resin. Proposed to be Carcinogen and Repro tox 1B | Carcinogen and Repro tox 1B proposed | N
|  | 111-77-3 | 2-(2-methoxyethoxy)ethanol, methyl carbitol | Solvent so will not occur in medical devices, proposed as Repro tox 1B | Repro tox 1B proposed | N

## All CMR
EC No. | CAS No. | Name all | Additional details | Hazard classification | Applicable to MDR | Uses
UPDATE January 2024
Candidate list 23 January 2024
211-989-5 | 732-26-3 | 2,4,6-tri-tert-butylphenol | SVHC - 23 January 2024 | Repr. 1B | 1 | Note: previously added as Repr1b by 18thATP in 2023 and classified as "N". Reclassified as "1" in Jan2024. Typically used as an intermediate chemical. However there is some evidence that although it is unlikely, this substance may be used directly as an antioxidant/stabilizer in plastic materials. As such, it is possible this substance may be present > 0.1% in medical devices and patient contact materials.
438-340-0 | 119344-86-4 | 2-(dimethylamino)-2-[(4-methylphenyl)methyl]-1-[4-(morpholin-4-yl)phenyl]butan-1-one | SVHC - 23 January 2024 | Repr. 1B | N | This substance is used as a photo-initiator for UV inks and coatings. In all applications due to the reactive nature of the substance, it is not expected to be present in the supplied article >0.1% w/w as it will be reacted and/or cured.
UPDATE December 2023
Candidate list 14 June 2023
278-355-8 | 75980-60-8 | diphenyl(2,4,6-trimethylbenzoyl)phosphine oxide | SVHC - 14 June 2023 | Repr. 1B | 1 | The main use of this substance is as a photo-initiator where it can be used in UV-curable coatings. The substance is a Norish type I (alpha-cleavage) photo initiator which means that during the reaction the molecule splits into two free radical fragments that become incorporated into the polymer. The substance will be consumed rapidly during the reaction process and any small residual amounts that may remain will be trapped within the polymer matrix. In most applications the substance is expected to be reacted and not present >0.1% w/w in supplied articles. However, the chemical has reported non-industrial uses as an auxiliary for dental technology.
20th ATP - Updated classifications CLP
215-540-4 | 1330-43-4 | disodium tetraborate, anhydrous; | 20th ATP - May 2023 | Repr. 1B | 1 | Flame retardant used in paper, card, paint, wood and also added to cellulose insulation. The substances can also be used in contact lens storage solutions and multi-purpose disinfecting solutions.
UPDATE Jan 2023
Candidate list 17 January 2023
201-236-9 | 79-94-7 | 2,2',6,6'-tetrabromo-4,4'-isopropylidenediphenol | SVHC - 17 January 2023 | Carc . 1B | 1 | As a reactive flame retardant and as an additive flame retardant in the manufacture of polymer resins, in products such as epoxy coated circuit boards, printed circuit boards, paper and textiles.
201-250-5 | 80-09-1 | 4,4'-sulphonyldiphenol | SVHC - 17 January 2023 | Rep. 1B ED Human health ED Environment | 1 | Used as a raw material or intermediate in the production of certain polymers and plastics. BPS is used in curing fast-drying epoxy glues and as a corrosion inhibitor. BPS is a building block in polyethersulfone and some epoxies. Substitution of BPA with BPS is also possible in other applications where BPA is used. It is not commonly used in medical devices.
237-222-4 | 13701-59-2 | Barium diboron tetraoxide | SVHC - 17 January 2023 | Repr. 1B | 1 | It is used as a refractory material and as a component of certain high-temperature ceramics. Barium Borate is UV resistant and may be used as a UV stabiliser in PVC materials in concentrations > 0.1% w/w. There is evidence this substance can also be present >0.1% w/w in resistive ink used in trimmer potentiometers.
|  | Bis(2-ethylhexyl) tetrabromophthalate covering any of the individual isomers and/or combinations thereof | SVHC - 17 January 2023 | ED Human Health | 1 | As a flame retardant and as a plasticiser for flexible polyvinylchloride and for use in wire and cable insulation, film and sheeting, carpet backing, coated fabrics, wall coverings and adhesives.
|  | Perfluoroheptanoic acid and its salts | SVHC - 17 January 2023 | Repr. 1B | N | Used as a raw material or intermediate in the production of certain fluoropolymers and other fluorinated compounds. It is not commonly used in medical devices.The substance is a degradation product of PFAS substances with carbon chains of 7 or more carbon atoms. PFHpA belongs to the chemical group of short-chain PFCAs. PFHpA and its salts can be expected to have very similar properties to PFHxA, HFPO-DA and PFOA. This substance may occur as a breakdown product of stain and grease-resistant coatings on products.
18th ATP - Updated classifications CLP
035-005-00-7 | 12124-97-9 | ammonium bromide | 18 ATP - May 2022 | Repr. 1B | N | Can be used as a raw material or intermediate in the manufacture of certain medical products and devices, but it is not typically used as a direct component of medical devices. It is used as a reagent in chemical synthesis and as a source of bromine in the production of brominated compounds for use in medicines and other medical products. It can also be used as a flame retardant in wood products and there are some studies showing it can be used as a flame retardant in polystyrene.
050-032-00-4 | 2781-10-4 | dibutyltin bis (2-ethylhexanoate) | 18 ATP - May 2022 | Repr. 1B | 1 | An organotin compound that is used as a catalyst in the production of polyvinyl chloride (PVC) and other polymers, and as a lubricant additive. DBT compounds are generally used in colourless and/or transparent plastic articles which could include valves, containers, flexible tubes, blood bags, surgical gloves
050-033-00-X | 1067-33-0 | dibutyltin di(acetate) | 18 ATP - May 2022 | Repr. 1B | 1 | An organotin compound that is used as a catalyst in the production of polyvinyl chloride (PVC) and other polymers, and as a lubricant additive. DBT compounds are generally used in colourless and/or transparent plastic articles which could include valves, containers, flexible tubes, blood bags, surgical gloves
052-001-00-0 | 13494-80-9 | tellurium | 18 ATP - May 2022 | Repr. 1B | 1 | Has antibacterial properties and could be used in antimicrobial coatings for medical devices. Tellurium compounds, such as tellurium dioxide, have the capability to be used in the synthesis of certain pharmaceuticals. It is not commonly used in medical devices.
052-002-00-6 | 7446-07-3 | tellurium dioxide | 18 ATP - May 2022 | Repr. 1B | 1 | Has antibacterial properties and could be used in antimicrobial coatings for medical devices. Tellurium compounds, such as tellurium dioxide, have the capability to be used in the synthesis of certain pharmaceuticals. It is not commonly used in medical devices.
056-005-00-3 | 13701-59-2 | barium diboron tetraoxide | 18 ATP - May 2022 | Repr. 1B | 1 | It is used as a refractory material and as a component of certain high-temperature ceramics. Barium Borate is UV resistant and may be used as a UV stabiliser in PVC materials in concentrations > 0.1% w/w. There is evidence this substance can also be present >0.1% w/w in resistive ink used in trimmer potentiometers. It is not commonly used in medical devices.
603-243-00-6 | 36483-57-5; 1522-92-5 | 2,2-dimethylpropan-1-ol, tribromo derivative; 3-bromo-2,2-bis (bromomethyl)propan- 1-ol | 18 ATP - May 2022 | Carc. 1B | N | Used as a reactive intermediate for high molecular weight flame retardants in combination with phosphorous and also as reactive flame retardants in polyurethanes. It may also used as a reactive processing aid in polymer production. In these applications, the substance is reacted and is not present in the final polymers and polyurethanes.
604-097-00-6 | 732-26-3 | 2,4,6-tri-tert-butylphenol | 18 ATP - May 2022 | Repr. 1B | 1 | Note: uses and "Applicable to MDR" updated in Jan2024. Previously indicated as "N". USes: Typically used as an intermediate chemical. However there is some evidence that although it is unlikely, this substance may be used directly as an antioxidant/stabilizer in plastic materials. As such, it is possible this substance may be present > 0.1% in medical devices and patient contact materials.
604-098-00-1 | 80-09-1 | 4,4'-sulphonyldiphenol; bisphenol S | 18 ATP - May 2022 | Repr. 1B | 1 | Used as a raw material or intermediate in the production of certain polymers and plastics. BPS is used in curing fast-drying epoxy glues and as a corrosion inhibitor. BPS is a building block in polyethersulfone and some epoxies. Substitution of BPA with BPS is also possible in other applications where BPA is used. It is not commonly used in medical devices.
606-153-00-5 | 119-61-9 | benzophenone | 18 ATP - May 2022 | Carc. 1B | 1 | Used as a reagent in chemical synthesis. It could be used as a component of certain coatings or materials that are used in medical devices. Benzophenone is a UV protection agent which can be added to plastics to prevent UV degradation of the plastic. Benzophenone is mainly used in plastic packaging to prevent photo-degradation of the packaging polymers or its contents.
607-761-00-3 | 375-85-9 | Perfluoroheptanoic acid; tridecafluoroheptanoic acid | 18 ATP - May 2022 | Repr. 1B | N | Used as a raw material or intermediate in the production of certain fluoropolymers and other fluorinated compounds. It is not commonly used in medical devices.The substance is a degradation product of PFAS substances with carbon chains of 7 or more carbon atoms. PFHpA belongs to the chemical group of short-chain PFCAs. PFHpA and its salts can be expected to have very similar properties to PFHxA, HFPO-DA and PFOA. This substance may occur as a breakdown product of stain and grease-resistant coatings on products.
607-763-00-4 |  | 6-[C12-18-alkyl- (branched, unsaturated)-2,5-dioxopyrrolidin-1-yl]hexanoic acid, sodium and tris (2-hydroxyethyl) ammonium salts | 18 ATP - May 2022 | Repr. 1B | N | This substance is used in the following products: hydraulic fluids, lubricants and greases and metal working fluids.
607-764-00-X | 2156592-54-8 | 6-[(C10-C13)-alkyl- (branched, unsaturated)-2,5-dioxopyrrolidin-1-yl]hexanoic acid | 18 ATP - May 2022 | Repr. 1B | N | Used as a raw material or intermediate in the production of various chemicals, including fragrances, flavours, and plastics. It is also used as a solvent for certain resins and coatings. It may have potential medicinal properties, as it has been shown to possess antimicrobial and anti-inflammatory properties.
607-765-00-5 |  | 6-[C12-18-alkyl- (branched, unsaturated)-2,5-dioxopyrrolidin-1-yl]hexanoic acid | 18 ATP - May 2022 | Repr. 1B | N | Used as a raw material or intermediate in the production of various chemicals, including fragrances, flavours, and plastics. It is also used as a solvent for certain resins and coatings. It may have potential medicinal properties, as it has been shown to possess antimicrobial and anti-inflammatory properties.
613-342-00-6 | 58-55-9 | theophylline; 1,3-dimethyl-3,7-dihydro-1H-purine-2,6-dione | 18 ATP - May 2022 | Repr. 1B | N | Theophylline is a cyclic phosphodiesterase inhibitor. Used for anti-inflammatory purposes.
616-238-00-9 | 874819-71-3 | N-(2-nitrophenyl) phosphoric triamide | 18 ATP - May 2022 | Repr. 1B | N | 2-Nitrophenyl)-phosphoric triamide, also known as 2-NPPA, is used in the synthesis of a variety of organic molecules. It is a common reagent in organic synthesis and has been used in the synthesis of a variety of drugs and other compounds.
616-240-00-X | 881685-58-1 | Reaction mass of 3-(difluoromethyl)- 1-methyl-N- [(1RS,4SR,9RS)-1,2,3,4-tetrahydro- 9-isopropyl-1,4-methanonaphthalen-5-yl]pyrazole- 4-carboxamide and 3-(difluoromethyl)- 1-methyl-N- [(1RS,4SR,9SR)-1,2,3,4-tetrahydro- 9-isopropyl-1,4-methanonaphthalen-5-yl]pyrazole- 4-carboxamide [≥ 78 % syn isomers ≤ 15 % anti isomers relative content]; isopyrazam | 18 ATP - May 2022 | Repr. 1B | N | A persisent and broad spectrum fungicide for cereals and other crops
023-001-00-8 | 1314-62-1 | divanadium pentaoxide; vanadium pentoxide | 18 ATP - May 2022 | Carc. 1B | N | Used as a catalyst in certain chemical reactions and as a pigment in ceramics and glass. It is not commonly used in medical devices.
601-024-00-X | 98-82-8 | Cumene | 18 ATP - May 2022 | Carc. 1B | N | Used as a raw material or intermediate in the production of certain chemicals, including plastics, resins, and pharmaceuticals. It is not commonly used in medical devices.
603-107-00-6 | 111-77-3 | 2-(2-methoxyethoxy) ethanol; diethylene glycol monomethyl ether | 18 ATP - May 2022 | Repr. 1B | N | Can be used as a solvent in a variety of applications and as an intermediate in the production of other chemicals.
607-230-00-6 |  | 2-ethylhexanoic acid and its salts, with the exception of those specified elsewhere in this Annex | 18 ATP - May 2022 | Repr. 1B | N | Used as a raw material or intermediate in the production of certain chemicals, including plasticizers, surfactants, and lubricants. It is also used as a solvent for certain resins and coatings. Its salts are also used in a variety of applications, including the production of plastics, detergents, and personal care products.
607-734-00-6 | 7216-95-7 | pentapotassium 2,2',2",2"',2""- (ethane-1,2-diylnitrilo) pentaacetate | 18 ATP - May 2022 | Repr. 1B | N | Used as a chelating agent in the following applications: Electroplating Paper & rubber industry Photography Printing inks
607-735-00-1 | 67-43-6 | N-carboxymethyliminobis (ethylenenitrilo)tetra(acetic acid) | 18 ATP - May 2022 | Repr. 1B | N | Used as a chelating agent. It has applications in magnetic resonance imaging (MRI) contrasting agents. It can also be used for treatment of radioactive materials such as plutonium, americium and other actinides. It is an active component of manganese and zinc fertilizers used in agrochemical industry. It acts as a color inhibitor and is used in the production of acrylon. It is also used in soaps as a water softener.
607-736-00-7 | 140-01-2 | pentasodium (carboxylatomethyl) iminobis(ethylenenitrilo) tetraacetate | 18 ATP - May 2022 | Repr. 1B | N | Used as a chelating agent. It is also used in soaps as a water softener and to prevent dyes and perfumes from integrating with metal ions in solution. Can also be used in the pulp industry as stabilizer in peroxide and hydrosulphite bleaching, in the textile industry as stabilizer in bleaching and dyeing baths and in the photographic industry as stabilizer in developing baths.
Candidate list January & June 2022
213-103-2 | 924-42-5 | N-(hydroxymethyl)acrylamide | SVCH - 10 July 2022 | Carc. 1B Muta. 1B |  | Note: Already included in the ATP15
213-934-0 | 1067-53-4 | tris(2-methoxyethoxy)vinylsilane | SVCH - 17 January 2022 | Repr. 1B |  | Note: Already included in the ATP15
204-327-1 | 119-47-1 | 6,6'-di-tert-butyl-2,2'-methylenedi-p-cresol | SVCH - 17 January 2022 | Repr. 1B |  | Note: Already included in the ATP17
253-242-6 | 36861-47-9 | (±)-1,7,7-trimethyl-3-[(4-methylphenyl)methylene]bicyclo[2.2.1]heptan-2-one covering any of the individual isomers and/or combinations thereof (4-MBC) | SVCH - 17 January 2022 | ED Human Health | N | This substance, also known as 4-methylbenzylidene-camphor (4-MBC) or Enzacamene and under trade names Eusolex 6300 and Parsol 5000, is used as a UV absorber/filter (specifically UVB) in cosmetic and personal care applications such as sunscreen and other products claiming an SPF value. Whilst not considered safe for use by the FDA or in Japan and Denmark, 4- MBC is approved by Health Canada and is allowed in the EU to be included in cosmetic products at concentrations up to 4% as a UV filter. A similar compound, 3-BC (which differs structurally by a single methyl group) was included on the Candidate List on 15 January 2019 and is banned for use in cosmetics within the EU. VT 62474 agreed not to include 3-BC on the IEC 62474 database because, similar to 4-MBC, the substance is a chemical UV filter which is used to provide UV protection in sunscreen lotions, sun blocks, hair care products and other cosmetic products claiming a Sun Protection Factor (SPF). None of these applications are relevant to parts and materials used in articles. Therefore, this substance is not present >0.1% w/w in supplied articles.
| 1782069-81-1 | (3E)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one | SVCH - 17 January 2022 | ED Human Health | N
| 95342-41-9 | (1R,3E,4S)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one | SVCH - 17 January 2022 | ED Human Health | N
| 852541-30-1 | (1S,3E,4R)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one | SVCH - 17 January 2022 | ED Human Health | N
| 852541-21-0 | (1R,3Z,4S)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one | SVCH - 17 January 2022 | ED Human Health | N
| 741687-98-9 | (1R,4S)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one | SVCH - 17 January 2022 | ED Human Health | N
| 852541-25-4 | (1S,3Z,4R)-1,7,7-trimethyl-3-(4-methylbenzylidene)bicyclo[2.2.1]heptan-2-one | SVCH - 17 January 2022 | ED Human Health | N
UPDATE Sept 2021
17th ATP - Updated classifications CLP
202-052-1 | 91-23-6 | 2-nitroanisole | 17th ATP - Sept 2021 | Carc. 1B | N | 2-nitroanisole (or o-nitroanisole) is one of three nitronisole isomers, where the para-isomer has previously been assessed as part of the BOMcheck screening of Proposition 65 substances in September 2019. 2-nitroanisole is primarily used as an intermediate in the synthesis of azo dyes, where it is used to produce the precursor o-anisidine. 2-nitroanisole can also be used as an intermediate in the manufacture of pharmaceutical products. In all applications, 2-Nitroanisole is an intermediate and is reacted and so is not present as a substance in the supplied article.
204-126-9 | 116-14-3 | tetrafluoroethylene | 17th ATP - Sept 2021 | Carc. 1B | N | Also known as TFE, this substance is a fluorocarbon - primarily used in industrial preparation of PTFE fluoropolymers such as Teflon. TFE is also used in the preparation of copolymers to produce fluoroplastics such as ETFE and PFA. PTFE and other fluorinated polymers have a broad range of applications that include non-stick coatings on cookware, cable insulation in electronics, bearings and components in small mechanical devices (e.g. Pumps), hose and tubing, waterproof clothing membranes. Medical applications of PTFE and other fluorinated polymers can include fabrics, heart valves and aorta implants, vascular grafts, bone replacements. However, in all these applications, TFE is used a monomer that is reacted during polymerisation and is not present in the supplied article.
204-327-1 | 119-47-1 | 6,6'-di-tert-butyl-2,2'-methylenedi-p-cresol [DBMC] | 17th ATP - Sept 2021 | Repr. 1B | 1 | Assessed as an SVHC as part of September 2021 screening. This substance can be used as an antioxidant additive in the manufacture of polymers, where it is added after the polymerisation process and so may be present in the supplied article in concentrations >0.1% by weight of the article. Polymers containing this substance as an antioxidant may be used in medical device materials and patient contact materials. https://www.solvay.com/en/chemical-categories/polymer-additives-uv-light-stabilizers-antioxidants-and-antistatic-agents/antioxidants-sterile-medical-applications-ppe
215-925-7 | 1453-58-3 | 3-methylpyrazole | 17th ATP - Sept 2021 | Repr. 1B | N | 3-methyl-pyrazole is used as a laboratory reagent and research chemical and also as nitrification inhibitor in nitrogen-based fertilisers.
235-541-3 | 12267-73-1 | tetraboron disodium heptaoxide, hydrate | 17th ATP - Sept 2021 | Repr. 1B | 1 | Flame retardant used in paper, card, paint, wood and also added to cellulose insulation
237-560-2 | 13840-56-7 | orthoboric acid, sodium salt | 17th ATP - Sept 2021 | Repr. 1B | 1 | Flame retardant used in paper, card, paint, wood and also added to cellulose insulation
215-540-4 | 1303-96-4, 1330-43-4, 12179-04-3 | disodium tetraborate decahydrate: borax decahydrate disodium tetraborate pentahydrate | 17th ATP - Sept 2021 | Repr. 1B | 1 | Flame retardant used in paper, card, paint, wood and also added to cellulose insulation. The substances can also be used in contact lens storage solutions and multi-purpose disinfecting solutions.
| 8018-01-7 | mancozeb (ISO); manganese ethylenebis(dithiocarbamate) (polymeric ) complex with zinc salt | 17th ATP - Sept 2021 | Repr. 1B | N | mancozeb (ISO); manganese ethylenebis(dithiocarbamate) (polymeric ) complex with zinc salt is a combination of two dithicarbamates; maneb and zineb and is used as a fungicide to protect agricultural crops against disease. Applications of this fungicide include use in many different plants such as fruits, vegetable, rice and cereals.
234-829-6 | 12035-72-2 | trinickel disulfide: nickel subsulfide | 17th ATP - Sept 2021 | Carc. 1A | N | Process chemical
| 12035-71-1 | heazlewoodite | 17th ATP - Sept 2021 | Carc. 1A | N | Nickel sulphide mineral
204-661-8 | 123-91-1 | 1,4-dioxane | 17th ATP - Sept 2021 | Carc. 1B | N | Added to REACH Candidate List in July 2021. 1,4-dioxane is used primarily as a solvent in manufacture of other chemicals and in paints, laquers, varnishes, cleaning and detergent preparations. It was historically used as a stabiliser with chlorinated solvents, particularly for 1,1,1-trichloroethane, which has since been phased-out. 1,4-dioxane can also be found as a trace impurity if used as a reaction medium during manufacture of other chemicals such as cosmetics and personal care products. For example, 1,4-dioxane can be found as an impurity in ethoxylated chemicals such as Polyethylene oxide aka ‘PEG’. 1,4-dioxane can also be released as an off gassing substance during the production of Polyethylene Terephthalate (PET). No uses of 1,4-dioxane are likely to result in this substance being present at > 0.1% in materials of medical devices or in patient contact materials or gas, liquid or solid transfer materials.
203-437-7 | 106-87-6 | 7-oxa-3-oxiranylbicyclo[4.1.0]heptane; 1,2-epoxy-4-epoxyethylcyclohexane; 4-vinylcyclohexene diepoxide | 17th ATP - Sept 2021 | Repr. 1B | N | 4-Vinylcyclohexene diepoxide is used in the production of epoxy resins in coatings and adhesives and as an intermediate chemical and reactive diluent. Reported applications of epoxy resins using this substance include the production of tissue embedding agents for use in electron microscopy. Epoxide groups are highly reactive species that polymerise in the presence of catalysts or when heated. This substance is expected to be reacted in all uses.
234-232-0 | 10605-21-7 | carbendazim (ISO): methyl benzimidazol-2-ylcarbamate | 17th ATP - Sept 2021 | Muta. 1B, Repr. 1B | N | Agricultural fungicide
404-200-2 | 110488-70-5 | dimethomorph (ISO); (E,Z)-4-(3-(4-chlorophenyl)-3-(3,4-dimethoxyphenyl)acryloyl)morpholine | 17th ATP - Sept 2021 | Repr. 1B | N | Dimethomorph is used as a fungicide to protect agricultural crops against disease and rot. Dimethomorph is a mixture of E and Z-isomers (1:1 ratio) where the Z-isomer is the active fungicidal isomer. It is used for many different crops including onions, potatoes, tomatoes and fruit.
206-022-9 | 288-88-0 | 1,2,4-triazole | 17th ATP - Sept 2021 | Repr. 1B | N | 1,2,4-triazole is one of a pair of isomeric molecules belonging to the triazoles group. 1,2,4-Triazole and its derivatives have a wide range of applications including pharmaceutical (antifungal, antibacterial, antituberical, anticancer, antioxodant, anti-inflammatory, antiviral and anticonvulsant) and agricultural (intermediate in synthesis of herbicides). 1,2,4-triazole can also be biocides and stabilisers in metal cutting oils. The cutting oils are washed from the metal part after the cutting operation and are not present on the finished metal article. Triazoles have reported industrial uses in photography to prevent fog formation in photographic emulsions, as LEDs in Electroluminescent devices and in increasing efficiency of cooling fluids (lubricant oil). None of these applications (which involve derivatives of 1,2,4-triazole) will result in 1,2,4-triazole being used in medical device materials or occuring in patient contact materials or gas, liquid or solid transfer materials.
Updated REACH candidate list 08/07/2021
|  | Phenol, alkylation products (mainly in para position) with C12-rich branched alkyl chains from oligomerisation, covering any individual isomers and/ or combinations thereof (PDDP) | SVHC - July 2021 | Repr. 1B ED Health ED Envi | N | This group of substances (PDDP) was added to the REACH Candidate List in July 2021 and includes PDB (Dodecylphenol, CAS 121158-58-5). https://www.products.pcc.eu/en/id/244261/dodecylphenol-2/ The main application of PDDB is as a chemical intermediate in the manufacture of lubricant additives and fuel system cleaners for use in petrol- and diesel-powered road vehicles and marine diesel engines. The lubricant oils can contain up to 2% of unreacted PDDP impurity. PDDB is also used as a chemical intermediate in the manufacture of chemicals, plastic products and rubber products such as tires. The boiling point range for PDDB substances is between 189-270°C and so PDDB will be completely reacted in these high temperature manufacturing processes. PDDB is used in certain resins and hardeners where it can be present up to 50%. However, in these applications the PDDB is reacted in the curing process and is not present > 0.1% in supplied articles. PDDP is also used as an intermediate in production of specialized paints, varnishes, resins and coatings (e.g. specialist hard floor coating https://www.nzspecialisedcoatings.co.nz/wp-content/uploads/MSDS/Sikafloor-264-Part-B-MSDS.pdf) and as monomer in phenol/formaldehyde resins that are then used in printing inks and paints and coatings. These specialist paints and coatings can typically contain about at 2.5-5% PDDP. However, the boiling point range for PDDP substances is between 189-270°C and so any unreacted amounts of these substances will evaporate in the drying/curing process. An example of this specialist use is hard flooring. None of the above applications will result in PDDP being present >0.1% in supplied articles
| 210555-94-5 | Phenol, 4-dodecyl, branched | SVHC - July 2021 | Repr. 1B ED Health ED Envi
| 27459-10-5 | 4-isododecylphenol | SVHC - July 2021 | Repr. 1B ED Health ED Envi
| 27147-75-7 | Phenol, 4-isododecyl- | SVHC - July 2021 | Repr. 1B ED Health ED Envi
310-154-3 | 121158-58-5 | Phenol, dodecyl-, branched | SVHC - July 2021 | Repr. 1B ED Health ED Envi
| 74499-35-7 | Phenol, (tetrapropenyl) derivatives | SVHC - July 2021 | Repr. 1B ED Health ED Envi
| 57427-55-1 | Phenol, tetrapropylene- | SVHC - July 2021 | Repr. 1B ED Health ED Envi
201-025-1 | 77-40-7 | 4,4'-(1-methylpropylidene)bisphenol | SVHC - July 2021 | ED Health ED Envi | 1 | BPB is structurally similar to BPA and so may be used to make polymers such as polycarbonate and epoxy resins, but usually present at <0.1%. Similarly, may also be used as alternative to BPA as an uncommon additive in PVC at about 0.1%
UPDATE Dec 2020
Substances identified as endocrine disruptors at EU level: List I (https://edlists.org/the-ed-lists/list-i-substances-identified-as-endocrine-disruptors-by-the-eu)
233-539-7 | 10222-01-2 | 2,2-dibromo-2-cyanoacetamide | Regulatory field: BPR | ED | N | 2,2-Dibromo-2-cyanoacetamide, also known as 2,2-dibromo-3-nitrilopropionamide (DBNPA) is a fast acting biocide/antimicrobial agent that has a number of industrial applications. Reported applications as a biocide include use in the paper industry, in industrial food processing vessels (including fermenters), water cooling systems and as a preservative in animal glues, metalworking cutting fluids, oil drilling fluids, emulsions, paints, and paper products. DBNPA is preferred for use as it is fast-acting but quickly degrades in the presence of water under certain conditions. These properties mean that this substance is used for short term-preservation applications and will not be used in medical device materials or occur in patient contact materials or gas, liquid or solid transfer materials.
200-673-2 | 67-97-0 | Cholecalciferol | Regulatory field: BPR | ED | N | Cholecalciferol, also known as Vitamin-D3, is made in the body naturally via UVB exposure and can also be found in food supplements. However, it is also used in rodenticide as it is toxic to animals (and humans in high does). Ingestion of toxic amounts of D3 causes overly high calcium and phosphorus levels in the body, resulting in severe, acute kidney failure, cardiovascular abnormalities, and tissue mineralization.
UPDATE SEPT 2020
15th ATP - Updated classifications CLP
206-991-8 | 409-21-2 | slicon carbide fibres (with diameter < 3 μm, length > 5 μm and aspect ratio ≥ 3:1) | 15th ATP - May 2020 | Carc. 1B | N | Silicon carbide fibres are a high strength and oxidation resistant ceramic material typically used in high temperature and reinforcement applications in aerospace, military and civil industrial fields. Note that this listing in the 15th ATP specifically refers to fibres with diameter < 3 μm, length > 5 μm and aspect ratio ≥ 3:1. This fibre form of silicon carbide is not used in medical devices.
213-934-0 | 1067-53-4 | tris(2-methoxyethoxy)vinylsilane; 6-(2-methoxyethoxy)-6-vinyl-2,5,7,10-tetraoxa-6-silaundecane | 15th ATP - May 2020 | Repr. 1B | 2 | tris(2-methoxyethoxy)vinylsilane, an organosilane, is used as a coupling agent and adhesion promoter in filled polymer systems to improve mechanical and electrical properties, particularly after wet conditioning. It can also be used as an intermediate and laboratory chemical. tris(2-methoxyethoxy)vinylsilane can be used as a co-monomer for the preparation of different polymers such as polyethylene or acrylics. A major application of polymers containing this substance is in wires, cables and pipes. However, tris(2-methoxyethoxy)vinylsilane will be reacted in all of the above applications and so is not likely to be found > 0.1% w/w in the supplied article. Coating used for outdoor decking - 0.3% w/w https://www.tor-coatings.com/wp-content/uploads/2020/01/RC5490001-Elastadeck-Walkway-Gritcoat-Base-gb.pdf
222-883-3 293-901-5 | 3648-18-8 91648-39-4 | dioctyltin dilaurate; [1] stannane, dioctyl-, bis(coco acyloxy) derivs. [2] | 15th ATP - May 2020 | Repr. 1B | 1 | Dioctyltin (DOT) compounds are added to PVC plastic (at a typical addition rate of between 0.001 and 1% w/w) to make it more stable to heating. DOT compounds prevent the plastic from degrading when it is heated to form different products (for example when PVC is injection moulded it may be heated to between 170°C and 210°C). DOT also protects the PVC plastic against degradation from sunlight. DOT compounds are generally used in colourless and/or transparent plastic articles. DOT compounds are used in rigid extrusions (e.g. trims, conduits), rigid mouldings (e.g. bottles) and medical applications (e.g. valves, containers, flexible tubes, blood bags, surgical gloves).
205-886-4 | 191-30-0 | dibenzo[def,p]chrysene dibenzo[a,l]pyrene | 15th ATP - May 2020 | Carc. 1B | 2 | dibenzo[def,p]chrysene,dibenzo[a,l]pyrene is a Polycyclic Aromatic Hydrocarbon (PAH) which can be present as an impurity in plastic and rubber parts, because they are present as impurities in some of the raw materials used in the production of these parts, in particular in extender oils and in carbon black. PAHs are not added intentionally to the articles and do not perform any specific function as constituents of the plastic or rubber parts. Extender oils can be used as plasticizers in rubber and plastics or in potting material used in electronic ballasts. PAHs may be present as contaminant in these oils. Black pigments and dyes used in rubber and plastic (e.g. Carbon Black) may contain PAH contaminated soot. In all of these applications the concentration of PAH in the supplied articles is typically below 0.02% and will always be considerably less than 0.1%.
| 125225-28-7 115850-69-6 115937-89-8 | ipconazole (ISO); (1RS,2SR,5RS;1RS,2SR,5SR)-2-(4-chlorobenzyl)-5-isopropyl-1-(1H-1,2,4-triazol-1-ylmethyl)cyclopentanol | 15th ATP - May 2020 | Repr. 1B | N | Agrochemical used as a fungicide to control seed diseases in root and leafy vegetables, cereals and rice.
205-594-7 | 143-24-8 | bis(2-(2-methoxyethoxy)ethyl)ether; tetraglyme | 15th ATP - May 2020 | Repr. 1B | N | Solvent which can be used in the electrolyte solution inside lithium-ion batteries
221-967-7 | 3296-90-0 | 2,2-bis(bromomethyl)propane-1,3-diol | 15th ATP - May 2020 | Carc. 1B Muta. 1B | 2 | 2,2-bis(bromomethyl)propane-1,3-diol, also known as BMP is a reactive brominated flame retardant that is used in the manufacture of unsaturated polyester resins for moulded products and rigid polyurethane foam with a typical addition rate of 1-5%. (https://www.industrialchemicals.gov.au/sites/default/files/1%2C3-Propanediol%2C%202%2C2-bis%28bromomethyl%29-_Human%20health%20tier%20II%20assessment.pdf and https://www.ncbi.nlm.nih.gov/books/NBK390857/). It can also be used as an intermediate in the production of other flame retardants. Reactive flame retardants are added during the polymerisation process and become an integral part of the polymer. The result is a modified polymer with flame retardant properties and different molecular structure compared to the original polymer molecule. As a result BMP is not present as a substance in the manufactured polymer > 0.1% w/w. BMP is also used as a reactive flame retardant in waterproofing coatings such as GACOFLEX® U-64 LIQUID POLYURETHANE COATING which is a two-part liquid polyurethane coating which contains 1-5% BMP (https://gaco.com/wp-content/uploads/U6402P-PEWTER-URE-CTG-POLY-SIDE-GHS04212015.pdf) and FIRESTONE® REGULAR URETHANE BLACK BASECOAT which also contains 1-5% BMP (https://41z32e1503dj3gyqe3151089-wpengine.netdna-ssl.com/wp-content/uploads/UB6421PC-%E2%80%93-REGULAR-URETHANE-BLACK-BASECOAT-VOC-COMPLIANT-POLYOL-COMPONENT-A-GHS07162018.pdf). As mentioned above, reactive flame retardants become an integral part of the polymer and therefore are not present in the coated material > 0.1% w/w.
201-289-8 | 80-54-6 | 2-(4-tert-butylbenzyl)propionaldehyde | 15th ATP - May 2020 | Repr. 1B | N | Also known via trade name lilial, this substance is a synthetic Lily of the Valley (flower) fragrance that is used as a fragrance ingredient in consumer products such as air care products, perfumes, polishes, waxes, cleaning and personal care products. It is also used as a perfume in cosmetic preparations and laundry powders, often under the name butylphenyl methylpropional. It can also be used in industrial coating and cleaning products. <1% w/w in car air freshener - http://s7d9.scene7.com/is/content/GenuinePartsCompany/1334652pdf?$PDF$
248-523-5 | 27554-26-3 | diisooctyl phthalate | 15th ATP - May 2020 | Repr. 1B | 1 | Plasticizer in plastic materials such as PVC and other plastic polymers. In flexible PVC the typical phthalate content can range from 20 to 50% w/w.
221-499-3 | 3121-61-7 | 2-methoxyethyl acrylate | 15th ATP - May 2020 | Repr. 1B | 3 | 2-methoxyethyl acrylate is a monomer used in the manufacture of specific polymers for use in coatings, paints, adhesives and plastics. The addition of this substance into the polymers provides improves flexibility at lower temperatures, improves UV and water resistance and also improves adhesion. In these applications the monomer is reacted during the manufacturing process and is not found in concentrations > 0.1% in the polymer. It can also be used as an additive to improve oil resistance. Neat Poly(2-methoxyethyl acrylate) is an antithrombogenic liquid used in medical fields but will not be contained in supplied articles.
274-357-8 | 70161-44-3 | sodium N-(hydroxymethyl)glycinate; [formaldehyde released from sodium N-(hydroxymethyl)glycinate] | 15th ATP - May 2020 | Carc. 1B | N | The substance is an anti-microbial which releases formaldehyde and is used as a preservative to prevent microbial growth and extend shelf life of products. The substance is used in consumer cosmetic products and cleaning agents and has industrial uses in paints and lacquers, adhesives, metalworking fluids and textiles.sodium hydroxymethyl glycinate is an anti-microbial. Cleaning wipes (0.1-1% w/w) - https://infoaboutthisproduct.eu/files/products/msds/shell-exterior-wipes-7920-20160302.pdf
236-671-3 | 13463-41-7 | pyrithione zinc; (T-4)-bis[1-(hydroxy-.kappa.O)pyridine-2(1H)-thionato-.kappa.S]zinc | 15th ATP - May 2020 | Repr. 1B | 1 | This substance is primarily used for its antibacterial properties as a biocide in cosmetic and personal care products. It can also be used to produce antimicrobial textiles and as a biocide in polyurethane foam, PVC and film preservatives. Zinc pyrithione can be added to PVC in a DIDP paste which contains 5% zinc pyrithione. The DIDP paste is added to flexible and plastisol PVC applications at 1% to 4% by weight of the total PVC formulation to provide protection against attack from a wide spectrum of fungi and bacteria. Therefore, the zinc pyrithione could be present up to 0.2% in the supplied PVC article - https://www.valtris.com/intercide-zinc-pyrithione-paste/ Powdered ZNP can also be added to Styrene-ethylene/butylene-styrene (SEBS) based thermoplastic elastomers (TPE) at 1.5% w/w, which are used in a wide range of products with elastic properties, such as remote control keyboards, cell phone covers, tooth brush cables, toys and others, to provide antimicrobial properties - https://www.scielo.br/pdf/mr/v20n5/1516-1439-mr-1980-5373-MR-2017-0137.pdf
262-661-3 | 61213-25-0 | flurochloridone (ISO); 3-chloro-4-(chloromethyl)-1-[3-(trifluoromethyl)phenyl]pyrrolidin-2-one | 15th ATP - May 2020 | Repr. 1B | N | This is an agrochemical which is used as a herbicide and pesticide to destroy pests. For example, it is used as a pesticide to treat sunflower crops. It is also used as a herbicide to control a range of weeds in umbelliferous crops (e.g. parsnips), cereals and potatoes
213-103-2 | 924-42-5 | N-(hydroxymethyl)acrylamide; methylolacrylamide; [NMA] | 15th ATP - May 2020 | Carc. 1B Muta. 1B | 2 | N-methylol acrylamide is a crosslinkable monomer widely used as in the manufacture of polymer products (adhesives, binders, surface coatings, resins and thermoplastics). N-methylolacrylamide is a bi-functional monomer with reactive vinyl and hydroxyethyl groups. Thermoplastic polymers can be formed by copolymerization of N-methylol-acrylamide with a variety of vinyl monomers by emulsion, solution and suspension techniques. The resulting products, which have pendant hydroxyethyl groups, can undergo cross-linking under moderate conditions, permitting conversion of thermoplastic backbone polymers to thermoset materials at the point of use in the absence of an external cross-linking agent. Conversely, the hydroxyethyl group can be reacted with a substrate like cellulose and subsequently cross-linked by free-radical polymerization. N-methylol acrylamide is also used as an intermediate in the production of other substances. N-methylol acrylamide is reacted in all of these uses and will not be present in supplied articles > 0.1% w/w
15th ATP Modification to already present substances
222-583-2 | 3542-36-7 | dichlorodioctylstannane | 15th ATP - May 2020 | Repr. 1B | 1 | Dioctyltin (DOT) compounds are added to PVC plastic (at a typical addition rate of between 0.001 and 1% w/w) to make it more stable to heating. DOT compounds prevent the plastic from degrading when it is heated to form different products (for example when PVC is injection moulded it may be heated to between 170°C and 210°C). DOT also protects the PVC plastic against degradation from sunlight. DOT compounds are generally used in colourless and/or transparent plastic articles. DOT compounds are used in rigid extrusions (e.g. trims, conduits), rigid mouldings (e.g. bottles) and medical applications (e.g. valves, containers, flexible tubes, blood bags, surgical gloves).
239-622-4 | 15571-58-1 | 2-ethylhexyl 10-ethyl-4,4-dioctyl-7-oxo-8-oxa-3,5-dithia-4-stannatetradecanoate; [DOTE] | 15th ATP - May 2020 | Repr. 1B | 1 | PVC stabiliser. Mainly used in rigid PVC including transparent PVC, but may also occur in flexible PVC
231-100-4 | 7439-92-1 | lead powder; [particle diameter < 1 mm] | 15th ATP - May 2020 | Repr. 1A | 2 | lead powder; [particle diameter < 1 mm] may be used in medical devices but not in patient contact materials or gas, liquid or solid transfer materials
231-100-4 | 7439-92-1 | lead massive: [particle diameter ≥ 1 mm] |  | Repr. 1A | 2 | lead massive: [particle diameter ≥ 1 mm] may be used in medical devices but not in patient contact materials or gas, liquid or solid transfer materials
202-987-5 | 101-90-6 | m-bis(2,3-epoxypropoxy)benzene; resorcinol diglycidyl ether | 15th ATP - May 2020 | Carc. 1B | 2 | resorcinol diglycidyl ether is a reactive diluent and modifier in epoxy resins (used in coatings, potting and tooling compounds, adhesives and composites) and is also used as a curing agent in polysulfide rubber. resorcinol diglycidyl ether is reacted or cured in all of these applications and so will not be present >0.1% w/w in a supplied article.
202-496-6 | 96-29-7 | butanone oxime; ethyl methyl ketoxime; ethyl methyl ketone oxime | 15th ATP - May 2020 | Carc. 1B | 2 | butanone oxime, known as MEKO in the paint industry, is widely used as an anti-skinning agent to prevent oxidative drying in the container of alkyd (oil-based) paints, primers, varnishes and other coatings that can be used in consumer or industrial applications. Concentrations of butanone oxime in paints is up to 1% w/w and 0.1-0.2% w/w in colorants or pigments used in printing inks. Other reported uses of butanone oxime include as an ingredient in pesticide products and wood preservatives, as a chemical intermediate and as a corrosion inhibitor in industrial boilers and water treatement systems. None of the above applications will result in the substance being present >0.1% w/w in a supplied article. When used in paints, the substance will evapourate during the drying process. butanone oxime can also be used in silicone sealants particularly room temperture vulcanisation (RTV) silicone. Example addition rates of butnanone oxime in silicone sealants include 0.5-3% w/w. However, in this application butanone oxime evaporates during the curng process and so will not be present in a supplied article >0.1% w/w. Canada Gov Study - https://www.ec.gc.ca/ese-ees/default.asp?lang=En&n=32AD1FD8-1 Primer 0.3% w/w - https://www.crowntrade.co.uk/wp-content/uploads/2014/12/Crown-Trade-Red-Oxide-Primer-CLP-SDS-180417.pdf Gloss paint 0.3-1% w/w - https://www.hammerite.co.uk/files/2017/02/HM_GB_EN_RADIATOR_PAINT_GLOSS.pdf Coating <1% w/w - https://marinecoatings.brand.akzonobel.com/m/6e00da8714a4f5f4/original/CLT229_0010eu_UA_EN_20170602_1.pdf Silicate sealant 0.5-1% w/w - https://www.techsil.co.uk/media/pdf/SDS/TBSI19058-sds.pdf Silicone sealant 1-3 % w/w - https://docs.rs-online.com/d07b/0900766b8134b277.pdf Study of emission of butanone oxime from sealants - https://www.sciencedirect.com/science/article/pii/S0026265X2031136X
201-279-3 | 80-43-3 | bis(α,α-dimethylbenzyl) peroxide |  | Repr. 1B | 2 | bis(α,α-dimethylbenzyl) peroxide also known as Dicumyl Peroxide is an organic peroxide used as a cross-linking agent in the maufacture of polymers and elastomers. Typical applications of Dicumyl Peroxide as cross linking agent include the vulcinaization of natural and synthetic rubbers and in the production of polyethylene and polymers used to produce hose, wires, rubber seals etc. In all of these applications Dicumyl Peroxide will be reacted during the curing process. Dicumyl Preoxide is also used as a flame retardant synergist in expanded polystyrene (EPS) where it is added in small quantities alongside halogonated flame retartardants to act as catalyst for the flame retardant action in case of heating. As it is an additive it may be present in supplied articles. However, EPS will not be used in medical devices in patient contact materials or gas, liquid or solid transfer materials.
Updated REACH candidate list 25/06/2020
245-152-0 | 22673-19-4 | Dibutylbis(pentane-2,4-dionato-O,O')tin | SVHC - 25/06/2020 | Repr. 1B | 1 | 14th ATP Dibutylbis(pentane-2,4-dionato-O,O')tin has similar properties and uses to dibutyltin dichloride which is already listed in the Candidate list as likely to be present in supplied articles at concentrations > 0.1% w/w. Butylin compounds are used as biocides and as stabilisers in plastics. Dibutylbis(pentane-2,4-dionato-O,O')tin has widespread uses as a catalyst and in the manufacturing of adhesives, sealants, coatings, dyes, polymer preparations, resins and rubber. It can be found in supplied articles in the following applications and concentration ranges: • Adhesive/sealant: 0.1-10% • Hardener/coating: 5-10%
202-318-7 | 94-26-8 | Butyl 4-hydroxybenzoate | SVHC - 25/06/2020 | ED | N | Butyl 4-hydroxybenzoate, also known as Butylparaben, or butyl p-hydroxybenzoate, is a member of the paraben family, a group of naturally occurring compounds found in fruit and vegetable products and produced by some microorganisms. Parabens are widely used as preservatives in cosmetics in concentrations between 0.01 and 0.3 %. Since the 1940’s Butylparaben has been widely used an antimicrobial additive for cosmetics and personal care products and in pharmaceuticals for its bactericidal, fungicidal properties. Butylparaben can be found at low concentrations in solid medication suspensions such as ibuprofen. It can also be used as a food preservative. Butylparaben can also be used as a heat/pressure transfer fluid in closed systems. None of these applications result in Butyl 4-hydroxybenzoate being present > 0.1% w/w in supplied articles.
211-765-7 | 693-98-1 | 2-methylimidazole | SVHC - 25/06/2020 | Repr. 1B | N | 14th ATP 2-methylimidazole can be used as a polymerisation crosslinking accelerator and hardener for epoxy resin systems for semiconductor potting compounds and soldering masks. 2-methylimidazole is fully reacted in these applications. 2-methylimidazole can also be used as an intermediate in the manufacture of a wide range of polymers including epoxy resin pastes, acrylic rubber-fluororubber laminates, films, adhesives, textile finishes and epoxy silane coatings. It has also been used as an auxiliary in the dyeing of acrylic fibers and plastic foams and as in intermediate in the manufacture of pharmaceuticals, photographic and photothermographic chemicals, dyes and pigments, agricultural chemicals and rubber. None of these applications result in 2-methylimidazole being present > 0.1% w/w in supplied articles.
214-012-0 | 1072-63-5 | 1-vinylimidazole | SVHC - 25/06/2020 | Repr. 1B | N | 13th ATP 1-vinylimidazole is used as UV curing agent in UV cured inks, UV cured coatings and UV cured adhesives. It is a highly reactive UV curing agent and becomes fully reacted during the UV curing process via copolymerization. 1-vinylimidazole is also used as an epoxy accelerator wherein the substance also becomes fully reacted. 1-vinylimidazole is also used as an ingredient in the manufacture of molecularly imprinted polymers, in the manufacture of magnetic-poly(divinylbenzene-1-vinylimidazole) microbeads, and in the manufacture of ion imprinted polyvinylimidazole-silica hybrid copolymer. None of these applications result in 1-vinylimidazole being present > 0.1% w/w in manufactured articles.
UPDATE APRIL 2020
Updated classifications CLP (Evaluated by BOMcheck Substance List Working Group on 11 March 2020)
231-158-0 | 7440-48-4 | cobalt | 14th ATP - February 2020 | Carc 1B, Repr 1B | 1 | Alloys (steels, hard coatings, etc.)
237-396-1 | 13770-89-3 | nickel bis(sulfamidate); nickel sulfamate | 14th ATP - February 2020 | Carc 1A; Repr 1B | N | Process chemical. The substance can also be used in contact lens storage solutions and multi-purpose disinfecting solutions.
200-849-9 | 75-21-8 | ethylene oxide; oxirane | 14th ATP - February 2020 | Carc 1B; Repr 1B; Muta 1B | N | Process chemical
404-360-3 | 119313-12-1 | 2-benzyl-2-dimethylamino-4'-morpholinobutyrophenone | 14th ATP - February 2020 | Repr 1B | N | The substance is classed as a type I photoinitiator and undergoes homolytic cleavage upon absorption of UV light to produce free radicals. The free radicals that are generated then initiate the polymerisation process. Further details about homolytic cleavage of type I photoinitiators is provided at the below webpage https://polymerinnovationblog.com/uv-curing-part-three-free-radical-photoinitiators/ As a type I photoinitiator, the substance is used in very fast curing applications including high-speed printing inks such as flexo, offset litho and UV ink jet. It is used to assist the UV curing process for inks, toners, adhesives, resins, paints and coatings, solder masks, etch resists and printing plates. The substance has a yellowing effect over time and so is not likely to be used in clear coatings. The typical addition rate for this substance in the coatings, adhesives and inks is 1-10 % before the UV curing process is carried out. Industry sources indicate that the efficiency of homolytic cleavage during the UV curing process is typically around 99% (incomplete homolytic cleavage would affect the chemical properties of the ink or surface coating and likely result in quality issues with the coated article). Therefore, this substance will not be present in concentrations > 0.1% by weight in materials used in medical devices.
204-064-2 | 115-09-3 | methylmercuric chloride | 14th ATP - February 2020 | Repr 1A | N | Previously, Methylmercuric salts were produced during the manufacture of acetaldehyde. However, this synthesis has since been replaced by alternative manufacturing methods. Methlymecuric salts were previously used as fungicides and pesticides. Mercury containing pesticides were banned by the US EPA in 1976 and are no longer used. There are no applications where methylmercuric chloride will be present >0.1% in supplied articles.
205-877-5 | 189-55-9 | benzo[rst]pentaphene | 14th ATP - February 2020 | Carc 1B | 2 | benzo[rst]pentaphene is a Polycyclic Aromatic Hydrocarbon (PAH) which can be present as an impurity in plastic and rubber parts, because they are present as impurities in some of the raw materials used in the production of these parts, in particular in extender oils and in carbon black. PAHs are not added intentionally to the articles and do not perform any specific function as constituents of the plastic or rubber parts. Extender oils can be used as plasticizers in rubber and plastics or in potting material used in electronic ballasts. PAHs may be present as contaminant in these oils. Black pigments and dyes used in rubber and plastic (e.g. Carbon Black) may contain PAH contaminated soot. In all of these applications the concentration of PAH in the supplied articles is typically below 0.02% and will always be considerably less than 0.1%.
205-878-0 | 189-64-0 | dibenzo[b,def]chrysene; dibenzo[a,h]pyrene | 14th ATP - February 2020 | Carc 1B | 2 | dibenzo[b,def]chrysene is a Polycyclic Aromatic Hydrocarbon (PAH) which can be present as an impurity in plastic and rubber parts, because they are present as impurities in some of the raw materials used in the production of these parts, in particular in extender oils and in carbon black. PAHs are not added intentionally to the articles and do not perform any specific function as constituents of the plastic or rubber parts. Extender oils can be used as plasticizers in rubber and plastics or in potting material used in electronic ballasts. PAHs may be present as contaminant in these oils. Black pigments and dyes used in rubber and plastic (e.g. Carbon Black) may contain PAH contaminated soot. In all of these applications the concentration of PAH in the supplied articles is typically below 0.02% and will always be considerably less than 0.1%.
308-208-6 | 97925-95-6 | ethanol, 2,2’-iminobis-, N- (C13-15-branched and linear alkyl) derivs. | 14th ATP - February 2020 | Repr 1B | 1 | Ethanol, 2,2'-iminobis-, N-(C13-15-branched and linear alkyl) derivatives are added during the manufacture of plastics as an antistatic agent to reduce or eliminate a static build up on the surface of the plastic. Anti-static additives work by lowering the resistivity of the plastic so that charges are mobile and therefore will not cause static issues such as dust attraction or electrostatic discharge. Typical addition rates for the substance are 1-4% in ABS plastic and 0.1-0.4% w/w in polyethylene.
276-090-2 | 71850-09-4 | diisohexyl phthalate | 14th ATP - February 2020 | Repr 1B | 1 | Plasticizer in plastic materials such as PVC and other plastic polymers. In flexible PVC the typical phthalate content can range from 20 to 50% w/w.
| 100784-20-1 | halosulfuron-methyl (ISO); methyl 3-chloro-5-{[(4,6- dimethoxypyrimidin-2-yl) carbamoyl]sulfamoyl}-1- methyl-1H-pyrazole-4- carboxylate | 14th ATP - February 2020 | Repr 1B | N | halosulfuron-methyl is used as a pesticide and herbicide. It is approval for use as a pesticide in the EU expires in 2023. https://www.montereylawngarden.com/wp-content/uploads/2018/09/SedgeHammer_Plus_2018_SDS.pdf There are no applications where halosulfuron-methyl will be present >0.1% in supplied articles.
211-765-7 | 693-98-1 | 2-methylimidazole | 14th ATP - February 2020 | Repr 1B | N | 2-methylimidazole can be used as a polymerisation crosslinking accelerator and hardener for epoxy resin systems for semiconductor potting compounds and soldering masks. 2-methylimidazole is fully reacted in these applications. 2-methylimidazole can also be used as an intermediate in the manufacture of a wide range of polymers including epoxy resin pastes, acrylic rubber-fluororubber laminates, films, adhesives, textile finishes and epoxy silane coatings. It has also been used as an auxiliary in the dyeing of acrylic fibers and plastic foams and as in intermediate in the manufacture of pharmaceuticals, photographic and photothermographic chemicals, dyes and pigments, agricultural chemicals and rubber. None of these applications result in 2-methylimidazole being present > 0.1% w/w in supplied articles.
245-152-0 | 22673-19-4 | dibutylbis(pentane-2,4- dionato-O,O’)tin | 14th ATP - February 2020 | Repr 1B | 1 | Dibutylbis(pentane-2,4-dionato-O,O')tin has similar properties and uses to dibutyltin dichloride which is already listed in the Candidate list as likely to be present in supplied articles at concentrations > 0.1% w/w. Butylin compounds are used as biocides and as stabilisers in plastics. Dibutylbis(pentane-2,4-dionato-O,O')tin has widespread uses as a catalyst and in the manufacturing of adhesives, sealants, coatings, dyes, polymer preparations, resins and rubber. It can be found in supplied articles in the following applications and concentration ranges: • Adhesive/sealant: 0.1-10% • Hardener/coating: 5-10%
UPDATE JANUARY 2020
Updated REACH Candidate list
276-090-2 | 71850-09-4 | Diisohexyl phthalate | Not registered under REACH | Repr 1B | 1 | Plasticizer in plastic materials such as PVC and other plastic polymers. In flexible PVC the typical phthalate content can range from 20 to 50% w/w.
404-360-3 | 119313-12-1 | 2-benzyl-2-dimethylamino-4'-morpholinobutyrophenone | The substance is used in polymer production | Repr 1B | N | The substance is classed as a type I photoinitiator and undergoes homolytic cleavage upon absorption of UV light to produce free radicals. The free radicals that are generated then initiate the polymerisation process. Further details about homolytic cleavage of type I photoinitiators is provided at the below webpage https://polymerinnovationblog.com/uv-curing-part-three-free-radical-photoinitiators/ As a type I photoinitiator, the substance is used in very fast curing applications including high-speed printing inks such as flexo, offset litho and UV ink jet. It is used to assist the UV curing process for inks, toners, adhesives, resins, paints and coatings, solder masks, etch resists and printing plates. The substance has a yellowing effect over time and so is not likely to be used in clear coatings. The typical addition rate for this substance in the coatings, adhesives and inks is 1-10 % before the UV curing process is carried out. Industry sources indicate that the efficiency of homolytic cleavage during the UV curing process is typically around 99% (incomplete homolytic cleavage would affect the chemical properties of the ink or surface coating and likely result in quality issues with the coated article). Therefore, this substance will not be present in concentrations > 0.1% by weight in materials used in medical devices.
400-600-6 | 71868-10-5 | 2-methyl-1-(4-methylthiophenyl)-2-morpholinopropan-1-one | The substance is used in polymer production | Repr 1B | 1 | Photoinitiator resulting in inclusion into a matrix, including application in coatings, adhesives and inks.
UPDATE SEPTEMBER 2019
Updates to the Registry of SVHC intentions
203-585-2 | 108-46-3 | Resorcinol |  |  |  | Endocrine disrupting properties (Article 57(f) – human health)
404-360-3 | 119313-12-1 | 2-benzyl-2-dimethylamino-4'-morpholinobutyrophenone |  |  |  | Toxic for reproduction (Article 57c)
400-600-6 | 71868-10-5 | 2-methyl-1-(4-methylthiophenyl)-2-morpholinopropan-1-one |  |  |  | Toxic for reproduction (Article 57c)
276-090-2 | 71850-09-4 | Diisohexyl phthalate |  |  |  | Toxic for reproduction (Article 57c)
tetraethylazanium nonafluorobutane-1-sulfonate EC / List no: 700-536-1 ¦ CAS no: 25628-08-4
1,1,2,2,3,3,4,4,4-nonafluoro-N-(2-hydroxyethyl)-N-methylbutane-1-sulphonamide EC / List no: 252-043-1 ¦ CAS no: 34454-97-2
1,1,2,2,3,3,4,4,4-nonafluorobutane-1-sulphonic acid EC / List no: 206-793-1 ¦ CAS no: 375-73-5
UPDATE JULY 2019
Updated REACH Candidate list
203-772-9 | 110-49-6 | 2-methoxyethyl acetate: methylglycol acetate | Not registered under REACH | Repr 1B | 2 | Solvent and intermediate; may be used in medical devices but only at < 0.1%
UPDATE APRIL 2019
Alignment with the IEC 62474 / BOMcheck screening of the 197 REACH Candidate List substances
215-125-8 | 1303-86-2 | diboron trioxide: boric oxide | Was originally considered as not relevant for MD (N). Has now been included in list (1). | Repr 1B | 1 | Diboron trioxide, also known as boric oxide, and is the anhydride form of boric acid. Assessment by BOMcheck and IEC 62474 identifies that Borates are used in wooden veneer sheets and pressed wooden panels, as a constituent within the starch adhesive, as a flame retardant and also as stabilizer in aminoplastic resins. Borates are used as a biocide in professional and industrial wood preservation. Diboron trioxide is used used as a flame retardant in wood products, paper, cotton and other plant-derived materials.
202-506-9 | 96-45-7 | ethylene thiourea: imidazolidine-2-thione: 2-imidazoline-2-thiol | Was originally considered as not relevant for MD (N). Has now been included in list (2). | Repr 1B | 2 | This substance is used as an accelerator in rubber vulcanising as the process with sulphur is extremely slow and impractical. Publications indicate that ethylene thiourea behaves as a catalyst and so, although it does take part in and is consumed by the chemical reactions, some is reformed and so this substance can remain (at >0.1%) in vulcanised rubber. Therefore, this substance may be present in devices in concentration > 0,1%. Assessment by BOMcheck and IEC 62474 identifies that Imidazolidine-2-thione is also used as a catalyst in some acrylic adhesive glues which may be used in adhesive tapes (for example, double sided adhesive tapes which may be used to hold the back-light in place in mobile phones). The adhesive glues are generally not cured at high temperatures. As a result, imidazolidine-2-thione may be detectable as a substance > 0.1% w/w in acrylic adhesive glue.
200-028-5 | 50-32-8, 63466-71-7 | benzo[a]pyrene: benzo[def]chrysene | Was originally considered as not relevant for MD (N). Has now been included in list (2). | Carc 1B, Repr 1B, Muta 1B | 2 | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1%
200-280-6 | 56-55-3, 1718-53-2 | benz[a]anthracene | Was originally considered as not relevant for MD (N). Has now been included in list (2). | Carc 1B | 2 | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1%
205-892-7 | 192-97-2 | benzo[e]pyrene | Was originally considered as not relevant for MD (N). Has now been included in list (2). | Carc 1B | 2 | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1%
205-910-3 | 205-82-3 | benzo[j]fluoranthene | Was originally considered as not relevant for MD (N). Has now been included in list (2). | Carc 1B | 2 | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1%
205-911-9 | 205-99-2 | benz[e]acephenanthrylene | Was originally considered as not relevant for MD (N). Has now been included in list (2). | Carc 1B | 2 | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1%
200-181-8 | 53-70-3 | dibenz[a,h]anthracene | Was originally considered as not relevant for MD (N). Has now been included in list (2). | Carc 1B | 2 | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1%
Information received from members
206-019-2 | 288-32-4 | imidazole | In January 2020 it was finally concluded it is not present in the cured adhesive and it has been removed from the list. | Repr 1B | N | Corrosion inhibitor (in water) and process chemical used to make imidazole derivatives
Updated classifications CLP (Evaluated by BOMcheck Substance List Working Group)
200-431-6 | 120-80-9 | 1,2-dihydroxybenzene; pyrocatechol | 13th ATP - October 2018 | Carc 1B | N | Commonly known as Catechol, this substance was added to the Proposition 65 List in July 2003 and is a process chemical which is used to produce anti-oxidants in rubber and polymerisation inhibitors. In these applications Catechol is reacted and is not present as a substance in the finished product. For example, when used as an intermediate in the production of stabilizers in the manufacture of polyurethane foam, vinyl acetate and butadiene, Catechol is reacted to form 4-tert-Butylcatechol (TBC). Catechol is also used as an antioxidant in tin plating processes to minimize the oxidation of the divalent tin by atmospheric oxygen. The tin plating processes are followed by several rinsing processes to remove excess process solution from the surface of the treated article. Therefore, Catechol is not detectable as a substance in, or on, the treated article.
200-836-8 | 75-07-0 | acetaldehyde; ethanal | 13th ATP - October 2018 | Carc 1B | N | Acetaldehyde; ethanal was added to the Proposition 65 List in April 1988. Acetaldehyde is manufactured for use as a process chemical and is also found naturally in fruit and vegetables. The first step of Ethanol breakdown by the liver produces acetaldehyde and it is the build-up of acetaldehyde in the body that contributes to the hangover symptoms which can occur after drinking alcohol. Acetaldehyde is used as a process intermediate and reagent in the synthesis of other chemicals that are used in adhesives, sealants, paints, coatings, paper products and plastic products. Unreacted acetaldehyde may be present as very low-level impurities in products which are manufactured from these synthesized chemicals. For example, PVAc flooring adhesives can contain up to 100-500 ppm acetaldehyde by weight of the adhesive. Acetaldehyde is also used as an ingredient in the manufacture of phenolic and urea resins. Another source of acetaldehyde is through thermal degradation of plastics such as PET (via oxidation of the vinyl end-group of the polymer chain) which can result in up to 3-4 ppm of acetaldehyde in the plastic. Manufacturers have developed ways of reducing the acetaldehyde content in PET because when PET is used in water bottles, even levels as low as 10-15 ppm acetaldehyde can be discernible in the taste of the water. Acetaldehyde is also found in liquids used in electronic cigarettes, in air fresheners, and in gasoline engine exhaust, and there have been several Proposition 65 settlement actions in California for these applications.
404-360-3 | 119313-12-1 | 2-benzyl-2-dimethylamino-4′- morpholinobutyrophenone | 13th ATP - October 2018 | Repr 1B | 1 | The substances is registered under REACH and is used as a UV curing agent in UV cured inks, UV cured coatings and UV cured adhesives. The substance can occur at levels up to 5% w/w in printing inks and up to 2.5 % w/w in adhesives. However, there is quite a lot of published research describing leaching out of unreacted photoinitiators from food packaging materials into food, which implies that reaction is not does not consume 100% of the reactants. It is quite possible, therefore, that some of the photoinitiator is not consumed and remains in its unreacted form. Polymer, adhesive and ink formulators need to ensure that they do not use too little photoinitiator as this will prevent full polymerisation and so to avoid this, the use of additional photoinitiator than is strictly required is a significant possibility, with a high likelihood that some will remain unreacted in the cured polymer. Also, it is possible that insufficient exposure to UV light of the correct wavelength will occur and would result in unreacted photoinitiator residues.
262-104-4 | 60207-90-1 | propiconazole (ISO); (2RS,4RS;2RS,4SR)-1-{[2-(2,4- dichlorophenyl)-4-propyl-1,3- dioxolan-2-yl]methyl}-1H1,2,4-triazole | 13th ATP - October 2018 | Repr 1B | N | Propiconazole is a triazole fungicide. It is used as an agricultural fungicide on various crops including corn, rice, peanuts and grasses grown for seed. It is also used as a wood preservative. There are no other recorded uses of the substance.
| 148477-71-8 | spirodiclofen (ISO); 3-(2,4-dichlorophenyl)-2-o xo1-oxaspiro[4.5]dec-3-en-4-yl 2,2-dimethylbutyrate | 13th ATP - October 2018 | Carc 1B | N | Spirodiclofen was added to the Proposition 65 List in October 2010. Spirodiclofen is an agricultural insecticide. It is used to kill mites on citrus, pome fruit, stone fruit and tree nuts. There have been no Proposition 65 settlement actions in California for Spirodiclofen. There are no other recorded uses of the substance. ?
214-012-0 | 1072-63-5 | 1-vinylimidazole | 13th ATP - October 2018 | Repr 1B | N | 1-vinylimidazole is used as UV curing agent in UV cured inks, UV cured coatings and UV cured adhesives. It is a highly reactive UV curing agent and becomes fully reacted during the UV curing process via copolymerization. 1-vinylimidazole is also used as an epoxy accelerator wherein the substance also becomes fully reacted. 1-vinylimidazole is also used as an ingredient in the manufacture of molecularly imprinted polymers, in the manufacture of magnetic-poly(divinylbenzene-1-vinylimidazole) microbeads, and in the manufacture of ion imprinted polyvinylimidazole-silica hybrid copolymer. None of these applications result in 1-vinylimidazole being present > 0.1% w/w in manufactured articles.
Updated REACH Candidate list
239-139-9 | 15087-24-8 | 1,7,7-trimethyl-3-(phenylmethylene)bicyclo[2.2.1]heptan-2-one | SVHC - 15/01/2019 | ED | N | Commonly known as 3-benzylidene camphor or 3-MBC, the substance is a chemical UV filter which is used to provide UV protection in sunscreen lotions, sunblocks, hair care products and other cosmetic products claiming a Sun Protection Factor (SPF)
401-720-1 | 6807-17-6 | 2,2-bis(4'-hydroxyphenyl)-4-methylpentane | SVHC - 15/01/2019 | Repr 1B | N | May be used as an alternative to Bisphenol A (BPA). As such, BisP-MIBK may be found in thermal paper applications up to 3% by weight of the paper. BisP-MIBK may also be used as an ingredient in the manufacture of specialist polycarbonate plastic and specialist epoxy resins, however the residual level of BisP-MIBK in the manufacture of these specialist materials is below 0.01%.
205-916-6 | 207-08-9 | Benzo[k]fluoranthene | SVHC - 15/01/2019 | Carc 1B | 2 | Polycyclic Aromatic Hydrocarbon (PAH) which may be present as impurities in extender oils which may be used as plasticisers in rubber and plastic, and as impurities in black colorants (e.g. Carbon Black) which may be used in rubber and plastic, but likely to be <<0.1%
UPDATE JUNE 2018
Updated REACH Candidate list
201-545-9 | 84-61-7 | DCHP (Dicyclohexyl phthalate) | SVHC - 27/06/2018 | ED (new) Repr 1B | 1 | Plasticiser
234-541-0 | 12008-41-2 | Disodium octaborate | SVHC - 27/06/2018 | Repr 1B | 1 | Flame retardant used in paper, card, paint, wood and also added to cellulose insulation
231-100-4 | 7439-92-1 | Lead metal | SVHC - 27/06/2018 | Repr 1A | 1 | See the COCIR study on uses of metallic lead in medical devices
RINA RESEARCH - February 2018
428-010-4 | 82413-20-5 | (<i>E</i>)-3-[1-[4-[2-(dimethylamino)ethoxy]phenyl]-2-phenylbut-1-enyl]phenol |  |  | N | Medicine
226-908-9 | 5543-58-8 | (<i>R</i>)-4-hydroxy-3-(3-oxo-1-phenylbutyl)-2-benzopyrone |  |  | N | Warfarin (drug)
226-907-3 | 5543-57-7 | (<i>S</i>)-4-hydroxy-3-(3-oxo-1-phenylbutyl)-2-benzopyrone |  |  | N | Warfarin (drug)
414-200-4 | 119738-06-6 | (±) tetrahydrofurfuryl (<i>R</i>)-2-[4-(6-chloroquinoxalin-2-yloxy)phenyloxy]propionate |  |  | N | Life science chemical
429-740-6 | 40722-80-3 | (2-chloroethyl)(3-hydroxypropyl)ammonium chloride |  |  | N | Water soluble surfactant
284-351-7 | 84852-39-1 | (2-ethylhexanoato-<i>O</i>)(isodecanoato-<i>O</i>)nickel |  |  | N | Catalyst (e.g. rubber)
287-470-2 | 85508-45-8 | (2-ethylhexanoato-<i>O</i>)(isononanoato-<i>O</i>)nickel |  |  | N | Catalyst (e.g. rubber)
285-698-7 | 85135-77-9 | (2-ethylhexanoato-<i>O</i>)(neodecanoato-<i>O</i>)nickel |  |  | N | Catalyst (e.g. rubber)
405-020-7 | 105024-66-6 | (4-ethoxyphenyl)(3-(4-fluoro-3-phenoxyphenyl)propyl)dimethylsilane |  |  | N | Pesticide
402-060-7 | 108225-03-2 | (6-(4-hydroxy-3-(2-methoxyphenylazo)-2-sulfonato-7-naphthylamino)-1,3,5-triazin-2,4-diyl)bis[(amino-1-methylethyl)ammonium] formate |  |  | 3 | Registered by one company (Archroma Germany GmbH) who are manufacturers of colours for textiles, paper, coatings (paint), adhesives and sealants
284-348-0 | 84852-36-8 | (isodecanoato-<i>O</i>)(isononanoato-<i>O</i>)nickel |  |  | N | Catalyst (e.g. rubber)
285-909-2 | 85166-19-4 | (isodecanoato-<i>O</i>)(isooctanoato-<i>O</i>)nickel |  |  | N | Catalyst (e.g. rubber)
287-471-8 | 85508-46-9 | (isononanoato-<i>O</i>)(isooctanoato-<i>O</i>)nickel |  |  | N | Catalyst (e.g. rubber)
287-592-6 | 85551-28-6 | (isononanoato-<i>O</i>)(neodecanoato-<i>O</i>)nickel |  |  | N | Catalyst (e.g. rubber)
284-347-5 | 84852-35-7 | (isooctanoato-<i>O</i>)(neodecanoato-<i>O</i>)nickel |  |  | N | Catalyst (e.g. rubber)
401-500-5 | 118658-99-4 | (methylenebis(4,1-phenylenazo(1-(3-(dimethylamino)propyl)-1,2-dihydro-6-hydroxy-4-methyl-2-oxopyridine-5,3-diyl)))-1,1'-dipyridinium dichloride dihydrochloride |  |  | 1 | Azo dye used in textiles, paper
265-748-4 | 65405-96-1 | [µ-[carbonato(2-)-<i>O</i>:<i>O</i>’]] dihydroxy trinickel |  |  | N | Process chemical
235-715-9 | 12607-70-4 | [carbonato(2-)] tetrahydroxytrinickel |  |  | N | May occur inside discharged NiCd batteries, process chemical
435-470-1 | 27366-72-9 | <i>N,N</i>-(dimethylamino)thioacetamide hydrochloride |  |  | N | Medicine & biochemical
202-959-2 | 101-61-1 | <i>N</i>,<i>N</i>,<i>N</i>',<i>N</i>'-tetramethyl-4,4'-methylendianiline |  |  | N | Michler's base - intermediate used to make dyes
210-338-2 | 613-35-4 | <i>N</i>,<i>N</i>'-diacetylbenzidine, N,N-diacetylacetamide |  |  | N | Process chemical
204-826-4 | 127-19-5 | <i>N</i>,<i>N</i>-dimethylacetamide, N, N-dimethylacetamide |  |  | N | Process chemical
200-679-5 | 68-12-2 | <i>N</i>,<i>N</i>-dimethylformamide: dimethyl formamide |  |  | N | Process chemical
200-316-0 | 57-14-7 | <i>N</i>,<i>N</i>-dimethylhydrazine |  |  | N | Process chemical
424-550-1 | 84245-12-5 | <i>N</i>-[6,9-dihydro-9-[[2-hydroxy-1-(hydroxymethyl)ethoxy]methyl]-6-oxo-1<i>H</i>-purin-2-yl]acetamide |  |  | N | Unknown, possibly as pharmaceutical intermediate
212-828-1 | 872-50-4 | <i>N</i>-methyl-2-pyrrolidone: 1-methyl-2-pyrrolidone, NMP |  |  | N | Solvent
201-182-6 | 79-16-3 | <i>N</i>-methylacetamide |  |  | N | Process chemical
204-624-6 | 123-39-7 | <i>N</i>-methylformamide |  |  | N | Process chemical
432-750-3 |  | <i>O</i>-hexyl-<i>N</i>-ethoxycarbonylthiocarbamate |  |  | N | Uncertain, possibly a biocide
434-350-4 | 103122-66-3 | <i>O</i>-isobutyl-<i>N</i>-ethoxy carbonylthiocarbamate |  |  | N | Uncertain, possibly a biocide
202-429-0 | 95-53-4 | <i>o</i>-toluidine: 2-aminotoluene |  |  | N | Process chemical
424-280-2 | 51594-55-9 | <i>R</i>-1-chloro-2,3-epoxypropane |  |  | N | Epichlorohydrin - process chemical
404-660-4 | 57044-25-4 | <i>R</i>-2,3-epoxy-1-propanol |  |  | N | Glycidol - process chemical
433-580-2 | 214353-17-0 | 1-(2-amino-5-chlorophenyl)-2,2,2-trifluoro-1,1-ethanediol, hydrochloride: [containing ≥ 0.1 % 4-chloroaniline (EC No 203-401-0)] |  |  | N | Intermediate
202-486-1 | 96-18-4 | 1,2,3-trichloropropane |  |  | N | Solvent and intermediate
271-093-5 | 68515-50-4 | 1,2-benzenedicarboxylic acid, dihexyl ester, branched and linear |  |  | 1 | Plasticiser
284-032-2 | 84777-06-0 | 1,2-benzenedicarboxylic acid, dipentylester, branched and linear |  |  | 1 | Plasticiser
276-158-1 | 71888-89-6 | 1,2-benzenedicarboxylic acid: di-C<sub>6-8</sub>-branched alkylesters, C<sub>7</sub>-rich |  |  | 1 | Plasticiser
271-084-6 | 68515-42-4 | 1,2-benzenedicarboxylic acid: di-C<sub>7-11</sub>-branched and linear alkylesters |  |  | 1 | Plasticiser
203-977-3 | 112-49-2 | 1,2-bis(2-methoxyethoxy)ethane: TEGDME: triethylene glycol dimethyl ether: triglyme |  |  | N | Solvent previously used inside lithium-ion batteries, now very uncommon
202-479-3 | 96-12-8 | 1,2-dibromo-3-chloropropane |  |  | N | Process chemical
203-444-5 | 106-93-4 | 1,2-dibromoethane |  |  | N | Process chemical
203-458-1 | 107-06-2 | 1,2-dichloroethane: ethylene dichloride |  |  | N | Process chemical
201-152-2 | 78-87-5 | 1,2-dichloropropane: propylene dichloride |  |  | N | Process chemical
211-076-1 | 629-14-1 | 1,2-diethoxyethane |  |  | 3 | SVHC: Ethylene glycol diethyl ether is a important solvent, used for ink, paint and coating industry. Mainly used in polymer, electrochemistry, boracium chemistry, resin, nitro cellulose, it is also used in surface treatment, halogen analysis in gasoline, acetic acid recycle in diluted acetic acid, depaint agent, thinner, flushing agent, stabilizer, anti oxidant , thickener of lube. Boiling point 120°C so it is not likely to be present in polymers or coatings made using this solvent. May be an additive in hydraulic and brake fluids.
203-794-9 | 110-71-4 | 1,2-dimethoxyethane: ethylene glycol dimethyl ether: EGDME |  |  | N | SVHC, Li-metal battery electrolyte (no patient contact)
| 540-73-8 | 1,2-dimethylhydrazine: salts of hydrazine |  |  | N | Process chemical
219-514-3 | 2451-62-9 | 1,3,5-tris(oxiranylmethyl)-1,3,5-triazine-2,4,6(1<i>H</i>,3<i>H</i>,5<i>H</i>)-trione: TGIC |  |  | N | Cross-linking / curing agent used to make epoxy resins and in polyester powder coatings. Cures when used so if fully reacted should not be present in finished product.
423-400-0 | 59653-74-6 | 1,3,5-tris-[(2<i>S</i> and 2<i>R</i>)-2,3-epoxypropyl]-1,3,5-triazine-2,4,6-(1<i>H</i>,3<i>H</i>,5<i>H</i>)-trione |  |  | N | Cross-linking / curing agent used to make epoxy resins and in polyester powder coatings. Cures when used so if fully reacted should not be present in finished product.
203-450-8 | 106-99-0 | 1,3-butadiene: buta-1,3-diene |  |  | N | Gas - monomer
202-491-9 | 96-23-1 | 1,3-dichloro-2-propanol |  |  | 3 | Process chemical, but can occur as an impurity breakdown product in paper and paints https://monographs.iarc.fr/ENG/Monographs/vol101/mono101-011.pdf
214-317-9 | 1120-71-4 | 1,3-propanesultone: 1,2-oxathiolane 2,2-dioxide |  |  | N | Li-ion battery electrolyte additive, usually at <0.1% (no patient contact)
219-603-7 | 2475-45-8 | 1,4,5,8-tetraaminoanthraquinone: C.I. Disperse Blue 1 |  |  | 1 | Dye used in plastics, fabrics, etc.
212-121-8 | 764-41-0 | 1,4-dichlorobut-2-ene |  |  | N | Monomer for chloroprene rubber
203-445-0 | 106-94-5 | 1-bromopropane: n-propyl bromide |  |  | N | Process chemical
203-439-8 | 106-89-8 | 1-chloro-2,3-epoxypropane: epichlorhydrin |  |  | N | Process chemical and biocide
200-730-1 | 70-25-7 | 1-methyl-3-nitro-1-nitrosoguanidine |  |  | N | Laboratory reagent
203-867-5 | 111-41-1 | 2-(2-aminoethylamino)ethanol: (AEEA) |  |  | N | Process chemical
214-237-4 | 1116-54-7 | 2,2'-(nitrosoimino)bisethanol |  |  | N | Research chemical
215-979-1 | 1464-53-5 | 2,2'-bioxirane: 1,2:3,4-diepoxybutane |  |  | N | Butadiene diepoxide is used as a research chemical, as a chemical intermediate, as a curing agent for polymers, as a cross-linking agent for textile fabrics, and to prevent microbial spoilage. It is also been used to synthesize erythritol and pharmaceuticals.
202-918-9 | 101-14-4 | 2,2'-dichloro-4,4'-methylenedianiline: 4,4'-methylene bis(2-chloroaniline): salts of 2,2'-dichloro-4,4'-methylenedianiline: salts of 4,4'-methylenebis(2-chloroaniline) |  |  | 1 | "MOCA", used mainly to make polyurethanes (some used also for epoxy resins), but residues may exceed 0.1% acccording to ECHA
202-480-9 | 96-13-9 | 2,3-dibromopropan-1-ol: 2,3-dibromo-1-propanol |  |  | N | Process chemical
210-013-5 | 602-01-7 | 2,3-dinitrotoluene |  |  | N | Process chemical (and used in explosives)
209-128-3 | 556-52-5 | 2,3-epoxypropan-1-ol: glycidol: oxiranemethanol |  |  | N | Process chemical
203-441-9 | 106-91-2 | 2,3-epoxypropyl methacrylate: glycidyl methacrylate |  |  | N | Process chemical primarily used as a reactive building block to produce polymer coatings, finishes, and adhesives.
221-221-0 | 3033-77-0 | 2,3-epoxypropyltrimethylammonium chloride ...%: glycidyl trimethylammonium chloride ...% |  |  | 2 | Cationisation of starches which are added in paper to give paper better surface quality. Used reactively so that only up to 25mg/kg remain, so <0.1%
205-282-0 | 137-17-7 | 2,4,5-trimethylaniline |  |  | N | Process chemical
| 21436-97-5 | 2,4,5-trimethylaniline hydrochloride |  |  | N | Process chemical
254-323-9 | 39156-41-7 | 2,4-diaminoanisole sulphate |  |  | N | Process chemical
210-406-1 | 615-05-4 | 2,4-diaminoanisole: 4-methoxy-<i>m</i>-phenylenediamine |  |  | N | Process chemical
204-450-0 | 121-14-2 | 2,4-dinitrotoluene |  |  | N | Process chemical (and used in explosives)
210-581-4 | 619-15-8 | 2,5-dinitrotoluene |  |  | N | Process chemical (and used in explosives)
210-106-0 | 606-20-2 | 2,6-dinitrotoluene |  |  | N | Process chemical (and used in explosives)
| 72319-19-8 | 2,7-naphthalenedisulfonic acid, nickel(II) salt |  |  | 1 | May be used as a textile dye but uncertain and unlikely
420-580-2 | 151798-26-4 | 2-[2-hydroxy-3-(2-chlorophenyl)carbamoyl-1-naphthylazo]-7-[2-hydroxy-3-(3-methylphenyl)carbamoyl-1-naphthylazo]fluoren-9-one |  |  | 1 | Pigment (used in some photocopiers)
200-855-1 | 75-26-3 | 2-bromopropane |  |  | N | Process chemical
425-150-8 | 94723-86-1 | 2-butyryl-3-hydroxy-5-thiocyclohexan-3-yl-cyclohex-2-en-1-one |  |  | N | Intermediate
433-890-8 | 2040-90-6 | 2-chloro-6-fluoro-phenol |  |  | N | Intermediate
203-804-1 | 110-80-5 | 2-ethoxyethanol: ethylene glycol monoethyl ether |  |  | N | SVHC: Solvent
203-839-2 | 111-15-9 | 2-ethoxyethyl acetate: ethylglycol acetate |  |  | N | Solvent and intermediate
231-480-1 | 7580-31-6 | 2-ethylhexanoic acid, nickel salt |  |  | N | Process chemical
239-622-4 | 15571-58-1 | 2-ethylhexyl 10-ethyl-4,4-dioctyl-7-oxo-8-oxa-3,5-dithia-4-stannatetradecanoate (DOTE) |  |  | 1 | PVC stabiliser (SVHC)
279-452-8 | 80387-97-9 | 2-ethylhexyl[[[3,5-bis(1,1-dimethylethyl)-4-hydroxyphenyl]methyl]thio]acetate |  |  | 4 | Unknown, not registered in the EU. Likely to be used as a process or laboratory chemical
201-963-1 | 90-04-0 | 2-methoxyaniline: <i>o</i>-anisidine |  |  | N | Process chemical
203-713-7 | 109-86-4, 109-87-5 | 2-methoxyethanol: ethylene glycol monomethyl ether |  |  | N | SVHC: Solvent and anti-icing additive in aviation fuels
203-772-9 | 110-49-6 | 2-methoxyethyl acetate: methylglycol acetate |  |  | N | Solvent and intermediate
216-455-5 | 1589-47-5 | 2-methoxypropanol |  |  | N | Solvent
274-724-2 | 70657-70-4 | 2-methoxypropyl acetate |  |  | N | Solvent
400-600-6 | 71868-10-5 | 2-methyl-1-(4-methylthiophenyl)-2-morpholinopropan-1-one |  |  | 1 | Photoinitiator resulting in inclusion into a matrix, including application in coatings, adhesives and inks.
200-878-7 | 75-55-8 | 2-methylaziridine: propyleneimine |  |  | N | Intermediate in the paper, textile, rubber, and pharmaceutical industries
202-080-4 | 91-59-8 | 2-naphthylamine |  |  | N | Banned by REACH. Has been used as an antioxidant and as a process chemical
209-474-5 | 581-89-5 | 2-nitronaphthalene |  |  | N | Lab chemical
201-209-1 | 79-46-9 | 2-nitropropane |  |  | N | Solvent and fuel
201-853-3 | 88-72-2 | 2-nitrotoluene |  |  | N | Process chemical
202-109-0 | 91-94-1 | 3,3'-dichlorobenzidine: 3,3'-dichlorobiphenyl-4,4'-ylenediamine: salts of 3,3'-dichlorobenzidine: salts of 3,3'-dichlorobiphenyl-4,4'-ylenediamine |  |  | N | Intermediate used to make yellow pigments, discontinued in the USA and EU
204-355-4 | 119-90-4 | 3,3'-dimethoxybenzidine: <i>o</i>-dianisidine: salts of 3,3'-dimethoxybenzidine: salts of <i>o</i>-dianisidine |  |  | N | Used almost exclusively as a chemical intermediate for producing dyes and pigments. Also used as a chemical intermediate to produce o-dianisidine diisocyanate for use in adhesives and as a component of polyurethanes. Dyes used for paper, plastics, rubber, and textiles
210-222-1 | 610-39-9 | 3,4-dinitrotoluene |  |  | N | Process chemical
210-566-2 | 618-85-9 | 3,5-dinitrotoluene |  |  | N | Process chemical
225-918-0 | 5146-66-7 | 3,7-dimethylocta-2,6-dienenitrile |  |  | N | Intermediate
205-057-7 | 132-32-1 | 3-amino-9-ethyl carbazole: 9-ethylcarbazol-3-ylamine: |  |  | N | Chromogenic stain for immunohistochemistry used to visualise secondary antibodies
421-150-7 | 143860-04-2 | 3-ethyl-2-methyl-2-(3-methylbutyl)-1,3-oxazolidine |  |  | N | Moisture scavenger. Reactive diluent for urethane prepolymers
200-340-1 | 57-57-8 | 3-propanolide: 1,3-propiolactone |  |  | N | Disinfectant used in vapor form and intermediate
209-321-2 | 569-61-9 | 4,4'-(4-iminocyclohexa-2,5-dienylidenemethylene)dianiline hydrochloride: C.I. Basic Red 9 |  |  | 1 | Dye
204-358-0 | 119-93-7 | 4,4'-bi-<i>o</i>-toluidine |  |  | N | Process chemical
202-027-5 | 90-94-8 | 4,4'-bis(dimethylamino)benzophenone: Michler's ketone |  |  | 1 | SVHC
202-974-4 | 101-77-9 | 4,4'-diaminodiphenylmethane: 4,4'-methylenedianiline |  |  | N | SVHC - MDA, adhesive curing agent, not present in product
401-720-1 | 6807-17-6 | 4,4-isobutylethylidenediphenol |  |  | N | Biochemistry
212-658-8 | 838-88-0 | 4,4'-methylenedi-<i>o</i>-toluidine |  |  | N | Intermediate
202-977-0 | 101-80-4 | 4,4'-oxydianiline and its salts: <i>p</i>-aminophenyl ether |  |  | N | Used to make various polymers
205-370-9 | 139-65-1 | 4,4'-thiodianiline and its salts |  |  | N | Used to make pigments and used as a medicine
202-591-2 | 97-56-3 | 4-<i>o</i>-tolylazo-<i>o</i>-toluidine: 4-amino-2',3-dimethylazobenzene: fast garnet GBC base: AAT: <i>o</i>-aminoazotoluene |  |  | N | Used to make dyes and used as a laboratory stain
202-696-3 | 98-73-7 | 4-<i>tert</i>-butylbenzoic acid |  |  | 1 | PVC stabiliser and resin modifier (alkyds)
402-230-0 | 399-95-1 | 4-amino-3-fluorophenol |  |  | N | Intermediate
200-453-6 | 60-09-3 | 4-aminoazobenzene: 4-phenylazoaniline |  |  | 1 | Also called aniline yellow. Used as a yellow pigments and in inks including inks for inkjet printersmicroscopy. Used in insecticides, lacquers, varnishes, waxes, oil stains, and styrene resins.Also used for vital staining, in pyrotechnics for yellow colored smokes. It is also an intermediate in synthesis of other dyes
|  | Benzidine based azo dyes: 4,4'-diarylazobiphenyl dyes, with the exception of those specified elsewhere in this Annex |  |  | 1 | Dyes of various colours, although some are banned by REACH in fabrics with skin contact
202-441-6 | 95-69-2 | 4-chloro-<i>o</i>-toluidine |  |  | N | Process chemical
221-627-8 | 3165-93-3 | 4-chloro-<i>o</i>-toluidine hydrochloride |  |  | N | Process chemical
203-401-0 | 106-47-8 | 4-chloroaniline: |  |  | N | Process chemical
|  | Technical product – reaction mass of 4-methyl-<i>m</i>-phenylene diamine (EC No 202-453-1) and 2-methyl-<i>m</i>-phenylene diamine (EC No 212-513-9)] |  |  | N | Process chemical
202-453-1 | 95-80-7 | 4-methyl-<i>m</i>-phenylenediamine: 2,4-toluenediamine |  |  | N | Intermediate
202-204-7 | 92-93-3 | 4-nitrobiphenyl |  |  | N | Banned by EU REACH, originally used as a process chemical
210-025-0 | 602-87-9 | 5-nitroacenaphthene |  |  | N | A PAH compound (no intentional uses and will be <0.1% if present)
400-340-3 | 85136-74-9 | 6-hydroxy-1-(3-isopropoxypropyl)-4-methyl-2-oxo-5-[4-(phenylazo)phenylazo]-1,2-dihydro-3-pyridinecarbonitrile |  |  | 1 | Also called Disperse Orange 149
204-419-1 | 120-71-8 | 6-methoxy-<i>m</i>-toluidine: <i>p</i>-cresidine |  |  | N | Intermediate used to make dyes
429-400-7 | 199327-61-2 | 7-methoxy-6-(3-morpholin-4-yl-propoxy)-3<i>H</i>-quinazolin-4-one: [containing < 0.5 % formamide (EC No 200-842-0)] |  |  | N | Intermediate
309-851-5 | 101316-45-4 | Absorption oils, bicyclo arom. and heterocyclic hydrocarbon fraction: Wash Oil Redistillate: [A complex combination of hydrocarbons obtained as a redistillate from the distillation of wash oil. It consists predominantly of 2-ringed aromatic and heterocyclic hydrocarbons boiling in the range of approximately 260 °C to 290 °C (500 °F to 554 °F).] |  |  | N | Petrochemical
| 34256-82-1 | Acetochlor |  |  | N | Herbicide
201-173-7 | 79-06-1, 122775-19-3 | acrylamide: prop-2-enamide |  |  | N | Process chemical
203-466-5 | 107-13-1 | acrylonitrile |  |  | N | Process chemical
| 15972-60-8 | Alachlor |  |  | N | Banned herbicide
292-454-3 | 90622-53-0 | Alkanes, C<sub>12-26</sub>-branched and linear |  |  | N | Hydrocarbon liquid
270-651-5 | 68475-57-0 | Alkanes, C1-2: Petroleum gas |  |  | N | Hydrocarbon gases
292-456-4 | 90622-55-2 | Alkanes, C1-4, C3-rich: Petroleum gas |  |  | N | Hydrocarbon gases
270-652-0 | 68475-58-1 | Alkanes, C2-3: Petroleum gas |  |  | N | Hydrocarbon gases
270-653-6 | 68475-59-2 | Alkanes, C3-4: petroleum gas |  |  | N | Hydrocarbon gases
270-654-1 | 68475-60-5 | Alkanes, C4-5: Petroleum gas |  |  | N | Hydrocarbon gases
232-143-1 | 7789-09-5 | ammonium dichromate |  |  | 1 | SVHC, mainly used as a process chemical, but may rarely be used as a mordant for dying fabrics and textiles (may possibly exceed 0.1%)
221-470-5 | 3108-42-7 | ammonium nonadecafluorodecanoate |  |  | 1 | SVHC, water-proofing textiles
249-415-0 | 29081-56-9 | ammonium perfluorooctane sulfonate: ammonium heptadecafluorooctanesulfonate |  |  | 1 | used as a processing aid for manufacture of fluoropolymers such as PVDF. Also used as an additive in lubricating oils, in cleaning agents, for waterproofing and textile coatings and in liquid crystal displays. Occurs in water and stain resistant textiles. Will not exceed 0.1% in a material, but could as an ultra-thin waterproof coating
223-320-4 | 3825-26-1 | ammonium pentadeca- fluorooctanoate |  |  | 1 | used as a processing aid for manufacture of fluoropolymers such as PVDF. Also used as an additive in lubricating oils, in cleaning agents, for waterproofing and textile coatings and in liquid crystal displays. Occurs in water and stain resistant textiles. Will not exceed 0.1% in a material, but could as an ultra-thin waterproof coating
200-857-2 | 75-28-5 | isobutane |  |  | N | Gas
295-274-3 | 91995-14-1 | Anthracene oil, acid ext.: Anthracene Oil Extract Residue: [A complex combination of hydrocarbons from the base-freed fraction obtained from the distillation of coal tar and boiling in the range of approximately 325 °C to 365 °C (617 °F to 689 °F). It contains predominantly anthracene and phenanthrene and their alkyl derivatives.] |  |  | N
295-275-9 | 91995-15-2 | Anthracene oil, anthracene paste, anthracene fraction: Anthracene Oil Fraction: [A complex combination of hydrocarbons from the distillation of anthracene obtained by the crystallization of anthracene oil from bituminous high temperature tar and boiling in the range of 330°C to 350°C (626°F to 662°F). It contains chiefly anthracene, carbazole and phenanthrene.] |  |  | N
295-276-4 | 91995-16-3 | Anthracene oil, anthracene paste, carbazole fraction: Anthracene Oil Fraction: [A complex combination of hydrocarbons from the distillation of anthracene obtained by crystallization of anthracene oil from bituminous coal high temperature tar and boiling in the approximate range of 350°C to 360°C (662°F to 680°F). It contains chiefly anthracene, carbazole and phenanthrene.] |  |  | N
295-278-5 | 91995-17-4 | Anthracene oil, anthracene paste, distn. lights: Anthracene Oil Fraction: [A complex combination of hydrocarbons from the distillation of anthracene obtained by crystallization of anthracene oil from bituminous high temperature tar and boiling in the range of approximately 290°C to 340°C (554°F to 644°F). It contains chiefly trinuclear aromatics and their dihydro derivatives.] |  |  | N
292-603-2 | 90640-81-6 | Anthracene oil, anthracene paste: Anthracene Oil Fraction: [The anthracene-rich solid obtained by the crystallization and centrifuging of anthracene oil. It is composed primarily of anthracene, carbazole and phenanthrene.] |  |  | N
292-604-8 | 90640-82-7 | Anthracene oil, anthracene-low: Anthracene Oil Fraction: [The oil remaining after the removal, by a crystallization process, of an anthracene-rich solid (anthracene paste) from anthracene oil. It is composed primarily of two, three and four membered aromatic compounds.] |  |  | N
292-602-7 | 90640-80-5 | Anthracene oil: Anthracene oil: [A complex combination of polycyclic aromatic hydrocarbons obtained from coal tar having an approximate distillation range of 300 °C ot 400 °C (572 °F to 752 °F). Composed primarily of phenanthrene, anthracene and carbazole.] |  |  | N
201-549-0 | 84-65-1 | anthraquinone |  |  | 3 | Mainly as an intermediate to make dyes and pharmaceutocals. Uncommonly used as an organic scintillator in radiation detectors. Patient contact is very unlikely
309-957-1 | 101794-75-6 | Aromatic hydrocarbons, C<sub>20-28</sub>, polycyclic, mixed coal-tar pitch-polyethylene pyrolysis-derived: Pyrolysis Products: [A complex combination of hydrocarbons obtained from mixed coal tar pitch-polyethylene pyrolysis. Composed primarily of polycyclic aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>28</sub> and having a softening point of 100 °C to 220 °C (212 °F to 428 °F) according to DIN 52025.] |  |  | N
309-956-6 | 101794-74-5 | Aromatic hydrocarbons, C<sub>20-28</sub>, polycyclic, mixed coal-tar pitch-polyethylene-polypropylene pyrolysis-derived: Pyrolysis Products: [A complex combination hydrocarbons obtained from mixed coal tar pitch-polyethylene-polypropylene pyrolysis. Composed primarily of polycyclic aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>28</sub> and having a softening point of 100 °C to 220 °C (212 °F to 428 °F) according to DIN 52025.] |  |  | N
309-958-7 | 101794-76-7 | Aromatic hydrocarbons, C<sub>20-28</sub>, polycyclic, mixed coal-tar pitch-polystyrene pyrolysis-derived: Pyrolysis Products: [A complex combination of hydrocarbons obtained from mixed coal tar pitch-polystyrene pyrolysis. Composed primarily of polycyclic aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>28</sub> and having a softening point of 100 °C to 220 °C (212 °F to 428 °F) according to DIN 52025.] |  |  | N
268-618-5 | 68131-49-7 | Aromatic hydrocarbons, C<sub>6-10</sub>, acid-treated, neutralized: Low boiling point naphtha - unspecified |  |  | N
292-697-5 | 90989-41-6 | Aromatic hydrocarbons, C<sub>6-10</sub>, C<sub>8</sub>-rich: Light Oil Redistillate, low boiling |  |  | N
270-658-3 | 68475-70-7 | Aromatic hydrocarbons, C<sub>6-8</sub>, naphtha-raffinate pyrolyzate-derived: Low boiling point thermally cracked naphtha: [A complex combination of hydrocarbons obtained by the fractionation pyrolysis at 816°C (1500°F) of naphtha and raffinate. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>8</sub>, including benzene.] |  |  | N
297-401-8 | 93571-75-6 | Aromatic hydrocarbons, C<sub>7-12</sub>, C<sub>8</sub>-rich: Low boiling point cat-reformed naphtha: [A complex combination of hydrocarbons obtained by separation from the platformate-containing fraction. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>12</sub> (primarily C<sub>8</sub>) and can contain nonaromatic hydrocarbons, both boiling in the range of approximately 130°C to 200°C (266°F to 392°F).] |  |  | N
292-698-0 | 90989-42-7 | Aromatic hydrocarbons, C<sub>7-8</sub>, dealkylation products, distn. residues: Low boiling point naphtha - unspecified |  |  | N
295-279-0 | 91995-18-5 | Aromatic hydrocarbons, C<sub>8</sub>, catalytic reforming-derived: Low boiling point cat-reformed naphtha |  |  | N
292-694-9 | 90989-38-1 | Aromatic hydrocarbons, C<sub>8</sub>: Light Oil Redistillate, high boiling |  |  | N
292-695-4 | 90989-39-2 | Aromatic hydrocarbons, C<sub>8-10</sub>: Low boiling point naphtha - unspecified |  |  | N
295-281-1 | 91995-20-9 | Aromatic hydrocarbons, C<sub>8-9</sub>, hydrocarbon resin polymn. by-product: Light Oil Redistillate, high boiling: [A complex combination of hydrocarbons obtained from the evaporation of solvent under vacuum from polymerized hydrocarbon resin. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>8</sub> through C<sub>9</sub> and boiling in the range of approximately 120°C to 215°C (248°F to 419°F).] |  |  | N
295-551-9 | 92062-36-7 | Aromatic hydrocarbons, C<sub>9-12</sub>, benzene distn.: Light Oil Redistillate, high boiling |  |  | N
601-649-8 | 12001-28-4 | asbestos |  |  | N | Thermal insulator, brake limings, thermal gaskets, Banned in the EU
601-650-3 | 12001-29-5, 132207-32-0 | asbestos |  |  | N | Thermal insulator, brake limings, thermal gaskets, Banned in the EU
601-801-3 | 12172-73-5 | asbestos |  |  | N | Thermal insulator, brake limings, thermal gaskets, Banned in the EU
| 132207-32-0 | asbestos |  |  | N | Thermal insulator, brake limings, thermal gaskets, Banned in the EU
616-471-6 | 77536-66-4 | asbestos |  |  | N | Thermal insulator, brake limings, thermal gaskets, Banned in the EU
616-472-1 | 77536-67-5 | asbestos |  |  | N | Thermal insulator, brake limings, thermal gaskets, Banned in the EU
616-473-7 | 77536-68-6 | asbestos |  |  | N | Thermal insulator, brake limings, thermal gaskets, Banned in the EU
| 1912-24-9 | Atrazine |  |  | N | Herbicide
620-425-0 | 68049-83-2 | azafenidin (ISO): 2-(2,4-dichloro-5-prop-2-ynyloxyphenyl)-5,6,7,8-tetrahydro-1,2,4-triazolo[4,3-a]pyridin-3(2<i>H</i>)-one |  |  | N | Herbicide
203-102-5 | 103-33-3 | azobenzene |  |  | N
201-622-7 | 85-68-7 | BBP: benzyl butyl phthalate |  |  | 1 | Plasticiser
241-775-7 | 17804-35-2 | benomyl (ISO): methyl 1-(butylcarbamoyl)benzimidazol-2-ylcarbamate |  |  | N | Agricultural fungicide
200-280-6 | 56-55-3, 1718-53-2 | benz[<i>a</i>]anthracene |  |  | N | PAH (>0.1% if present)
205-911-9 | 205-99-2 | benz[<i>e</i>]acephenanthrylene |  |  | N | PAH (>0.1% if present)
200-753-7 | 71-43-2, 1076-43-3 | benzene |  |  | N | Solvent, should not occur in medical devices as will evaporate during production
202-199-1 | 92-87-5 | benzidine: 1,1'-biphenyl-4,4'-diamine: 4,4'-diaminobiphenyl: biphenyl-4,4'-ylenediamine |  |  | N | Process chemical
200-028-5 | 50-32-8, 63466-71-7 | benzo[<i>a</i>]pyrene: benzo[<i>def</i>]chrysene |  |  | N | PAH (>0.1% if present)
205-892-7 | 192-97-2 | benzo[<i>e</i>]pyrene |  |  | N | PAH (>0.1% if present)
205-910-3 | 205-82-3 | benzo[<i>j</i>]fluoranthene |  |  | N | PAH (>0.1% if present)
205-916-6 | 207-08-9 | benzo[<i>k</i>]fluoranthene |  |  | N | PAH (>0.1% if present)
266-023-5 | 65996-88-5 | Benzol forerunnings (coal): Light Oil Redistillate, low boiling: [The distillate from coke oven light oil having an approximate distillation range below 100°C (212°F). Composed primarily of C<sub>4</sub> to C<sub>6</sub> aliphatic hydrocarbons.] |  |  | N
215-133-1 | 1304-56-9 | beryllium oxide |  |  | 3 | No patient contact. Beryllium oxide is used as a Beryllium oxide ceramics having excellent electrical insulation properties and a high thermal conductivity and is used in some power semiconductors. One report lists uses as guide bores of laser scalpels and parts of MRI and CAT scanners.
231-150-7 | 7440-41-7 | beryllium metal |  |  | 1 | CuBe alloy is the main alloy used For its high electrical and/or thermal conductivity, high strength and hardness, non-magnetic properties, and good corrosion and fatigue resistance, but is unlikely in patient contact materials. Nickel-beryllium alloy is used as springs in thermostats, bellows, diaphragms, burn-in connectors and sockets. Beryllium metal is also used as X-ray tube windows. BeAl lightweight high strength alloys are used in hard disc drives, aerospace, space, etc. One unusual repoprted use is in loudspeakers.
|  | beryllium compounds with the exception of aluminium beryllium silicates, and with those specified elsewhere in this Annex |  |  | N | No, all uses are as process chemicals only, except oxide
207-612-9 | 485-31-4 | binapacryl (ISO): 2-<i>sec</i>-butyl-4,6-dinitrophenyl-3-methylcrotonate |  |  | N | Miticide and fungicide
202-110-6 | 91-95-2 | biphenyl-3,3',4,4'-tetrayltetraamine: diaminobenzidine |  |  | N | Process chemical
202-177-1 | 92-67-1 | biphenyl-4-ylamine: xenylamine: 4-aminobiphenyl: salts of biphenyl-4-ylamine: salts of xenylamine: salts of 4-aminobiphenyl |  |  | N | Process chemical
276-205-6 | 71957-07-8 | bis(.sc.d.sc.-gluconato-<i>O</i><sup>1</sup>,<i>O</i><sup>2</sup>)nickel |  |  | N | Pharmaceutical and veterinary drug
204-211-0 | 117-81-7 | bis(2-ethylhexyl) phthalate: di-(2-ethylhexyl) phthalate: DEHP |  |  | 1 | Plasticiser
203-924-4 | 111-96-6 | bis(2-methoxyethyl) ether |  |  | 2 | SVHC: Solvent for battery electrolytes (lithium ion and other). May also be used as a component of inks, but will not be present in printed material.
204-212-6 | 117-82-8 | bis(2-methoxyethyl) phthalate |  |  | 1 | SVHC: Phthalate plasticiser in polymeric materials and paints, lacquers and varnishes, including printing inks
208-832-8 | 542-88-1 | bis(chloromethyl) ether: oxybis(chloromethane) |  |  | N | Intermediate (largely discontinued)
201-245-8 | 80-05-7 | bisphenol A: 4,4’-isopropylidenediphenol |  | Repr. 1B | 1 | Used to make polymers such as polycarbonate and epoxy resins, but usually present at <0.1%. Uncommon additive in PVC at about 0.1%
233-139-2 | 10043-35-3 | boric acid |  | Repr. 1B | 1 | Flame retardant in paper and wood and also used paints, inks and adhesives. Used as a preservative in wood preservative, preservatives for textile, paper, leather, rubber and polymerised materials. Also used in lubricants and other materials.The substance can also be used in contact lens storage solutions and multi-purpose disinfecting solutions.
234-343-4 | 11113-50-1 | boric acid |  | Repr. 1B | 1 | Flame retardant in paper and wood and also used paints, inks and adhesives. Used as a preservative in wood preservative, preservatives for textile, paper, leather, rubber and polymerised materials. Also used in lubricants and other materials.The substance can also be used in contact lens storage solutions and multi-purpose disinfecting solutions.
259-980-5 | 56073-10-0 | brodifacoum (ISO): 4-hydroxy-3-(3-(4'-bromo-4-biphenylyl)-1,2,3,4-tetrahydro-1-naphthyl)coumarin |  |  | N | Rat poison
249-205-9 | 28772-56-7 | bromadiolone (ISO); 3-[3-(4′-bromobiphenyl-4-yl)-3-hydroxy-1-phenylpropyl]-4-hydroxy-2H-chromen-2-one |  |  | N | Rat poison
209-800-6 | 593-60-2 | bromoethylene |  |  | N | Gas
| 34492-97-2 | bunsenite |  |  | N | Natural nickel oxide
203-448-7 | 106-97-8 | butane |  |  | N | Gas
208-953-6 | 548-62-9 | C.I. Basic Violet 3: 4-[4,4'-bis(dimethylamino) benzhydrylidene]cyclohexa-2,5-dien-1-ylidene]dimethylammonium chloride | C.I. Basic Violet 3 with ≥ 0.1 % of Michler's ketone (EC no. 202-027-5) |  | 1 | Pigment
231-152-8 | 7440-43-9 | cadmium (pyrophoric) | cadmium (non-pyrophoric) |  | 1 | Banned by RoHS, so should be absent in electrical equipment. Also restricted by REACH but with exemptions. Used as cadmium plating fasteners, connectors and as a component of very flexible alloys and low melting point alloys.
208-168-9 | 513-78-0 | cadmium carbonate |  |  | N | Process chemical
233-296-7 | 10108-64-2, 35658-65-2 | cadmium chloride |  |  | N | Process chemical
232-222-0 | 7790-79-6 | cadmium fluoride |  |  | N | Process chemical
244-168-5 | 21041-95-2 | cadmium hydroxide: cadmium dihydroxide |  |  | N | Process chemical
233-710-6 | 10022-68-1, 10325-94-7 | cadmium nitrate: cadmium dinitrate |  |  | N | Process chemical
215-146-2 | 1306-19-0 | cadmium oxide (non-pyrophoric) |  |  | 2 | Power contacts (RoHS exempt), but no patient contact
233-331-6 | 7790-84-3, 10124-36-4, 31119-53-6 | cadmium sulphate |  |  | N | Process chemical
215-147-8 | 1306-23-6 | cadmium sulphide |  |  | 1 | Banned by RoHS, so should be absent in electrical equipment. Also restricted by REACH but with exemptions. Used as a yellow pigment in plastics, ceramics and in inks.
237-366-8 | 13765-19-0 | calcium chromate |  |  | 1 | Pigment, corrosion inhibitor, electroplating chemical, photochemical processing and industrial waste treatment. Restricted by RoHS
219-363-3 | 2425-06-1 | captafol (ISO): 1,2,3,6-tetrahydro-<i>N</i>-(1,1,2,2-tetrachloroethylthio)phthalimide |  |  | N | Agricultural fungicide
229-879-0 | 6804-07-5 | carbadox (INN): methyl 3-(quinoxalin-2-ylmethylene)carbazate 1,4-dioxide: 2-(methoxycarbonylhydrazonomethyl)quinoxaline 1,4-dioxide |  |  | N | Vetenary drug
234-232-0 | 10605-21-7 | carbendazim (ISO): methyl benzimidazol-2-ylcarbamate |  |  | N | Agricultural fungicide
240-286-6 | 16118-49-3 | carbetamide (ISO): (R)-1-(ethylcarbamoyl)ethyl carbanilate; (2R)-1-(ethylamino)-1-oxopropan-2-yl phenylcarbamate |  |  | N | Herbicide
211-128-3 | 630-08-0 | carbon monoxide |  |  | N | Gas
240-408-8 | 16337-84-1 | carbonic acid, nickel salt |  |  | N | Process chemical
| 12789-03-6 | Chlordane |  |  | N | Pesticide
| 57-74-9 | Chlordane, – cis- or trans- |  |  | N | Pesticide
203-480-1 | 107-30-2 | chlormethyl methyl ether: chlorodimethyl ether |  |  | N | Process chemical and solvent
425-970-6 | 3724-43-4 | chloro-<i>N,N</i>-dimethylformiminium chloride |  |  | N | Process chemical
223-003-0 | 3691-35-8 | chlorophacinone (ISO): 2-[(4-chlorophenyl)(phenyl)acetyl]-1H-indene-1,3(2H)-dione |  |  | N | Rat poison
204-818-0 | 126-99-8 | chloroprene (stabilised): 2-chlorobuta-1,3-diene (stabilised) |  |  | N | Monomer used to make rubbers
215-607-8 | 1333-82-0 | chromium (VI) trioxide |  |  | N | Process chemical
239-056-8 | 14977-61-8 | chromyl dichloride: chromic oxychloride |  |  | N | Process chemical
205-923-4 | 218-01-9, 1719-03-5 | chrysene |  |  | 2 | PAH, impurity in rubber and plastic but likely to be <<0.1%
242-161-1 | 18283-82-4 | citric acid, ammonium nickel salt |  |  | N | Process chemical and catalyst
245-119-0 | 22605-92-1 | citric acid, nickel salt |  |  | N | Process chemical and catalyst
265-064-6 | 64741-62-4 | Clarified oils (petroleum), catalytic cracked: Heavy Fuel oil: [A complex combination of hydrocarbons produced as the residual fraction from distillation of the products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub> and boiling above approximately 350 °C (662 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
269-782-0 | 68333-26-6 | Clarified oils (petroleum), hydrodesulfurized catalytic cracked: Heavy Fuel oil; [A complex combination of hydrocarbons obtained by treating catalytic cracked clarified oil with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists of hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub> and boiling above approximately 350 °C (662 °F). This stream is likely to contain 5 wt. % or more of 4-to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
302-682-8 | 94114-47-3 | Coal liquids, liq. solvent extn. soln.: [The product obtained by filtration of coal mineral matter and undissolved coal from coal extract solution produced by digesting coal in a liquid solvent. A black, viscous, highly complex liquid combination composed primarily of aromatic and partly hydro-genated aromatic hydrocarbons, aromatic nitrogen compounds, aromatic sulfur compounds, phenolic and other aromatic oxygen compounds and their alkyl derivatives.] |  |  | N
302-683-3 | 94114-48-4 | Coal liquids, liq. solvent extn.: [The substantially solvent-free product obtained by the distillation of the solvent from filtered coal extract solution produced by digesting coal in a liquid solvent. A black semi-solid, composed primarily of a complex combination of condensed-ring aromatic hydrocarbons, aromatic nitrogen compounds, aromatic sulfur compounds, phenolic compounds and other aromatic oxygen compounds, and their alkyl derivatives.] |  |  | N
208-169-4 | 513-79-1 | cobalt carbonate |  |  | N | Process chemical
200-755-8 | 71-48-7, 6147-53-1 | cobalt di(acetate) |  |  | 1 | Paint and varnish dryer, anodised aluminium sealing.
231-589-4 | 7646-79-9, 7791-13-1 | cobalt dichloride |  |  | 3 | Moisture indicator cards, uncommon ink drier
268-169-5 | 68016-03-5 | cobalt dimolybdenum nickel octaoxide |  |  | 1 | Pigment
233-402-1 | 10026-22-9, 10141-05-6 | cobalt dinitrate |  |  | N | Process chemical
442-750-5 |  | cobalt lithium nickel oxide |  |  | N | Li-ion batteries (no patient contact)_
261-346-8 | 58591-45-0 | cobalt nickel dioxide |  |  | 2 | Reported to be used in supercapacitors
269-051-6 | 68186-89-0 | cobalt nickel gray periclase: C.I. Pigment Black 25: C.I. 77332 |  |  | 1 | Pigment
620-395-9 | 12737-30-3 | cobalt nickel oxide |  |  | 2 | Reported to be used in supercapacitors
233-334-2 | 10026-24-1, 10124-43-3 | cobalt sulfate |  |  | N | Process chemical (reported to be used in storage batteries (so no patient contact)
200-598-5 | 64-86-8 | colchicine |  |  | N | Medicinal drug
227-424-0 | 5836-29-3 | coumatetralyl (ISO): 4-hydroxy-3-(1,2,3,4-tetrahydro-1-naphthyl)coumarin |  |  | N | rat poison
292-606-9 | 90640-85-0 | Creosote oil, acenaphthene fraction, acenaphthene-free: Wash Oil Redistillate: [The oil remaining after removal by a crystallization process of acenaphthene from acenaphthene oil from coal tar. Composed primarily of naphthalene and alkylnaphthalenes.] |  |  | N
292-605-3 | 90640-84-9 | Creosote oil, acenaphthene fraction: Wash Oil: [A complex combination of hydrocarbons produced by the distillation of coal tar and boiling in the range of approximately 240°C to 280°C (464°F to 536°F). Composed primarily of acenaphthene, naphthalene and alkyl naphthalene.] |  |  | N
274-565-9 | 70321-79-8 | Creosote oil, high-boiling distillate: Wash Oil: [The high-boiling distillation fraction obtained from the high temperature carbonization of bituminous coal which is further refined to remove excess crystalline salts. It consists primarily of creosote oil with some of the normal polynuclear aromatic salts, which are components of coal tar distillates, removed. It is crystal free at approximately 5°C (41°F).] |  |  | N
274-566-4 | 70321-80-1 | Creosote oil, low-boiling distillate: Wash Oil: [The low-boiling distillation fraction obtained from the high temperature carbonization of bituminous coal, which is further refined to remove excess crystalline salts. It consists primarily of creosote oil with some of the normal polynuclear aromatic salts, which are components of coal tar distillate, removed. It is crystal free at approximately 38°C (100°F).] |  |  | N
263-047-8 | 61789-28-4 | Creosote oil: [A complex combination of hydrocarbons obtained by the distillation of coal tar. It consists primarily of aromatic hydrocarbons and may contain appreciable quantities of tar acids and tar bases. It distills at the approximate range of 200°C to 325°C (392°F to 617°F).] |  |  | N
232-287-5 | 8001-58-9 | Creosote: [The distillate of coal tar produced by the high temperature carbonization of bituminous coal. It consists primarily of aromatic hydrocarbons, tar acids and tar bases.] |  |  | N
427-230-8 | 5571-36-8 | cyclic 3-(1,2-ethanediylacetale)-estra-5(10),9(11)-diene-3,17-dione |  |  | N | Medical and research
200-636-0 | 66-81-9 | cycloheximide (ISO): 4-{}{(2<i>R</i>)-2-[(1<i>S</i>,3<i>S</i>,5<i>S</i>)-3,5-dimethyl-2-oxocyclohexyl]-2-hydroxyethyl}}piperidine-2,6-dione |  |  | N | Medical research
619-020-1 | 94361-06-5 | cyproconazole (ISO): (2RS,3RS;2RS,3SR)-2-(4-chlorophenyl)-3-cyclopropyl-1-(1H-1,2,4-triazol-1-yl)butan-2-ol |  |  | N | Agricultural fungicide
| 50-29-3 | DDT (technical) (Clofenotane) |  |  | N | Pesticide
234-454-8 | 12004-35-2 | dialuminium nickel tetraoxide |  |  | N | Catalyst
239-793-5 | 15699-18-0 | diammonium nickel bis(sulfate) |  |  | N | Process chemical
| 74195-78-1 | diammonium nickel hexacyanoferrate |  |  | N | Electrochemical sesnors, ion exchange media?
215-116-9 | 1303-28-2, 12044-50-7 | diarsenic pentaoxide: arsenic pentoxide: arsenic oxide: arsenic acid and its salts with the exception of those specified elsewhere in this Annex |  |  | N | SVHC
215-481-4 | 1327-53-3, 7440-38-2 | diarsenic trioxide: arsenic trioxide |  |  | N | SVHC
206-382-7 | 334-88-3 | diazomethane |  |  | N | Laboratory reagent
200-181-8 | 53-70-3 | dibenz[<i>a</i>,<i>h</i>]anthracene |  |  | N | PAH
215-125-8 | 1303-86-2 | diboron trioxide: boric oxide |  |  | N | Process chemical
201-557-4 | 84-74-2, 93952-11-5 | dibutyl phthalate: DBP |  |  | 1 | Plasticiser
211-670-0 | 683-18-1 | dibutyltin dichloride: (DBTC) |  |  | 1 | PVC stabiliser and polyurethane foam catalyst . Has also been an additive for rubbers.
201-039-8 | 77-58-7 | dibutyltin dilaurate: dibutyl[bis(dodecanoyloxy)]stannane |  |  | 1 | Dibutyltin dilaurate is used as an accelerator for solvent based, chemical cross-linking, two-component coatings like polyurethane systems. PVC stabiliser
401-040-5 | 75113-37-0 | dibutyltin hydrogen borate |  |  | N | Biocide? Not likely
246-356-2 | 24613-89-6 | dichromium tris(chromate): chromium III chromate: chromic chromate: Chromium (VI) compounds, with the exception of barium chromate and of compounds specified elsewhere in this Annex |  |  | N | Process chemical
274-460-8 | 70225-14-8 | diethanolamine perfluorooctane sulfonate |  |  | N | Insecticide
200-589-6 | 64-67-5 | diethyl sulphate |  |  | N | Solvent & process chemical
259-978-4 | 56073-07-5 | difenacoum (ISO): 3-(3-biphenyl-4-yl-1,2,3,4-tetrahydro-1-naphthyl)-4-hydroxycoumarin |  |  | N | Rat poison
600-594-7 | 104653-34-1 | difethialone (ISO); 3-[3-(4′-bromobiphenyl-4-yl)-1,2,3,4-tetrahydronaphthalen-1-yl]-4-hydroxy-2H-1-benzothiopyran-2-one |  |  | N | Rat poison
201-559-5 | 84-75-3 | dihexyl phthalate |  |  | 1 | Plasticiser
201-553-2 | 84-69-5 | diisobutyl phthalate |  |  | 1 | Plasticiser
210-088-4 | 605-50-5 | diisopentylphthalate |  |  | 1 | Plasticiser
201-058-1 | 77-78-1 | dimethyl sulphate |  |  | N | Solvent & process chemical
201-208-6 | 79-44-7 | dimethylcarbamoyl chloride |  |  | N | Process chemical
301-323-2 | 93983-68-7 | dimethylhexanoic acid nickel salt |  |  | N | Process chemical
200-549-8 | 62-75-9 | dimethylnitrosoamine: <i>N</i>-nitrosodimethylamine |  |  | N | Process chemical
236-412-4 | 13360-57-1 | dimethylsulfamoylchloride |  |  | N | Process chemical
234-494-6 | 12007-01-1 | dinickel boride |  |  | 1 | Process catalyst and in hard coatings
238-426-6 | 14448-18-1 | dinickel diphosphate |  |  | N | Process chemical
238-946-3 | 14874-78-3 | dinickel hexacyanoferrate |  |  | N | R&D into electrodes
237-411-1 | 13775-54-7 | dinickel orthosilicate |  |  | N | Mineral form
234-828-0 | 12035-64-2 | dinickel phosphide |  |  | 1 | Semiconductor used in laser diodes. Occurs in electroless nickel plating (if phosphorous reducing agent is used)
235-033-1 | 12059-14-2 | dinickel silicide |  |  | N | R&D into semiconductors
215-217-8 | 1314-06-3 | dinickel trioxide |  |  | 1 | Uncommonly used as a pigment in ceramics and glass. Also used as a process chemical
246-836-1 | 25321-14-6 | dinitrotoluene |  |  | N | Process chemical
254-408-0 | 39300-45-3 | dinocap (ISO): (<i>RS</i>)-2,6-dinitro-4-octylphenyl crotonates and (<i>RS</i>)-2,4-dinitro-6-octylphenyl crotonates in which “octyl” is a reaction mass of 1-methylheptyl, 1-ethylhexyl and 1-propylpentyl groups |  |  | N | Biocide
201-861-7 | 88-85-7 | dinoseb (ISO): 6-<i>sec</i>-butyl-2,4-dinitrophenol: salts and esters of dinoseb, with the exception of those specified elsewhere in this Annex |  |  | N | Biocide
215-813-8 | 1420-07-1 | dinoterb (ISO): 2-<i>tert</i>-butyl-4,6-dinitrophenol: salts and esters of dinoterb |  |  | N | Biocide
205-017-9 | 131-18-0 | di-n-pentyl phthalate |  |  | 1 | Plasticiser
251-087-9 | 32536-52-0 | diphenylether; octabromo derivate |  |  | 1 | Octa-BDE FR banned by RoHS, POPs and by REACH
| 19372-20-4 | diphosphoric acid, nickel(II) salt |  |  | N | Process chemical
240-221-1 | 16071-86-6 | disodium {}{5-[(4'-((2,6-hydroxy-3-((2-hydroxy-5-sulphophenyl)azo)phenyl)azo)(1,1'-biphenyl)-4-yl)azo]salicylato(4-)}}cuprate(2-): CI Direct Brown 95 |  |  | 1 | Pigment
209-358-4 | 573-58-0 | disodium 3,3'-[[1,1'-biphenyl]-4,4'-diylbis(azo)]bis(4-aminonaphthalene-1-sulphonate): C.I. Direct Red 28: <i>o</i>-dianisidine based azo dyes: 4,4'-diarylazo-3,3'-dimethoxybiphenyl dyes with the exception of those mentioned elsewhere in this Annex: <i>o</i>-tolidine based dyes: 4,4'-diarylazo-3,3'-dimethylbiphenyl dyes, with the exception of those mentioned elsewhere in this Annex |  |  | 1 | Pigment
217-710-3 | 1937-37-7 | disodium 4-amino-3-[[4'-[(2,4-diaminophenyl)azo][1,1'-biphenyl]-4-yl]azo]-5-hydroxy-6-(phenylazo)naphtalene-2,7-disulphonate: C.I. Direct Black 38 |  |  | 1 | Pigment
234-541-0 | 12008-41-2, 12280-03-4 | disodium octaborate anhydrous: disodium octaborate tetrahydrate |  |  | 1 | Flame retardant used in paper, card, paint, wood and also added to cellulose insulation
215-540-4 | 1303-96-4, 1330-43-4, 12179-04-3 | disodium tetraborate decahydrate: borax decahydrate |  |  | 1 | Flame retardant used in paper, card, paint, wood and also added to cellulose insulation
309-984-9 | 101896-26-8 | Distillates (coal tar), benzole fraction, BTX-rich: Light Oil Redistillate, low boiling: [A residue from the distillation of crude benzole to remove benzole fronts. Composed primarily of benzene, toluene and xylenes boiling in the range of approximately 75°C to 200°C (167°F to 392°F).] |  |  | N
310-165-3 | 121620-46-0 | Distillates (coal tar), benzole fraction, distn. residues: Wash Oil: [A complex combination of hydrocarbons obtained from the distillation of crude benzole (high temperature coal tar). It may be a liquid with the approximate distillation range of 150°C to 300°C (302°F to 572°F) or a semi-solid or solid with a melting point up to 70°C (158°F). It is composed primarily of naphthalene and alkyl naphthalenes.] |  |  | N
283-482-7 | 84650-02-2 | Distillates (coal tar), benzole fraction: Light Oil: [A complex combination of hydrocarbons obtained by the distillation of coal tar. It consists of hydrocarbons having carbon numbers primarily in the range of C<sub>4</sub> to C<sub>10</sub> and distilling in the approximate range of 80 °C to 160 °C (175 °F to 320 °F).] |  |  | N
295-304-5 | 91995-42-5 | Distillates (coal tar), heavy oils, pyrene fraction: Heavy Anthracene Oil Redistillate: [The redistillate obtained from the fractional distillation of pitch distillate boiling in the range of approximately 350 °C to 400 °C (662 °F to 752 °F). Consists predominantly of tri- and polynuclear aromatics and heterocyclic hydrocarbons.] |  |  | N
292-607-4 | 90640-86-1 | Distillates (coal tar), heavy oils: Heavy Anthracene Oil: [Distillate from the fractional distillation of coal tar of bituminous coal, with boiling range of 240 °C to 400 °C (464 °F to 752 °F). Composed primarily of tri- and polynuclear hydrocarbons and heterocyclic compounds.] |  |  | N
292-609-5 | 90640-87-2 | Distillates (coal tar), light oils, acid exts.; Light Oil Extract Residues, high boiling: [This oil is a complex reaction mass of aromatic hydrocarbons, primarily indene, naphthalene, coumarone, phenol, and <i>o</i>-, <i>m</i>- and <i>p</i>-cresol and boiling in the range of 140°C to 215°C (284°F to 419°F).] |  |  | N
292-610-0 | 90640-88-3 | Distillates (coal tar), light oils, alk. exts.: Alkaline Extract: [The aqueous extract from carbolic oil produced by an alkaline wash such as aqueous sodium hydroxide. Composed primarily of the alkali salts of various phenolic compounds.] |  |  | N
309-971-8 | 101794-90-5 | Distillates (coal tar), light oils, neutral fraction; Light Oil Extract Residues, high boiling: [A distillate from the fractional distillation of high temperature coal tar. Composed primarily of alkyl-substituted one ring aromatic hydrocarbons boiling in the range of approximately 135°C to 210°C (275°F to 410°F). May also include unsaturated hydrocarbons such as indene and coumarone.] |  |  | N
283-483-2 | 84650-03-3 | Distillates (coal tar), light oils; Carbolic Oil: [A complex combination of hydrocarbons obtained by distillation of coal tar. It consists of aromatic and other hydrocarbons, phenolic compounds and aromatic nitrogen compounds and distills at the approximate range of 150°C to 210°C (302°F to 410°F).] |  |  | N
295-310-8 | 91995-49-2 | Distillates (coal tar), naphthalene oil crystn. mother liquor: Naphthalene Oil Redistillate: [A complex combination of organic compounds obtained as a filtrate from the crystallization of the naphthalene fraction from coal tar and boiling in the range of approximately 200°C to 230°C (392°F to 446°F). Contains chiefly naphthalene, thionaphthene and alkylnaphthalenes.] |  |  | N
295-309-2 | 91995-48-1 | Distillates (coal tar), naphthalene oils, acid exts.: Methylnaphthalene Oil Extract Residue: [A complex combination of hydrocarbons obtained by debasing the methylnaphthalene fraction obtained by the distillation of coal tar and boiling in the range of approximately 230°C to 255°C (446°F to 491°F). Contains chiefly 1(2)-methylnaphthalene, naphthalene, dimethylnaphthalene and biphenyl.] |  |  | N
292-611-6 | 90640-89-4 | Distillates (coal tar), naphthalene oils, alk. exts.: Alkaline Extract: [The aqueous extract from naphthalene oil produced by an alkaline wash such as aqueous sodium hydroxide. Composed primarily of the alkali salts of various phenolic compounds.] |  |  | N
309-972-3 | 101794-91-6 | Distillates (coal tar), naphthalene oils, indole-methylnaphthalene fraction: Methylnaphthalene Oil: [A distillate from the fractional distillation of high temperature coal tar. Composed primarily of indole and methylnaphthalene boiling in the range of approximately 235°C to 255°C (455°F to 491°F).] |  |  | N
309-985-4 | 101896-27-9 | Distillates (coal tar), naphthalene oils, methylnaphthalene fraction: Methylnaphthalene Oil: [A distillate from the fractional distillation of high temperature coal tar. Composed primarily of substituted two ring aromatic hydrocarbons and aromatic nitrogen bases boiling in the range of approximately 225°C to 255°C (437°F to 491°F).] |  |  | N
292-612-1 | 90640-90-7 | Distillates (coal tar), naphthalene oils, naphthalene-free, alk. exts.: Naphthalene Oil Extract Residue: [The oil remaining after the removal of phenolic compounds (tar acids) from drained naphthalene oil by an alkali wash. Composed primarily of naphthalene and alkyl naphthalenes.] |  |  | N
284-898-1 | 84989-09-3 | Distillates (coal tar), naphthalene oils, naphthalene-low: Naphthalene Oil Redistillate: [A complex combination of hydrocarbons obtained by crystallization of naphthalene oil. Composed primarily of naphthalene, alkyl naphthalenes and phenolic compounds.] |  |  | N
283-484-8 | 84650-04-4 | Distillates (coal tar), naphthalene oils: Naphthalene Oil: [A complex combination of hydrocarbons obtained by the distillation of coal tar. It consists primarily of aromatic and other hydrocarbons, phenolic compounds and aromatic nitrogen compounds and distills in the approximate range of 200°C to 250°C (392°F to 482°F).] |  |  | N
295-312-9 | 91995-51-6 | Distillates (coal tar), pitch, heavy oils: Heavy Anthracene Oil: [The distillate from the distillation of the pitch obtained from bituminous high temperature tar. Composed primarily of tri- and polynuclear aromatic hydrocarbons and boiling in the range of approximately 300 °C to 470 °C (572 °F to 878 °F). The product may also contain heteroatoms.] |  |  | N
295-313-4 | 91995-52-7 | Distillates (coal tar), pitch, pyrene fraction: Heavy Anthracene Oil Redistillate: [The redistillate obtained from the fractional distillation of pitch distillate and boiling in the range of approximately 380 °C to 410 °C (7160 to 770 °F). Composed primarily of tri- and polynuclear aromatic hydrocarbons and heterocyclic compounds.] |  |  | N
309-855-7 | 101316-49-8 | Distillates (coal tar), pitch: Heavy Anthracene Oil: [The oil obtained from condensation of the vapors from the heat treatment of pitch. Composed primarily of two- to four-ring aromatic compounds boiling in the range of 200 °C to greater than 400 °C (392 °F to greater than 752 °F).] |  |  | N
284-899-7 | 84989-10-6 | Distillates (coal tar), upper, fluorene-free: Wash Oil Redistillate: [A complex combination of hydrocarbons obtained by the crystallization of tar oil. It consists of aromatic polycyclic hydrocarbons, primarily diphenyl, dibenzofuran and acenaphthene.] |  |  | N
284-900-0 | 84989-11-7 | Distillates (coal tar), upper, fluorene-rich: Wash Oil Redistillate: [A complex combination of hydrocarbons obtained by the crystallization of tar oil. It consists af aromatic and polycyclic hydrocarbons primarily fluorene and some acenaphthene.] |  |  | N
266-026-1 | 65996-91-0 | Distillates (coal tar), upper: Heavy Anthracene Oil: [The distillate from coal tar having an approximate distillation range of 220 °C to 450 °C (428 °F to 842 °F). Composed primarily of three to four membered condensed ring aromatic hydrocarbons and other hydrocarbons.] |  |  | N
266-027-7 | 65996-92-1 | Distillates (coal tar): Heavy Anthracene Oil: [The distillate from coal tar having an approximate distillation range of 100 °C to 450 °C (212 °F to 842 °F). Composed primarily of two to four membered condensed ring aromatic hydrocarbons, phenolic compounds, and aromatic nitrogen bases.] |  |  | N
295-295-8 | 91995-35-6 | Distillates (coal), coal tar-residual pyrolysis oils, naphthalene oils: Redistillates: [The redistillate obtained from the fractional distillation of bituminous coal high temperature tar and pyrolysis residual oils and boiling in the range of approximately 190°C to 270°C (374°F to 518°F). Composed primarily of substituted dinuclear aromatics.] |  |  | N
285-076-5 | 85029-51-2 | Distillates (coal), coke-oven light oil, naphthalene cut: Naphthalene Oil: [The complex combination of hydrocarbons obtained from prefractionation (continuous distillation) of coke oven light oil. It consists predominantly of naphthalene, coumarone and indene and boils above 148°C (298°F).] |  |  | N
302-688-0 | 94114-52-0 | Distillates (coal), liq. solvent extn., primary: [The liquid product of condensation of vapors emitted during the digestion of coal in a liquid solvent and boiling in the range of approximately 30°C to 300°C (86°F to 572°F). Composed primarily of partly hydrogenated condensed-ring aromatic hydrocarbons, aromatic compounds containing nitrogen, oxygen and sulfur, and their alkyl derivatives having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>14</sub>.] |  |  | N
302-693-8 | 94114-57-5 | Distillates (coal), solvent extn., hydrocracked hydrogenated middle: [Distillate from the hydrogenation of hydrocracked middle distillate from coal extract or solution produced by the liquid solvent extraction or supercritical gas extraction processes and boiling in the range of approximately 180°C to 280°C (356°F to 536°F). Composed primarily of hydrogenated two- ring carbon compounds and their alkyl derivatives having carbon numbers predominantly in the range of C<sub>9</sub> through C<sub>14</sub>.] |  |  | N
302-692-2 | 94114-56-4 | Distillates (coal), solvent extn., hydrocracked middle: [Distillate obtained from the hydrocracking of coal extract or solution produced by the liquid solvent extraction or supercritical gas extraction processes and boiling in the range of approximately 180°C to 300°C (356°F to 572°F. Composed primarily of two-ring aromatic, hydrogenated aromatic and naphthenic compounds, their alkyl derivatives and alkanes having carbon numbers predominantly in the range of C<sub>9</sub> through C<sub>14</sub>. Nitrogen, sulfur and oxygen-containing compounds are also present.] |  |  | N
302-689-6 | 94114-53-1 | Distillates (coal), solvent extn., hydrocracked: [Distillate obtained by hydrocracking of coal extract or solution produced by the liquid solvent extraction or supercritical gas extraction processes and boiling in the range of approximately 30°C to 300°C (86°F to 572°F). Composed primarily of aromatic, hydrogenated aromatic and naphthenic compounds, their alkyl derivatives and alkanes with carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>14</sub>. Nitrogen, sulfur and oxygen-containing aromatic and hydrogenated aromatic compounds are also present.] |  |  | N
269-159-3 | 68188-48-7 | Distillates (coal-petroleum), condensed-ring arom: Distillates: [The distillate from a mixture of coal and tar and aromatic petroleum streams having an approximate distillation range of 220 °C to 450 °C (428 °F to 842 °F). Composed primarily of 3- to 4-membered condensed ring aromatic hydrocarbons.] |  |  | N
295-294-2 | 91995-34-5 | Distillates (petroleum) catalytic reformer, heavy arom. conc.: Gasoil - unspecified: [A complex combination of hydrocarbons obtained from the distillation of a catalytically reformed petroleum cut. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>10</sub> through C<sub>16</sub> and boiling in the range of approximately 200 °C to 300 °C (392 °F to 572 °F).] |  |  | N
265-117-3 | 64742-18-3 | Distillates (petroleum), acid-treated heavy naphthenic: Unrefined or mildly refined baseoil: [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-119-4 | 64742-20-7 | Distillates (petroleum), acid-treated heavy paraffinic: Unrefined or mildly refined baseoil: [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid process. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil having a viscosity of a least 100 SUS at 100 °F (19cSt at 40 °C).] |  |  | N
265-118-9 | 64742-19-4 | Distillates (petroleum), acid-treated light naphthenic: Unrefined or mildly refined baseoil: [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-121-5 | 64742-21-8 | Distillates (petroleum), acid-treated light paraffinic: Unrefined or mildly refined baseoil: [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil having a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C).] |  |  | N
265-114-7 | 64742-14-9 | Distillates (petroleum), acid-treated light: Gasoil - unspecified: [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>9</sub> through C<sub>16</sub> and boiling in the range of approximately 150 °C to 290 °C (302 °F to 554 °F).] |  |  | N
265-113-1 | 64742-13-8 | Distillates (petroleum), acid-treated middle: Gasoil - unspecified: [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>11</sub> through C<sub>20</sub> and boiling in the range of approximately 205 °C to 345 °C (401 °F to 653 °F).] |  |  | N
295-292-1 | 91995-31-2 | Distillates (petroleum), alkene-alkyne manuf. pyrolysis oil, mixed with high-temp. coal tar, indene fraction: Redistillates: [A complex combination of hydrocarbons obtained as a redistillate from the fractional distillation of bituminous coal high temperature tar and residual oils that are obtained by the pyrolytic production of alkenes and alkynes from petroleum products or natural gas. It consists predominantly of indene and boils in a range of approximately 160°C to 190°C (320°F to 374°F).] |  |  | N
270-725-7 | 68477-34-9 | Distillates (petroleum), C<sub>3-5</sub>, 2-methyl-2-butene-rich: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons from the distillation of hydrocarbons usually ranging in carbon numbers from C<sub>3</sub> through C<sub>5</sub>, predominantly isopentane and 3-methyl-1-butene. It consists of saturated and unsaturated hydrocarbons having carbon numbers in the range of C<sub>3</sub> through C<sub>5</sub>, predominantly 2-methyl-2-butene.] |  |  | N
296-903-4 | 93165-19-6 | Distillates (petroleum), C<sub>6</sub>-rich: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained from the distillation of a petroleum feedstock. It consists predominantly of hydrocarbons having carbon numbers of C<sub>5</sub> through C<sub>7</sub>, rich in C<sub>6</sub>, and boiling in the range of approximately 60°C to 70°C (140°F to 158°F).] |  |  | N
309-862-5 | 101316-56-7 | Distillates (petroleum), C<sub>7-9</sub>, C<sub>8</sub>-rich, hydrodesulfurized dearomatized: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by the distillation of petroleum light fraction, hydrodesulfurized and dearomatized. It consists predominantly of hydrocarbons having carbon numbers in the range of C<sub>7</sub> through C<sub>9</sub>, predominantly C<sub>8</sub> paraffins and cycloparaffins, boiling in the range of approximately 120°C to 130°C (248°F to 266°F).] |  |  | N
270-726-2 | 68477-35-0 | Distillates (petroleum), C3-6, piperylene-rich: Petroleum gas: [A complex combination of hydrocarbons from the distillation of saturated and unsaturated aliphatic hydrocarbons usually ranging in the carbon numbers C3 through C6. It consists of saturated and unsaturated hydrocarbons having carbon numbers in the range of C3 through C6, predominantly piperylenes.] |  |  | N
309-667-5 | 100683-97-4 | Distillates (petroleum), carbon-treated light paraffinic: Gasoil - unspecified: [A complex combination of hydrocarbons obtained by the treatment of a petroleum oil fraction with activated charcoal for the removal of traces of polar constituents and impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>12</sub> through C<sub>28</sub>.] |  |  | N
270-660-4 | 68475-79-6 | Distillates (petroleum), catalytic reformed depentanizer: Low boiling point cat-reformed naphtha: [A complex combination of hydrocarbons from the distillation of products from a catalytic reforming process. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C<sub>3</sub> through C<sub>6</sub> and boiling in the range of approximately -49°C to 63°C (-57°F to 145°F).] |  |  | N
285-509-8 | 85116-58-1 | Distillates (petroleum), catalytic reformed hydrotreated light, C<sub>8-12</sub> arom. fraction: Low boiling point cat-reformed naphtha: [A complex combination of alkylbenzenes obtained by the catalytic reforming of petroleum naphtha. It consists predominantly of alkylbenzenes having carbon numbers predominantly in the range of C<sub>8</sub> through C<sub>10</sub> and boiling in the range of approximately 160°C to 180°C (320°F to 356°F).] |  |  | N
271-008-1 | 68513-63-3 | Distillates (petroleum), catalytic reformed straight-run naphtha overheads: Low boiling point cat-reformed naphtha: [A complex combination of hydrocarbons obtained by the catalytic reforming of straight-run naphtha followed by the fractionation of the total effluent. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C<sub>2</sub> through C<sub>6</sub>.] |  |  | N
270-719-4 | 68477-29-2 | Distillates (petroleum), catalytic reformer fractionator residue, high-boiling: Gasoil - unspecified; [A complex combination of hydrocarbons from the distillation of catalytic reformer fracftionator residue. It boils in the range of approximately 343 °C to 399 °C (650 °F to 750 °F).] |  |  | N
270-721-5 | 68477-30-5 | Distillates (petroleum), catalytic reformer fractionator residue, intermediate-boiling: Gasoil - unspecified: [A complex combination of hydrocarbons from the distillation of catalytic reformer fractionator residue. It boils in the range of approximately 288 °C to 371 °C (550 °F to 700 °F).] |  |  | N
270-722-0 | 68477-31-6 | Distillates (petroleum), catalytic reformer fractionator residue, low-boiling: Gasoil - unspecified: [The complex combination of hydrocarbons from the distillation of catalytic reformer fractionator residue. It boils approximately below 288 °C (550 °F).] |  |  | N
265-135-1 | 64742-34-3 | Distillates (petroleum), chemically neutralized heavy naphthenic: Unrefined or mildly refined baseoil: [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-127-8 | 64742-27-4 | Distillates (petroleum), chemically neutralized heavy paraffinic: Unrefined or mildly refined baseoil: [A complex combination of hydrocarbons obtained from a treating process to remove acidic materials. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of aliphatic hydrocarbons.] |  |  | N
265-136-7 | 64742-35-4 | Distillates (petroleum), chemically neutralized light naphthenic: Unrefined or mildly refined baseoil: [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS a 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-128-3 | 64742-28-5 | Distillates (petroleum), chemically neutralized light paraffinic: Unrefined or mildly refined baseoil: [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity less than 100 SUS at 100 °F (19cSt at 40 °C).] |  |  | N
265-130-4 | 64742-30-9 | Distillates (petroleum), chemically neutralized middle: Gasoil - unspecified: [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>11</sub> through C<sub>20</sub> and boiling in the range of approximately 205 °C to 345 °C (401 °F to 653 °F).] |  |  | N
265-146-1 | 64742-44-5 | Distillates (petroleum), clay-treated heavy naphthenic: Baseoil - unspecified: [A complex combination of hydrocarbons resulting from treatment of a petroleum fraction with natural or modified clay in either a contacting or percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-147-7 | 64742-45-6 | Distillates (petroleum), clay-treated light naphthenic: Baseoil - unspecified: [A complex combination of hydrocarbons resulting from treatment of a petroleum fraction with natural or modified clay in either a contacting or percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-138-8 | 64742-37-6 | Distillates (petroleum), clay-treated light paraffinic: Baseoil - unspecified: [A complex combination of hydrocarbons resulting from treatment of a petroleum fraction with natural or modified clay in either a contacting or percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of saturated hydrocarbons.] |  |  | N
265-139-3 | 64742-38-7 | Distillates (petroleum), clay-treated middle: Gasoil - unspecified: [A complex combination of hydrocarbons resulting from treatment of a petroleum fraction with natural or modified clay, usually in a percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>9</sub> through C<sub>20</sub> and boiling in the range of approximately 150 °C to 345 °C (302 °F to 653 °F).] |  |  | N
265-137-2 | 64742-36-5 | Distillates (petroleum), clay-treated paraffinic: Baseoil - unspecified: [A complex combination of hydrocarbons resulting from treatment of a petroleum fraction with natural or modified clay in either a contacting or percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of saturated hydrocarbons.] |  |  | N
292-613-7 | 90640-91-8 | Distillates (petroleum), complex dewaxed heavy paraffinci: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by dewaxing heavy paraffinic distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of equal to or greater than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
292-614-2 | 90640-92-9 | Distillates (petroleum), complex dewaxed light paraffinic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by dewaxing light paraffinic distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>12</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
270-727-8 | 68477-38-3 | Distillates (petroleum), cracked steam-cracked petroleum distillates: Cracked gasoil: [A complex combination of hydrocarbons produced by distilling cracked steam cracked distillate and/or its fractionation products. It consists of hydrocarbons having carbon numbers predominently in the range of C<sub>10</sub> to low molecular weight polymers.] |  |  | N
270-771-8 | 68477-89-4 | Distillates (petroleum), depentanizer overheads: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained from a catalytic cracked gas stream. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>6</sub>.] |  |  | N
295-300-3 | 91995-39-0 | Distillates (petroleum), dewaxed heavy paraffinic, hydrotreated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained from an intensive treatment of dewaxed distillate by hydrogenation in the presence of a catalyst. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>25</sub> through C<sub>39</sub> and produces a finished oil with a viscosity of approximately 44 cSt at 50 °C.] |  |  | N
295-301-9 | 91995-40-3 | Distillates (petroleum), dewaxed light paraffinic, hydrotreated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained from an intensive treatment of dewaxed distillate by hydrogenation in the presence of a catalyst. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>21</sub> through C<sub>29</sub> and produces a finished oil with a viscosity of approximately 13 cSt at 50 °C.] |  |  | N
295-302-4 | 91995-41-4 | Distillates (petroleum), heat-soaked steam-cracked naphtha, C<sub>5</sub>-rich: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by distillation of heat-soaked steam-cracked naphtha. It consists predominantly of hydrocarbons having carbon numbers in the range of C<sub>4</sub> through C<sub>6</sub>, predominantly C<sub>5</sub>.] |  |  | N
267-563-4 | 67891-79-6 | Distillates (petroleum), heavy arom.: Low boiling point thermally cracked naphtha: [The complex combination of hydrocarbons from the distillation of the products from the thermal cracking of ethane and propane. This higher boiling fraction consists predominantly of C<sub>5-7</sub> aromatic hydrocarbons with some unsaturated aliphatic hydrocarbons having carbon number predominantly of C<sub>5</sub>. This stream may contain benzene.] |  |  | N
265-063-0 | 64741-61-3 | Distillates (petroleum), heavy catalytic cracked: Heavy Fuel oil: [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>35</sub> and boiling in the range of approximately 260 °C to 500 °C (500 °F to 932 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
265-077-7 | 64741-76-0 | Distillates (petroleum), heavy hydrocracked: Baseoil - unspecified: [A complex combination of hydrocarbons from the distillation of the products from a hydrocracking process. It consists predominantly of saturated hydrocarbons having carbon numbers in the range of C<sub>15</sub>-C<sub>39</sub> and boiling in the range of approximately 260 °C to 600 °C (500 °F to 1112 °F).] |  |  | N
265-054-1 | 64741-53-3 | Distillates (petroleum), heavy naphthenic: Unrefined or mildly refined baseoil: [A complex combination of hydrocarbons produced by vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-052-0 | 64741-51-1 | Distillates (petroleum), heavy paraffinic: Unrefined or mildly refined baseoil: [A complex combination of hydrocarbons produced by vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of saturated aliphatic hydrocarbons.] |  |  | N
309-939-3 | 101631-14-5 | Distillates (petroleum), heavy steam-cracked: Cracked gasoil: [A complex combination of hydrocarbons obtained by distillation of steam cracking heavy residues. It consists predominantly of highly alkylated heavy aromatic hydrocarbons boiling in the range of approximately 250 °C to 400 °C (482 °F to 752 °F).] |  |  | N
265-082-4 | 64741-81-7 | Distillates (petroleum), heavy thermal cracked: Heavy Fuel oil: [A complex combination of hydrocarbons from the distillation of the products from a thermal cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>36</sub> and boiling in the range of approximately 260 °C to 480 °C (500 °F to 896 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
292-615-8 | 90640-93-0 | Distillates (petroleum), highly refined middle: Gasoil - unspecified: [A complex combination of hydrocarbons obtained by the subjection of a petroleum fraction to several of the following steps: filtration, centrifugation, atmospheric distillation, vacuum distillation, acidification, neutralization and clay treatment. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>10</sub> through C<sub>20</sub>.] |  |  | N
307-010-7 | 97488-73-8 | Distillates (petroleum), hydrocracked solvent-refined light: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by the solvent treatment of a distillate from hydrocracked petroleum distillates. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>18</sub> through C<sub>27</sub> and boiling in the range of approximately 370 °C to 450 °C (698 °F to 842 °F.] |  |  | N
295-306-6 | 91995-45-8 | Distillates (petroleum), hydrocracked solvent-refined, dewaxed: Baseoil - unspecified: [A complex combination of liquid hydrocarbons obtained by recrystallization of dewaxed hydrocracked solvent-refined petroleum distillates.] |  |  | N
309-863-0 | 101316-57-8 | Distillates (petroleum), hydrodesulfurized full-range middle: Heavy Fuel oil: [A complex combination of hydrocarbons obtained by treating a petroleum stock with hydrogen. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>9</sub> through C<sub>25</sub> and boiling in the range of approximately 150 °C to 400 °C (302 °F to 752 °F).] |  |  | N
269-784-1 | 68333-28-8 | Distillates (petroleum), hydrodesulfurized heavy catalytic cracked: Heavy Fuel oil: [A complex combination of hydrocarbons obtained by treatment of heavy catalytic cracked distillates with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>35</sub> and boiling in the range of approximately 260 °C to 500 °C (500 °F to 932 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
269-783-6 | 68333-27-7 | Distillates (petroleum), hydrodesulfurized intermediate catalytic cracked: Heavy Fuel oil: [A complex combination of hydrocarbons obtained by treating intermediate catalytic cracked distillates with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>11</sub> through C<sub>30</sub> and boiling in the range of approximately 205 °C to 450 °C (401 °F to 842 °F). It contains a relatively large proportion of tricyclic aromatic hydrocarbons.] |  |  | N
269-781-5 | 68333-25-5 | Distillates (petroleum), hydrodesulfurized light catalytic cracked: Cracked gasoil: [A complex combination of hydrocarbons obtained by treating light catalytic cracked distillates with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>9</sub> through C<sub>25</sub> and boiling in the range of approximately 150 °C to 400 °C (302 °F to 752 °F). It contains a relatively large proportion of bicyclic aromatic hydrocarbons.] |  |  | N
309-865-1 | 101316-59-0 | Distillates (petroleum), hydrodesulfurized middle coker: Cracked gasoil: [A complex combination of hydrocarbons by fractionation from hydrodesulfurised coker distillate stocks. Is consists of hydro-carbons having carbon numbers predominantly in the range of C<sub>12</sub> through C<sub>21</sub> and boiling in the range of approximately 200 °C to 360 °C (392 °F to 680 °F).] |  |  | N
265-183-3 | 64742-80-9 | Distillates (petroleum), hydrodesulfurized middle: Gasoil - unspecified: [A complex combination of hydrocarbons obtained from a petroleum stock by treating with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>11</sub> through C<sub>25</sub> and boiling in the range of approximately 205 °C to 400 °C (401 °F to 752 °F).] |  |  | N
285-505-6 | 85116-53-6 | Distillates (petroleum), hydrodesulfurized thermal cracked middle: Cracked gasoil: [A complex combination of hydrocarbons obtained by fractionation from hydrodesulfurized themal cracker distillate stocks. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>11</sub> to C<sub>25</sub> and boiling in the range of approximately 205 °C to 400 °C (401 °F to 752 °F).] |  |  | N
270-094-8 | 68410-98-0 | Distillates (petroleum), hydrotreated heavy naphtha, deisohexanizer overheads: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained by distillation of the products from a heavy naphtha hydrotreating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>3</sub> through C<sub>6</sub> and boiling in the range of approximately -49°C to 68°C (-57°F to 155°F).] |  |  | N
265-155-0 | 64742-52-5 | Distillates (petroleum), hydrotreated heavy naphthenic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-157-1 | 64742-54-7 | Distillates (petroleum), hydrotreated heavy paraffinic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of saturated hydrocarbons.] |  |  | N
265-156-6 | 64742-53-6 | Distillates (petroleum), hydrotreated light naphthenic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-158-7 | 64742-55-8 | Distillates (petroleum), hydrotreated light paraffinic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of saturated hydrocarbons.] |  |  | N
270-092-7 | 68410-96-8 | Distillates (petroleum), hydrotreated middle, intermediate boiling: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained by the distillation of products from a middle distillate hydrotreating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>10</sub> and boiling in the range of approximately 127°C to 188°C (262°F to 370°F).] |  |  | N
265-148-2 | 8042-47-5, 64742-46-7, 72623-86-0, 97862-82-3 | Distillates (petroleum), hydrotreated middle: Gasoil - unspecified: [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>11</sub> through C<sub>25</sub> and boiling in the range of approximately 205 °C to 400 °C (401 °F to 752 °F).] |  |  | N
295-990-6 | 92201-59-7 | Distillates (petroleum), intermediate catalytic cracked, thermally degraded: Heavy Fuel oil: [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process which has been used as a heat transfer fluid. It consists predominantly of hydrocarbons boiling in the range of approximately 220 °C to 450 °C (428 °F to 842 °F). This stream is likely to contain organic sulfur compounds.] |  |  | N
265-062-5 | 64741-60-2 | Distillates (petroleum), intermediate catalytic cracked: Cracked gasoil: [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>11</sub> through C<sub>30</sub> and boiling in the range of approximately 205 °C to 450 °C (401 °F to 842 °F). It contains a relatively large proportion of tricyclic aromatic hydrocarbons.] |  |  | N
309-668-0 | 100683-98-5 | Distillates (petroleum), intermediate paraffinic, carbon-treated: Gasoil - unspecified: [A complex combination of hydrocarbons obtained by the treatment of petroleum with activated charcoal for the removal of trace polar constituents and impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>16</sub> through C<sub>36</sub>.] |  |  | N
309-669-6 | 100683-99-6 | Distillates (petroleum), intermediate paraffinic, clay-treated: Gasoil - unspecified: [A complex combination of hydrocarbons obtained by the treatment of petroleum with bleaching earth for the removal of trace polar constituents and impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>16</sub> through C<sub>36</sub>.] |  |  | N
274-683-0 | 70592-76-6 | Distillates (petroleum), intermediate vacuum: Heavy Fuel oil: [A complex combination of hydrocarbons produced by the vacuum, distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>14</sub> through C<sub>42</sub> and boiling in the range of approximately 250 °C to 545 °C (482 °F to 1013 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
267-565-5 | 67891-80-9 | Distillates (petroleum), light arom.: Low boiling point thermally cracked naphtha: [The complex combination of hydrocarbons from the distillation of the products from the thermal cracking of ethane and propane. This lower boiling fraction consists predominantly of C<sub>5-7</sub> aromatic hydrocarbons with some unsaturated aliphatic hydrocarbons having a carbon number predominantly of C<sub>5</sub>. This stream may contain benzene.] |  |  | N
295-991-1 | 92201-60-0 | Distillates (petroleum), light catalytic cracked, thermally degraded: Cracked gasoil: [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process which has been used as a heat transfer fluid. It consists predominantly of hydrocarbons boiling in the range of approximately 190 °C to 340 °C (374 °F to 644 °F). This stream is likely to contain organic sulfur compounds.] |  |  | N
265-060-4 | 64741-59-9 | Distillates (petroleum), light catalytic cracked: Cracked gasoil: [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>9</sub> through C<sub>25</sub> and boiling in the range of approximately 150 °C to 400 °C (302 °F to 752 °F). It contains a relatively large proportion of bicyclic aromatic hydrocarbons.] |  |  | N
270-093-2 | 68410-97-9 | Distillates (petroleum), light distillate hydrotreating process, low-boiling: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained by the distillation of products from the light distillate hydrotreating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>9</sub> and boiling in the range of approximately 3°C to 194°C (37°F to 382°F).] |  |  | N
265-053-6 | 64741-52-2 | Distillates (petroleum), light naphthenic: Unrefined or mildly refined baseoil: [A complex combination of hydrocarbons produced by vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-051-5 | 64741-50-0 | Distillates (petroleum), light paraffinic: Unrefined or mildly refined baseoil: [A complex combination of hydrocarbons produced by vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains a relatively large proportion of saturated aliphatic hydrocarbons normally present in this distillation range of crude oil.] |  |  | N
270-662-5 | 68475-80-9 | Distillates (petroleum), light steam-cracked naphtha: Cracked gasoil: [A complex combination of hydrocarbons from the multiple distillation of products from a steam cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>10</sub> through C<sub>18</sub>.] |  |  | N
272-931-2 | 68921-08-4 | Distillates (petroleum), light straight-run gasoline fractionation stabilizer overheads: Low boiling point naphtha: [A complex combination of hydrocarbons obtained by the fractionation of light straight-run gasoline. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C<sub>3</sub> through C<sub>6</sub>.] |  |  | N
273-266-0 | 68955-29-3 | Distillates (petroleum), light thermal cracked, debutanized arom.: Low boiling point thermally cracked naphtha: [A complex combination of hydrocarbons produced by the distillation of products from a thermal cracking process. It consists predominantly of aromatic hydrocarbons, primarily benzene.] |  |  | N
265-084-5 | 64741-82-8 | Distillates (petroleum), light thermal cracked: Cracked gasoil: [A complex combination of hydrocarbons from the distillation of the products from a thermal cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C<sub>10</sub> through C<sub>22</sub> and boiling in the range of approximately 160 °C to 370 °C (320 °F to 698 °F).] |  |  | N
274-684-6 | 70592-77-7 | Distillates (petroleum), light vacuum: Heavy Fuel oil: [A complex combination of hydrocarbons produced by the vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>11</sub> through C<sub>35</sub> and boiling in the range of approximately 250 °C to 545 °C (482 °F to 1013 °F).] |  |  | N
295-311-3 | 91995-50-5 | Distillates (petroleum), naphtha steam cracking-derived, hydrotreated light arom.: Low boiling point cat-cracked naphtha.: [A complex combination of hydrocarbons obtained by treating a light distillate from steam-cracked naphtha. It consists predominantly of aromatic hydrocarbons.] |  |  | N
295-315-5 | 91995-53-8 | Distillates (petroleum), naphtha steam cracking-derived, solvent-refined light hydrotreated: Low boiling point modified naphtha: [A complex combination of hydrocarbons obtained as the raffinates from a solvent extraction process of hydrotreated light distillate from steam-cracked naphtha.] |  |  | N
272-932-8 | 68921-09-5 | Distillates (petroleum), naphtha unifiner stripper: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons produced by stripping the products from the naphtha unifiner. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C<sub>2</sub> through C<sub>6</sub>.] |  |  | N
270-344-6 | 68425-29-6 | Distillates (petroleum), naphtha-raffinate pyrolyzate-derived, gasoline-blending: Low boiling point thermally cracked naphtha: [The complex combination of hydrocarbons obtained by the pyrolysis fractionation at 816°C (1500°F) of naphtha and raffinate. It consists predominantly of hydrocarbons having a carbon number of C<sub>9</sub> and boiling at approximately 204°C (400°F).] |  |  | N
273-263-4 | 68955-27-1 | Distillates (petroleum), petroleum residues vacuum: Heavy Fuel oil: [A complex combination of hydrocarbons produced by the vacuum distillation of the residuum from the atmospheric distillation of crude oil.] |  |  | N
270-735-1 | 68477-50-9 | Distillates (petroleum), polymd. steam-cracked petroleum distillates, C<sub>5-12</sub> fraction: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained from the distillation of polymerized steam-cracked petroleum distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>12</sub>.] |  |  | N
292-616-3 | 90640-94-1 | Distillates (petroleum), solvent dewaxed heavy paraffinic, clay-treated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treating dewaxed heavy paraffinic distillate with neutral or modified clay in either a contacting or percolation process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub>.] |  |  | N
292-618-4 | 90640-96-3 | Distillates (petroleum), solvent dewaxed light paraffinic, clay-treated: Baseoil - unspecified: [A complex combination of hydrocarbons resulting from treatment of dewaxed light paraffinic distillate with natural or modified clay in either a contacting or percolation process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub>.] |  |  | N
292-620-5 | 90640-97-4 | Distillates (petroleum), solvent dewaxed light paraffinic, hydrotreated: Baseoil - unspecified: [A complex combination of hydrocarbons produced by treating a dewaxed light paraffinic distillate with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub>.] |  |  | N
265-167-6 | 64742-63-8 | Distillates (petroleum), solvent-dewaxed heavy naphthenic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by removal of normal paraffins from a petroleum fraction by solvent crystallization. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>20 .</sub> through C<sub>50</sub> and produces a finished oil of not less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-169-7 | 64742-65-0 | Distillates (petroleum), solvent-dewaxed heavy paraffinic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by removal of normal paraffins from a petroleum fraction by solvent crystallization. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity not less than 100 SUS at 100 °F (19cSt at 40 °C).] |  |  | N
265-168-1 | 64742-64-9 | Distillates (petroleum), solvent-dewaxed light naphthenic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by removal of normal paraffins from a petroleum fraction by solvent crystallization. It consists of hydrocarbons having carbon numbers predominantly in the range C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-159-2 | 64742-56-9 | Distillates (petroleum), solvent-dewaxed light paraffinic: Baseoil - unspecified: [A complex comination of hydrocarbons obtained by removal of normal paraffins from a petroleum fraction by solvent crystallization. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C).] |  |  | N
265-097-6 | 64741-96-4 | Distillates (petroleum), solvent-refined heavy naphthenic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt a 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-090-8 | 64741-88-4 | Distillates (petroleum), solvent-refined heavy paraffinic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C).] |  |  | N
305-589-0 | 94733-09-2 | Distillates (petroleum), solvent-refined hydrocracked light: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by solvent dearomatization of the residue of hydrocracked petroleum. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>18</sub> through C<sub>27</sub> and boiling in the range of approximately 370 °C to 450 °C (698 °F to 842 °F).] |  |  | N
307-011-2 | 97488-74-9 | Distillates (petroleum), solvent-refined hydrogenated heavy: Baseoil - unspecified: [A complex combination of hydrocarbons, obtained by the treatment of a hydrogenated petroleum distillate with a solvent. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>19</sub> through C<sub>40</sub> and boiling in the range of approximately 390 °C to 550 °C (734 °F to 1022 °F).] |  |  | N
305-588-5 | 94733-08-1 | Distillates (petroleum), solvent-refined hydrotreated heavy, hydrogenated: Baseoil - unspecified |  |  | N
295-316-0 | 91995-54-9 | Distillates (petroleum), solvent-refined light naphthenic, hydrotreated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst and removing the aromatic hydrocarbons by solvent extraction. It consists predominantly of naphthenic hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of between 13-15cSt at 40 °C.] |  |  | N
265-098-1 | 64741-97-5 | Distillates (petroleum), solvent-refined light naphthenic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-091-3 | 64741-89-5 | Distillates (petroleum), solvent-refined light paraffinic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C).] |  |  | N
265-093-4 | 64741-91-9 | Distillates (petroleum), solvent-refined middle: Gasoil - unspecified: [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C<sub>9</sub> through C<sub>20</sub> and boiling in the range of approximately 150 °C to 345 °C (302 °F to 653 °F).] |  |  | N
270-738-8 | 68477-55-4 | Distillates (petroleum), steam-cracked, C<sub>5-10</sub> fraction, mixed with light steam-cracked petroleum naphtha C<sub>5</sub> fraction: Low boiling point naphtha - unspecified |  |  | N
270-736-7 | 68477-53-2 | Distillates (petroleum), steam-cracked, C<sub>5-12</sub> fraction: Low boiling point naphtha - unspecified: [A complex combination of organic compounds obtained by the distillation of products from a steam cracking process. It consists of unsaturated hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>12</sub>.] |  |  | N
305-750-5 | 95009-23-7 | Distillates (petroleum), steam-cracked, C<sub>8-12</sub> fraction, polymd., distn. lights: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by distillation of the polymerized C<sub>8</sub> through C<sub>12</sub> fraction from steam-cracked petroleum distillates. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>8</sub> through C<sub>12</sub>.] |  |  | N
270-077-5 | 68410-05-9 | Distillates (petroleum), straight-run light: Low boiling point naphtha: [A complex combination of hydrocarbons produced by the distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>2</sub> through C<sub>7</sub> and boiling in the range of approximately -88°C to 99°C (-127°F to 210°F).] |  |  | N
265-088-7 | 64741-86-2 | Distillates (petroleum), sweetened middle: Gasoil - unspecified: [A complex combination of hydrocarbons obtained by subjecting a petroleum distillate to a sweetening process to convert mercaptans or to remove acidic impurities. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>9</sub> through C<sub>20</sub> and boiling in the range of approximately 150 °C to 345 °C (302 °F to 653 °F).] |  |  | N
271-632-4 | 68603-01-0 | Distillates (petroleum), thermal cracked naphtha and gas oil, C<sub>5</sub>-dimer-contg.: Low boiling point thermally cracked naphtha: [A complex combination of hydrocarbons produced by the extractive distillation of thermal cracked naphtha and/or gas oil. It consists predominantly of hydrocarbons having a carbon number of C<sub>5</sub> with some dimerized C<sub>5</sub> olefins and boiling in the range of approximately 33°C to 184°C (91°F to 363°F).] |  |  | N
271-634-5 | 68603-03-2 | Distillates (petroleum), thermal cracked naphtha and gas oil, extractive: Low boiling point thermally cracked naphtha: [A complex combination of hydrocarbons produced by the extractive distillation of thermal cracked naphtha and/or gas oil. It consists of paraffinic and olefinic hydrocarbons, predominantly isoamylenes such as 2-methyl-1-butene and 2-methyl-2-butene and boiling in the range of approximately 31°C to 40°C (88°F to 104°F).] |  |  | N
271-631-9 | 68603-00-9 | Distillates (petroleum), thermal cracked naphtha and gas oil: Low boiling point thermally cracked naphtha: [A complex combination of hydrocarbons produced by distillation of thermally cracked naphtha and/or gas oil. It consists predominantly of olefinic hydrocarbons having a carbon number of C<sub>5</sub> and boiling in the range of approximately 33°C to 60°C (91°F to 140°F).] |  |  | N
274-685-1 | 70592-78-8 | Distillates (petroleum), vacuum: Heavy Fuel oil: [A complex combination of hydrocarbons produced by the vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having numbers predominantly in the range of C<sub>15</sub> through C<sub>50</sub> and boiling in the range of approximately 270 °C to 600 °C (518 °F to 1112 °F). This stream is likely to contain 5 wt.% or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
406-850-2 | 106325-08-0, 133855-98-8 | epoxiconazole (ISO): (2RS,3SR)-3-(2-chlorophenyl)-2-(4-fluorophenyl)-[( 1H-1,2,4-triazol-1-yl)methyl]oxirane |  |  | N | Agricultural fungicide
620-440-2 | 12510-42-8 | erionite |  |  | N | Type of zeolite
253-704-7 | 37894-46-5 | etacelasil (ISO): 6-(2-chloroethyl)-6-(2-methoxyethoxy)-2,5,7,10-tetraoxa-6-silaundecane |  |  | N | Pesticide
401-290-5 | 103112-35-2 | ethyl 1-(2,4-dichlorophenyl)-5-(trichloromethyl)-1<i>H</i>-1,2,4-triazole-3-carboxylate |  |  | N | Pharmaceutical
275-897-7 | 71720-48-4 | ethyl hydrogen sulfate, nickel(II) salt |  |  | N | Process chemical
200-849-9 | 75-21-8 | ethylene oxide: oxirane |  |  | N | Process chemical, also used as a steriliser gas (boiling point 10°C)
202-506-9 | 96-45-7 | ethylene thiourea: imidazolidine-2-thione: 2-imidazoline-2-thiol |  |  | N | Used as a vulcanising agent in rubbers but should not be present in vulcanised rubber
205-793-9 | 151-56-4 | ethyleneimine: aziridine |  |  | N | Process chemical
284-901-6 | 84989-12-8 | Extract oils (coal), acidic, tar-base free: Methylnaphthalene Oil Extract Residue: [The extract oil boiling in the range of approximately 220°C to 265°C (428°F to 509°F) from coal tar alkaline extract residue produced by an acidic wash such as aqueous sulfuric acid after distillation to remove tar bases. Composed primarily of alkylnaphthalenes.] |  |  | N
310-171-6 | 122070-80-8 | Extract oils (coal), coal tar residual pyrolysis oils, naphthalene oil, distn. residues: Redistillates: [Residue from the distillation of dephenolated and debased methylnaphthalene oil (from bituminous coal tar and pyrolysis residual oils) with a boiling range of 240°C to 260°C (464°F to 500°F). Composed primarily of substituted dinuclear aromatic and heterocyclic hydrocarbons.] |  |  | N
295-329-1 | 91995-66-3 | Extract oils (coal), coal tar-residual pyrolysis oils, naphthalene oil, redistillate: Redistillates: [The redistillate from the fractional distillation of dephenolated and debased methylnaphthalene oil obtained from bituminous coal high temperature tar and pyrolysis residual oils boiling in the approximate range of 220°C to 230°C (428°F to 446°F). It consists predominantly of unsubstituted and substituted dinuclear aromatic hydrocarbons.] |  |  | N
310-170-0 | 122070-79-5 | Extract oils (coal), coal tar-residual pyrolysis oils, naphthalene oils: Redistillates: [A neutral oil obtained by debasing and dephenolating the oil obtained from the distillation of high temperature tar and pyrolysis residual oils which has a boiliing range of 225°C to 255°C (437°F to 491°F). Composed primarily of substituted dinuclear aromatic hydrocarbons.] |  |  | N
292-622-6 | 90640-99-6 | Extract oils (coal), light oil: Acid Extract: [The aqueous extract produced by an acidic wash of alkali-washed carbolic oil. Composed primarily of acid salts of various aromatic nitrogen bases including pyridine, quinoline and their alkyl derivatives.] |  |  | N
292-623-1 | 90641-00-2 | Extract oils (coal), naphthalene oils: Acid Extract: [The aqueous extract produced by an acidic wash of alkali-washed naphthalene oil. Composed primarily of acid salts of various aromatic nitrogen bases including pyridine, quinoline and their alkyl derivatives.] |  |  | N
273-077-3 | 68937-63-3 | Extract oils (coal), tar base, collidine fraction: Distillate Bases: [The extract produced by the acidic extraction of bases from crude coal tar aromatic oils, neutralization, and distillation of the bases. Composed primarily of collidines, aniline, toluidines, lutidines, xylidines.] |  |  | N
266-020-9 | 65996-86-3 | Extract oils (coal), tar base: Acid Extract: [The extract from coal tar oil alkaline extract residue produced by an acidic wash such as aqueous sulfuric acid after distillation to remove naphthalene. Composed primarily of the acid salts of various aromatic nitrogen bases including pyridine, quinoline, and their alkyl derivatives.] |  |  | N
309-868-8 | 101316-63-6 | Extract residues (coal tar), benzole fraction alk., acid ext.: Light Oil Extract Residues, low boiling: [A complex combination of hydrocarbons obtained by the redistillation of the distillate of high temperature coal tar (tar acid and tar base free). It consists predominantly of unsubstituted and substituted mononuclear aromatic hydrocarbons boiling in the range of 85°C to 195°C (185°F to 383°F).] |  |  | N
298-725-2 | 93821-38-6 | Extract residues (coal), benzole fraction acid: Light Oil Extract Residues, low boiling: [An acid sludge by-product of the sulfuric acid refining of crude high temperature coal. Composed primarily of sulfuric acid and organic compounds.] |  |  | N
295-323-9 | 91995-61-8 | Extract residues (coal), benzole fraction alk., acid ext.: Light Oil Extract Residues, low boiling: [The redistillate from the distillate, freed of tar acids and tar bases, from bituminous coal high temperature tar boiling in the approximate range of 90°C to 160°C (194°F to 320°F). It consists predominantly of benzene, toluene and xylenes.] |  |  | N
294-285-0 | 91697-23-3 | Extract residues (coal), brown: Coal Tar Extract: [The residue from extraction of dried coal.] |  |  | N
310-189-4 | 122384-77-4 | Extract residues (coal), creosote oil acid: Wash Oil Extract Residue: [A complex combination of hydrocarbons from the base-freed fraction from the distillation of coal tar, boiling in the range of approximately 250°C to 280°C (482°F to 536°F). It consists predominantly of biphenyl and isomeric diphenylnaphthalenes.] |  |  | N
309-867-2 | 101316-62-5 | Extract residues (coal), light oil alk., acid ext., indene fraction: Light Oil Extract Residues, intermediate boiling |  |  | N
292-624-7 | 90641-01-3 | Extract residues (coal), light oil alk., acid ext.: Carbolic Oil Extract Residue: [The oil resulting from the acid washing of alkali-washed carbolic oil to remove the minor amounts of basic compounds (tar bases). Composed primarily of indene, indan and alkylbenzenes.] |  |  | N
292-625-2 | 90641-02-4 | Extract residues (coal), light oil alk., distn. overheads: Light Oil Extract Residues, low boiling: [The first fraction from the distillation of aromatic hydrocarbons, coumarone, naphthalene and indene rich prefractionator bottoms or washed carbolic oil boiling substantially below 145°C (293°F). Composed primarily of C<sub>7</sub> and C<sub>8</sub> aliphatic and aromatic hydrocarbons.] |  |  | N
292-626-8 | 90641-03-5 | Extract residues (coal), light oil alk., indene naphtha fraction: Light Oil Extract Residues, high boiling: [The distillate from aromatic hydrocarbons, coumarone, naphthalene and indene rich prefractionator bottoms or washed carbolic oils, having an approximate boiling range of 155°C to 180°C (311°F to 356°F). Composed primarily of indene, indan and trimethylbenzenes.] |  |  | N
310-191-5 | 122384-78-5 | Extract residues (coal), low temp. coal atar alk.: [The residue from low temperature coal tar oils after an alkaline wash, such as aqueous sodium hydroxide, to remove crude coal tar acids. Composed primarily of hydrocarbons and aromatic nitrogen bases.] |  |  | N
292-627-3 | 90641-04-6 | Extract residues (coal), naphthalene oil alk., distn. overheads: Naphthalene Oil Extract Residue: [The distillate from alkali-washed naphthalene oil having an approximate distillation range of 180°C to 220°C (356°F to 428°F). Composed primarily of naphthalene, alkylbenzenes, indene and indan.] |  |  | N
292-628-9 | 90641-05-7 | Extract residues (coal), naphthalene oil alk., distn. residues: Methylnaphthalene Oil Extract Residue: [The residue from the distillation of alkali-washed naphthalene oil having an approximate distillation range of 220°C to 300°C (428°F to 572°F). Composed primarily of naphthalene, alkylnaphthalenes and aromatic nitrogen bases.] |  |  | N
310-167-4 | 121620-48-2 | Extract residues (coal), naphthalene oil, alk., naphthalene-low: Naphthalene Oil Extract Residue: [A complex combination of hydrocarbons remaining after the removal of naphthalene from alkali-washed naphthalene oil by a crystallization process. It is composed primarily of naphthalene and alkyl naphthalenes.] |  |  | N
310-166-9 | 121620-47-1 | Extract residues (coal), naphthalene oil, alk.: Naphthalene Oil Extract Residue: [A complex combination of hydrocarbons obtained from the alkali washing of naphthalene oil to remove phenolic compounds (tar acids). It is composed of naphthalene and alkyl naphthalenes.] |  |  | N
292-629-4 | 90641-06-8 | Extract residues (coal), tar oil alk., carbonated, limed: Crude Phenols: [The product obtained by treatment of coal tar oil alkaline extract with CO<sub>2</sub> and CaO. Composed primarily of CaCO<sub>3</sub>, Ca(OH)<sub>2</sub>, Na<sub>2</sub>CO<sub>3</sub> and other organic and inorganic impurities.] |  |  | N
277-567-8 | 73665-18-6 | Extract residues (coal), tar oil alk., naphthalene distn. residues: Naphthalene Oil Extract Residue: [The residue obtained from chemical oil extracted after the removal of naphthalene by distillation composed primarily of two to four membered condensed ring aromatic hydrocarbons and aromatic nitrogen bases.] |  |  | N
266-021-4 | 65996-87-4 | Extract residues (coal), tar oil alk.: Carbolic Oil Extract Residue: [The residue obtained from coal tar oil by an alkaline wash such as aqueous sodium hydroxide after the removal of crude coal tar acids. Composed primarily of naphthalenes and aromatic nitrogen bases.] |  |  | N
308-261-5 | 97926-43-7 | Extracts (petroleum) heavy naphtha solvent, clay-treated: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by the treatment of heavy naphthic solvent petroleum extract with bleaching earth. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>10</sub> and boiling in the range of approximately 80°C to 180°C (175°F to 356°F).] |  |  | N
295-331-2 | 91995-68-5 | Extracts (petroleum), catalytic reformed light naphtha solvent: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained as the extract from the solvent extraction of a catalytically reformed petroleum cut. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>8</sub> and boiling in the range of approximately 100°C to 200°C (212°F to 392°F).] |  |  | N
270-741-4 | 68477-61-2 | Extracts (petroleum), cold-acid, C<sub>4-6</sub>: Low boiling point naphtha - unspecified: [A complex combination of organic compounds produced by cold acid unit extraction of saturated and unsaturated aliphatic hydrocarbons usually ranging in carbon numbers from C<sub>3</sub> through C<sub>6</sub>, predominantly pentanes and amylenes. It consists predominantly of saturated and unsaturated hydrocarbons having carbon numbers in the range of C<sub>4</sub> through C<sub>6</sub>, predominantly C<sub>5</sub>.] |  |  | N
265-111-0 | 64742-11-6 | Extracts (petroleum), heavy naphthenic distillate solvent |  |  | N
272-175-3 | 68783-00-6 | Extracts (petroleum), heavy naphthenic distillate solvent, arom. conc.: Distillate aromatic extract (treated): [An aromatic concentrate produced by adding water to heavy naphthenic distillate solvent extract and extraction solvent.] |  |  | N
297-827-4 | 93763-10-1 | Extracts (petroleum), heavy naphthenic distillate solvent, hydrodesulfurized: Distillate aromatic extract (treated): [A complex combination of hydrocarbons obtained from a petroleum stock by treating with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of greater than 19cSt at 40 °C.] |  |  | N
292-631-5 | 90641-07-9 | Extracts (petroleum), heavy naphthenic distillate solvent, hydrotreated: Distillate aromatic extract (treated): [A complex combination of hydrocarbons obtained by treating a heavy naphthenic distillate solvent extract with hydrogen in the presence of a catalyst. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil of at least 19cSt at 40 °C (100 SUS at 100 °F).] |  |  | N
265-103-7 | 64742-04-7 | Extracts (petroleum), heavy paraffinic distillate solvent |  |  | N
296-437-1 | 92704-08-0 | Extracts (petroleum), heavy paraffinic distillate solvent, clay-treated: Distillate aromatic extract (treated): [A complex combination of hydrocarbons resulting from treatment of a petroleum fraction with natural or modified clay in either a contact or percolation process to remove the trace amounts of polar compounds and impurities present. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub>. This stream is likely to contain 5 wt.% or more 4-6 membered ring aromatic hydrocarbons.] |  |  | N
292-632-0 | 90641-08-0 | Extracts (petroleum), heavy paraffinic distillate solvent, hydrotreated: Distillate aromatic extract (treated): [A complex combination of hydrocarbons produced by treating a heavy paraffinic distillate solvent extract with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>21</sub> through C<sub>33</sub> and boiling in the range of approximately 350 °C to 480 °C (662 °F to 896 °F). |  |  | N
272-342-0 | 68814-89-1 | Extracts (petroleum), heavy paraffinic distillates, solvent-deasphalted: Distillate aromatic extract (treated): [A complex combination of hydrocarbons obtained as the extract from a solvent extraction of heavy paraffinic distillate.] |  |  | N
295-335-4 | 91995-73-2 | Extracts (petroleum), hydrotreated light paraffinic distillate solvent: Distillate aromatic extract (treated): [A complex combination of hydrocarbons obtained as the extract from solvent extraction of intermediate paraffinic top solvent distillate that is treated with hydrogen in the presence of a catalyst. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>16</sub> through C<sub>36</sub>.] |  |  | N
265-102-1 | 64742-03-6 | Extracts (petroleum), light naphthenic distillate solvent |  |  | N
295-338-0 | 91995-75-4 | Extracts (petroleum), light naphthenic distillate solvent, hydrodesulfurized: Distillate aromatic extract (treated): [A complex combination of hydrocarbons obtained by treating the extract, obtained from a solvent extraction process, with hydrogen in the presence of a catalyst under conditions primarily to remove sulfur compounds. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub>. This stream is likely to contain 5 wt.% or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
265-104-2 | 64742-05-8 | Extracts (petroleum), light paraffinic distillate solvent |  |  | N
295-339-6 | 91995-76-5 | Extracts (petroleum), light paraffinic distillate solvent, acid-treated: Distillate aromatic extract (treated): [A complex combination of hydrocarbons obtained as a fraction of the distillation of an extract from the solvent extraction of light paraffinic top petroleum distillates that is subjected to a sulfuric acid refining. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>16</sub> through C<sub>32</sub>.] |  |  | N
309-672-2 | 100684-02-4 | Extracts (petroleum), light paraffinic distillate solvent, carbon-treated: Distillate aromatic extract (treated): [A complex combination of hydrocarbons obtained as a fraction from distillation of an extract recovered by solvent extraction of light paraffinic top petroleum distillate treated with activated charcoal to remove traces of polar constituents and impurities. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>16</sub> through C<sub>32</sub>.] |  |  | N
309-673-8 | 100684-03-5 | Extracts (petroleum), light paraffinic distillate solvent, clay-treated: Distillate aromatic extract (treated): [A complex combination of hydrocarbons obtained as a fraction from distillation of an extract recovered by solvent extraction of light paraffinic top petroleum distillates treated with bleaching earth to remove traces of polar constituents and impurities. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>16</sub> through C<sub>32</sub>.] |  |  | N
295-340-1 | 91995-77-6 | Extracts (petroleum), light paraffinic distillate solvent, hydrodesulfurized: Distillate aromatic extract (treated): [A complex combination of hydrocarbons obtained by solvent extraction of a light paraffin distillate and treated with hydrogen to convert the organic sulfur to hydrogen sulfide which is eliminated. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>40</sub> and produces a finished oil with a viscosity of greater than 10cSt at 40 °C.] |  |  | N
292-633-6 | 90641-09-1 | Extracts (petroleum), light paraffinic distillate solvent, hydrotreated: Distillate aromatic extract (treated): [A complex combination of hydrocarbons produced by treating a light paraffinic distillate solvent extract with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>17</sub> through C<sub>26</sub> and boiling in the range of approximately 280 °C to 400 °C (536 °F to 752 °F).] |  |  | N
295-341-7 | 91995-78-7 | Extracts (petroleum), light vacuum gas oil solvent |  |  | N
309-675-9 | 100684-05-7 | Extracts (petroleum), light vacuum gas oil solvent, clay-treated: Distillate aromatic extract (treated): [A complex combination of hydrocarbons obtained by solvent extraction of light vacuum petroleum gas oils treated with bleaching earth for removal of trace polar constituents and impurities. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>13</sub> through C<sub>30</sub>.] |  |  | N
295-342-2 | 91995-79-8 | Extracts (petroleum), light vacuum gas oil solvent, hydrotreated: Distillate aromatic extract (treated): [A complex combination of hydrocarbons, obtained by solvent extraction from light vacuum petroleum gas oils and treated with hydrogen in the presence of a catalyst. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>13</sub> through C<sub>30</sub>.] |  |  | N
309-674-3 | 100684-04-6 | Extracts (petroleum), light vacuum, gas oil solvent, carbon-treated: Distillate aromatic extract (treated): [A complex combination of hydrocarbons obtained by solvent extraction of light vacuum petroleum gas oil treated with activated charcoal for the removal of trace polar constituents and impurities. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>13</sub> through C<sub>30</sub>.] |  |  | N
297-829-5 | 93763-11-2 | Extracts (petroleum), solvent-dewaxed heavy paraffinic distillate solvent, hydrodesulfurized: Distillate aromatic extract (treated): [A complex combination of hydrocarbons obtained from a solvent dewaxed petroleum stock by treating with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of greater than 19cSt at 40 °C.] |  |  | N
272-180-0 | 68783-04-0 | Extracts (petroleum), solvent-refined heavy paraffinic distillate solvent: Distillate aromatic extract (treated): [A complex combination of hydrocarbons obtained as the extract from the re-extraction of solvent-refined heavy paraffinic distillate. It consists of saturated and aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub>.] |  |  | N
266-017-2 | 65996-83-0 | Extracts, coal tar oil alk.: Alkaline Extract: [The extract from coal tar oil produced by an alkaline wash such as aqueous sodium hydroxide. Composed primarily of the alkali salts of various phenolic compounds.] |  |  | N
294-302-1 | 91697-41-5 | fatty acids, C<sub>6-19</sub>-branched, nickel salts |  |  | N | Process catalyst
283-972-0 | 84776-45-4 | fatty acids, C<sub>8-18</sub> and C<sub>18</sub>-unsaturated, nickel salts |  |  | N | Process catalyst
421-960-0 | 90035-08-8 | flocoumafen (ISO): reaction mass of: <i>cis</i>-4-hydroxy-3-(1,2,3,4-tetrahydro-3-(4-(4-trifluoromethylbenzyloxy)phenyl)-1-naphthyl)coumarin and trans-4-hydroxy-3-(1,2,3,4-tetrahydro-3-(4-(4-trifluoromethylbenzyloxy)phenyl)-1-naphthyl)coumarin |  |  | N | Rat poison
274-125-6 | 69806-50-4, 79241-46-6 | fluazifop-butyl (ISO): butyl (<i>RS</i>)-2-[4-(5-trifluoromethyl-2-pyridyloxy)phenoxy]propionate |  |  | N | Agricultural herbicide
600-425-7 | 103361-09-7 | flumioxazin (ISO): 2-[7-fluoro-3-oxo-4-(prop-2-yn-1-yl)-3,4-dihydro-2H-1,4-benzoxazin-6-yl]-4,5,6,7-tetrahydro-1H-isoindole-1,3(2H)-dione |  |  | N | Agricultural herbicide
617-717-5 | 85509-19-9 | flusilazole (ISO): bis(4-fluorophenyl)(methyl)(1<i>H</i>-1,2,4-triazol-1-ylmethyl)silane |  |  | N | Agricultural fungicide
300-225-7 | 93924-31-3 | Foots oil (petroleum), acid-treated: Foots oil: [A complex combination of hydrocarbons obtained by treatment of Foot's oil with sulfuric acid. It consists predominantly of branched-chain hydrocarbons with carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub>.] |  |  | N
308-126-0 | 97862-76-5 | Foots oil (petroleum), carbon-treated: Foots oil: [A complex combination of hydrocarbons obtained by the treatment of Foots oil with activated carbon for the removal of trace constituents and impurities. It consists predominantly of saturated straight chain hydrocarbons having carbon numbers predominantly greater than C<sub>12</sub>.] |  |  | N
300-226-2 | 93924-32-4 | Foots oil (petroleum), clay-treated: Foots oil: [A complex combination of hydrocarbons obtained by treatment of Foot's oil with natural or modified clay in either a contacting or percolation process to remove the trace amounts of polar compounds and impurities present. It consists predominantly of branched chain hydrocarbons with carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub>.] |  |  | N
295-394-6 | 92045-12-0 | Foots oil (petroleum), hydrotreated: Foots oil: Refractory Ceramic Fibres, Special Purpose Fibres, with the exception of those specified elsewhere in this Annex: [Man-made vitreous (silicate) fibres with random orientation with alkaline oxide and alkali earth oxide (Na<sub>2</sub>O+K<sub>2</sub>O+CaO+ MgO+BaO) content less or equal to 18 % by weight] |  |  | N
308-127-6 | 97862-77-6 | Foots oil (petroleum), silicic acid-treated: Foots oil: [A complex combination of hydrocarbons obtained by the treatment of Foots oil with silicic acid for removal of trace constituents and impurities. It consists predominantly of straight chain hydrocarbons having carbon numbers predominantly greater than C<sub>12</sub>.] |  |  | N
265-171-8 | 64742-67-2 | Foots oil (petroleum): Foots oil: [A complex combination of hydrocarbons obtained as the oil fraction from a solvent deoiling or a wax sweating process. It consists predominantly of branched chain hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub>.] |  |  | N
200-001-8 | 50-00-0 | formaldehyde …% |  |  | 3 | Water soluble gas used as a process chemical. Can be generated by degradation of some types of polymers.
200-842-0 | 75-12-7 | formamide |  |  | N | Solvent and process chemical
268-755-0 | 68134-59-8 | formic acid, copper nickel salt |  |  | N | Process chemical
239-946-6 | 15843-02-4 | formic acid, nickel salt |  |  | N | Process chemical and catalyst
270-670-9 | 68476-29-9 | Fuel gases, crude oil of distillates: Petroleum gas: [A complex combination of light gases produced by distillation of crude oil and by catalytic reforming of naphtha. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C4 and boiling in the range of approximately – 217 °C to – 12 °C (– 423 °F to 10 °F).] |  |  | N
270-667-2 | 68476-26-6 | Fuel gases: Petroleum gas: [A combination of light gases. It consists predominantly of hydrogen and/or low molecular weight hydrocarbons.] |  |  | N
295-396-7 | 92045-14-2 | Fuel oil, heavy, high-sulfur: Heavy Fuel oil: [A complex combination of hydrocarbons obtained by the distillation of crude petroleum. It consists predominantly of aliphatic, aromatic and cycloaliphatic hydrocarbons having carbon numbers predominantly higher than C<sub>25</sub> and boiling above approximately 400 °C (752 °F).] |  |  | N
271-384-7 | 68553-00-4 | Fuel oil, No 6: Heavy Fuel oil: [A distillate oil having a minimum viscosity of 900 SUS at 37.7 °C (100 °F) to a maximum of 9000 SUS at 37.7 °C (100 °F).] |  |  | N
270-675-6 | 68476-33-5 | Fuel oil, residual: Heavy Fuel oil: [The liquid product from various refinery streams, usually residues. The composition is complex and varies with the source of the crude oil.] |  |  | N
270-674-0 | 68476-32-4 | Fuel oil, residues-straight-run gas oils, high-sulfur: Heavy Fuel oil |  |  | N
203-727-3 | 110-00-9 | furan |  |  | N | Process chemical
215-114-8 | 1303-00-0 | gallium arsenide |  |  | 2 | Semiconductor used inside ICs and LEDs. Also used in solar cells Always encapsulated so that physical contact should not be possible
| 58-89-9 | Gamma-HCH (Lindane) |  |  | N | Pesticide
265-112-6 | 64742-12-7 | Gas oils (petroleum), acid-treated: Gasoil - unspecified: [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>13</sub> through C<sub>25</sub> and boiling in the range of approximately 230 °C to 400 °C (446 °F to 752 °F).] |  |  | N
265-129-9 | 64742-29-6 | Gas oils (petroleum), chemically neutralized: Gasoil - unspecified: [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>13</sub> through C<sub>25</sub> and boiling in the range of approximately 230 °C to 400 °C (446 °F to 752 °F).] |  |  | N
272-184-2 | 68783-08-4 | Gas oils (petroleum), heavy atmospheric: Heavy Fuel oil: [A complex combination of hydrocarbons obtained by the distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>35</sub> and boiling in the range of approximately 121 °C to 510 °C (250 °F to 950 °F).] |  |  | N
265-058-3 | 64741-57-7 | Gas oils (petroleum), heavy vacuum: Heavy Fuel oil: [A complex combination of hydrocarbons produced by the vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and boiling in the range of approximately 350 °C to 600 °C (662 °F to 1112 °F). This stream is likely to contain 5 wt. % or more of 4-to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
285-555-9 | 85117-03-9 | Gas oils (petroleum), hydrodesulfurized coker heavy vacuum: Heavy Fuel oil: [A complex combination of hydrocarbons obtained by hydrodesulfurization of heavy coker distillate stocks, It consists predominantly of hydrocarbons having carbon numbers predominantly in the range C<sub>18</sub> to C<sub>44</sub> and boiling in the range of approximately 304 °C to 548 °C (579 °F to 1018 °F). Likely to contain 5 % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
265-189-6 | 64742-86-5 | Gas oils (petroleum), hydrodesulfurized heavy vacuum: Heavy Fuel oil: [A complex combination of hydrocarbons obtained from a catalytic hydrodesulfurization process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and boiling in the range of approximately 350 °C to 600 °C (662 °F to 1112 °C). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
265-182-8 | 64742-79-6 | Gas oils (petroleum), hydrodesulfurized: Gasoil - unspecified: [A complex combination of hydrocarbons obtained from a petroleum stock by treating with hydrogen to convert organic sulfur to hydrogen sulfide which is removed. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>13</sub> through C<sub>25</sub> and boiling in the range of approximately 230 °C to 400 °C (446 °F to 752 °F).] |  |  | N
265-162-9 | 64742-59-2 | Gas oils (petroleum), hydrotreated vacuum: Heavy Fuel oil: [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>13</sub> through C<sub>50</sub> and boiling in the range of approximately 230 °C to 600 °C (446 °F to 1112 °F). This stream is likely to contain 5 wt.% or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
308-278-8 | 97926-59-5 | Gas oils (petroleum), light vacuum, thermal-cracked hydrodesulfurized: Cracked gasoil: [A complex combination of hydrocarbons obtained by catalytic dehydrosulfurization of thermal-cracked light vacuum petroleum. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>14</sub> through C<sub>20</sub> and boiling in the range of approximately 270 °C to 370 °C (518 °F to 698 °F).] |  |  | N
265-092-9 | 64741-90-8 | Gas oils (petroleum), solvent-refined: Gasoil - unspecified: [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C<sub>11</sub> through C<sub>25</sub> and boiling in the range of approximately 205 °C to 400 °C (401 °F to 752 °F).] |  |  | N
271-260-2 | 68527-18-4 | Gas oils (petroleum), steam-cracked: Cracked gasoil: [A complex combination of hydrocarbons produced by distillation of the products from a steam cracking process. It consists of hydrocarbons having carbon numbers predominantly greater than C<sub>9</sub> and boiling in the range of from approximately 205 °C to 400 °C (400 °F to 752 °F).] |  |  | N
295-411-7 | 92045-29-9 | Gas oils (petroleum), thermal-cracked, hydrodesulfurized: Cracked gasoil |  |  | N
308-128-1 | 97862-78-7 | Gas oils, hydrotreated: Gasoil - unspecified: [A complex combination of hydrocarbons obtained from the redistillation of the effluents from the treatment of paraffins with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>17</sub> through C<sub>27</sub> and boiling in the range of approximately 330 °C to 340 °C (626 °F to 644 °F).] |  |  | N
300-227-8 | 93924-33-5 | Gas oils, paraffinic: Gasoil - unspecified: [A distillate obtained from the redistillation of a complex combination of hydrocarbons obtained by the distillation of the effluents from a severe catalytic hydrotreatment of paraffins. It boils in the range of approximately 190 °C to 330 °C (374 °F to 594 °F).] |  |  | N
271-737-5 | 68606-27-9 | Gases (petroleum), alkylation feed: Petroleum gas: [A complex combination of hydrocarbons produced by the catalytic cracking of gas oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C3 through C4.] |  |  | N
270-746-1 | 68477-65-6 | Gases (petroleum), amine system feed: Refinery gas: [The feed gas to the amine system for removal of hydrogen sulfide. It consists of hydrogen. Carbon monoxide, carbon dioxide, hydrogen sulfide and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5 may also be present.] |  |  | N
270-747-7 | 68477-66-7 | Gases (petroleum), benzene unit hydrodesulfurizer off: Refinery gas: [Off gases produced by the benzene unit. It consists primarily of hydrogen. Carbon monoxide and hydrocarbons having carbon numbers predominantly in the range of C1 through C6, including benzene, may also be present.] |  |  | N
271-623-5 | 68602-82-4 | Gases (petroleum), benzene unit hydrotreater depentanizer overheads: Refinery gas: [A complex combination produced by treating the feed from the benzene unit with hydrogen in the presence of a catalyst followed by depentanizing. It consists primarily of hydrogen, ethane and propane with various small amounts of nitrogen, carbon monoxide, carbon dioxide and hydrocarbons having carbon numbers predominantly in the range of C1 through C6. It may contain trace amounts of benzene.] |  |  | N
270-748-2 | 68477-67-8 | Gases (petroleum), benzene unit recycle, hydrogen-rich: Refinery gas: [A complex combination of hydrocarbons obtained by recycling the gases of the benzene unit. It consists primarily of hydrogen with various small amounts of carbon monoxide and hydrocarbons having carbon numbers in the range of C1 through C6.] |  |  | N
270-749-8 | 68477-68-9 | Gases (petroleum), blend oil, hydrogen-nitrogen-rich: Refinery gas: [A complex combination of hydrocarbons obtained by distillation of a blend oil. It consists primarily of hydrogen and nitrogen with various small amounts of carbon monoxide, carbon dioxide, and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
270-750-3 | 68477-69-0 | Gases (petroleum), butane splitter overheads: Petroleum gas: [A complex combination of hydrocarbons obtained from the distillation of the butane stream. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C3 through C4.] |  |  | N
271-624-0 | 68602-83-5 | Gases (petroleum), C1-5, wet: Petroleum gas: [A complex combination of hydrocarbons produced by the distillation of crude oil and/or the cracking of tower gas oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
270-751-9 | 68477-70-3 | Gases (petroleum), C2-3-: Petroleum gas: [A complex combination of hydrocarbons produced by the distillation of products from a catalytic fractionation process. It contains predominantly ethane, ethylene, propane, and propylene.] |  |  | N
272-205-5 | 68783-65-3 | Gases (petroleum), C2-4, sweetened: Petroleum gas: [A complex combination of hydrocarbons obtained by subjecting a petroleum distillate to a sweetening process to convert mercaptans or to remove acidic impurities. It consists predominantly of saturated and unsaturated hydrocarbons having carbon numbers predominantly in the range of C2 through C4 and boiling in the range of approximately – 51 °C to – 34 °C (– 60 °F to – 30 °F).] |  |  | N
270-766-0 | 68477-84-9 | Gases (petroleum), C2-return stream: Refinery gas: [A complex combination of hydrocarbons obtained by the extraction of hydrogen from a gas stream which consists primarily of hydrogen with small amounts of nitrogen, carbon monoxide, methane, ethane, and ethylene. It contains predominantly hydrocarbons such as methane, ethane, and ethylene with small amounts of hydrogen, nitrogen and carbon monoxide.] |  |  | N
270-724-1 | 68477-33-8 | gases (petroleum), C3-4, isobutane-rich: Petroleum gas: [A complex combination of hydrocarbons from the distillation of saturated and unsaturated hydrocarbons usually ranging in carbon numbers from C3 through C6, predominantly butane and isobutane. It consists of saturated and unsaturated hydrocarbons having carbon numbers in the range of C3 through C4, predominantly isobutane.] |  |  | N
268-629-5 | 68131-75-9 | Gases (petroleum), C3-4: Petroleum gas: [A complex combination of hydrocarbons produced by distillation of products from the cracking of crude oil. It consists of hydrocarbons having carbon numbers in the range of C3 through C4, predominantly of propane and propylene, and boiling in the range of approximately – 51 °C to – 1 °C (– 60 °F to 30 °F.)] |  |  | N
270-765-5 | 68477-83-8 | Gases (petroleum), C3-5 olefinic-paraffinic alkylation feed: Petroleum gas: [A complex combination of olefinic and paraffinic hydrocarbons having carbon numbers in the range of C3 through C5 which are used as alkylation feed. Ambient temperatures normally exceed the critical temperature of these combinations.] |  |  | N
270-767-6 | 68477-85-0 | Gases (petroleum), C4-rich: Petroleum gas: [A complex combination of hydrocarbons produced by distillation of products from a catalytic fractionation process. It consists of aliphatic hydrocarbons having carbon numbers in the range of C3 through C5, predominantly C4.] |  |  | N
270-763-4 | 68477-82-7 | Gases (petroleum), C6-8 catalytic reformer recycle, hydrogen-rich: Refinery gas |  |  | N
270-761-3 | 68477-80-5 | Gases (petroleum), C6-8 catalytic reformer recycle: Refinery gas: [A complex combination of hydrocarbons produced by distillation of products from catalytic reforming of C6-C8 feed and recycled to conserve hydrogen. It consists primarily of hydrogen. It may also contain various small amounts of carbon monoxide, carbon dioxide, nitrogen, and hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
270-762-9 | 68477-81-6 | Gases (petroleum), C6-8 catalytic reformer: Refinery gas: [A complex combination of hydrocarbons produced by distillation of products from catalytic reforming of C6-C8feed. It consists of hydrocarbons having carbon numbers in the range of C1 through C5 and hydrogen.] |  |  | N
273-169-3 | 68952-76-1 | Gases (petroleum), catalytic cracked naphtha debutanizer: Petroleum gas: [A complex combination of hydrocarbons obtained from fractionation of catalytic cracked naphtha. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
270-755-0 | 68477-73-6 | Gases (petroleum), catalytic cracked naphtha depropanizer overhead, C3-rich acid-free: Petroleum gas: [A complex combination of hydrocarbons obtained from fractionation of catalytic cracked hydrocarbons and treated to remove acidic impurities. It consists of hydrocarbons having carbon numbers in the range of C2 through C4, predominantly C3.] |  |  | N
270-071-2 | 68409-99-4 | Gases (petroleum), catalytic cracked overheads: Petroleum gas: [A complex combination of hydrocarbons produced by the distillation of products from the catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C3 through C5 and boiling in the range of approximately – 48 °C to 32 °C (– 54 °F to 90 °F).] |  |  | N
270-757-1 | 68477-75-8 | Gases (petroleum), catalytic cracker, C1-5-rich: Petroleum gas: [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process. It consists of aliphatic hydrocarbons having carbon numbers in the range of C1 through C6, predominantly C1 through C5.] |  |  | N
270-756-6 | 68477-74-7 | Gases (petroleum), catalytic cracker: Petroleum gas: [A complex combination of hydrocarbons produced by the distillation of the products from a catalytic cracking process. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
272-203-4 | 68783-64-2 | Gases (petroleum), catalytic cracking: Petroleum gas: [A complex combination of hydrocarbons produced by the distillation of the products from a catalytic cracking process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C3 through C5.] |  |  | N
270-758-7 | 68477-76-9 | Gases (petroleum), catalytic polymd. naphtha stabilizer overhead, C2-4-rich: Petroleum gas: [A complex combination of hydrocarbons obtained from the fractionation stabilization of catalytic polymerized naphtha. It consists of aliphatic hydrocarbons having carbon numbers in the range of C2 through C6, predominantly C2 through C4.] |  |  | N
270-759-2 | 68477-77-0 | Gases (petroleum), catalytic reformed naphtha stripper overheads: Refinery gas: [A complex combination of hydrocarbons obtained from stabilization of catalytic reformed naphtha. Its consists of hydrogen and saturated hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
270-999-8 | 68513-14-4 | Gases (petroleum), catalytic reformed straight-run naphtha stabilizer overheads: Refinery gas: [A complex combination of hydrocarbons obtained from the catalytic reforming of straight-run naphtha followed by fractionation of the total effluent. It consists of hydrogen, methane, ethane and propane.] |  |  | N
270-760-8 | 68477-79-2 | Gases (petroleum), catalytic reformer, C1-4-rich: Petroleum gas: [A complex combination of hydrocarbons produced by distillation of products from a catalytic reforming process. It consists of hydrocarbons having carbon numbers in the range of C1 through C6, predominantly C1 through C4.] |  |  | N
270-752-4 | 68477-71-4 | Gases (petroleum), catalytic-cracked gas oil depropanizer bottoms, C4-rich acid-free: Petroleum gas: [A complex combination of hydrocarbons obtained from fractionation of catalytic cracked gas oil hydrocarbon stream and treated to remove hydrogen sulfide and other acidic components. It consists of hydrocarbons having carbon numbers in the range of C3 through C5, predominantly C4.] |  |  | N
270-754-5 | 68477-72-5 | Gases (petroleum), catalytic-cracked naphtha debutanizer bottoms, C3-5-rich: Petroleum gas: [A complex combination of hydrocarbons obtained from the stabilization of catalytic cracked naphtha. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C3 through C5.] |  |  | N
273-563-5 | 68989-88-8 | Gases (petroleum), crude distn. and catalytic cracking: Refinery gas: [A complex combination produced by crude distillation and catalytic cracking processes. It consists of hydrogen, hydrogen sulfide, nitrogen, carbon monoxide and paraffinic and olefinic hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
272-871-7 | 68918-99-0 | Gases (petroleum), crude oil fractionation off: Petroleum gas: [A complex combination of hydrocarbons produced by the fractionation of crude oil. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
270-768-1 | 68477-86-1 | Gases (petroleum), deethanizer overheads: Petroleum gas: [A complex combination of hydrocarbons produced from distillation of the gas and gasoline fractions from the catalytic cracking process. It contains predominantly ethane and ethylene.] |  |  | N
272-872-2 | 68919-00-6 | Gases (petroleum), dehexanizer off: Petroleum gas: [A complex combination of hydrocarbons obtained by the fractionation of combined naphtha streams. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
270-769-7 | 68477-87-2 | Gases (petroleum), deisobutanizer tower overheads: Petroleum gas: [A complex combination of hydrocarbons produced by the atmospheric distillation of a butane-butylene stream. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C3 through C4.] |  |  | N
271-742-2 | 68606-34-8 | Gases (petroleum), depropanizer bottoms fractionation off: Petroleum gas: [A complex combination of hydrocarbons obtained from the fractionation of depropanizer bottoms. It consists predominantly of butane, isobutane and butadiene.] |  |  | N
270-772-3 | 68477-90-7 | Gases (petroleum), depropanizer dry, propene-rich: Petroleum gas: [A complex combination of hydrocarbons produced by the distillation of products from the gas and gasoline fractions of a catalytic cracking process. It consists predominantly of propylene with some ethane and propane.] |  |  | N
270-773-9 | 68477-91-8 | Gases (petroleum), depropanizer overheads: Petroleum gas: [A complex combination of hydrocarbons produced by distillation of products from the gas and gasoline fractions of a catalytic cracking process. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C4.] |  |  | N
272-873-8 | 68919-01-7 | Gases (petroleum), distillate unifiner desulfurization stripper off: Refinery gas: [A complex combination stripped from the liquid product of the unifiner desulfurization process. It consists of hydrogen sulfide, methane, ethane, and propane.] |  |  | N
270-774-4 | 68477-92-9 | Gases (petroleum), dry sour, gas-concn.-unit-off: Refinery gas: [The complex combination of dry gases from a gas concentration unit. It consists of hydrogen, hydrogen sulfide and hydrocarbons having carbon numbers predominantly in the range of C1 through C3.] |  |  | N
272-874-3 | 68919-02-8 | Gases (petroleum), fluidized catalytic cracker fractionation off: Refinery gas: [A complex combination produced by the fractionation of the overhead product of the fluidized catalytic cracking process. It consists of hydrogen, hydrogen sulfide, nitrogen, and hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
272-875-9 | 68919-03-9 | Gases (petroleum), fluidized catalytic cracker scrubbing secondary absorber off: Refinery gas: [A complex combination produced by scrubbing the overhead gas from the fluidized catalytic cracker. It consists of hydrogen, nitrogen, methane, ethane and propane.] |  |  | N
272-893-7 | 68919-20-0 | Gases (petroleum), fluidized catalytic cracker splitter overheads: Petroleum gas: [A complex combination of hydrocarbons produced by the fractionation of the charge to the C3-C4 splitter. It consists predominantly of C3 hydrocarbons.] |  |  | N
271-000-8 | 68513-15-5 | Gases (petroleum), full-range straight-run naphtha dehexanizer off: petroleum gas: [A complex combination of hydrocarbons obtained by the fractionation of the full-range straight-run naphtha. It consists of hydrocarbons having carbon numbers predominantly in the range of C2 through C6.] |  |  | N
270-776-5 | 68477-93-0 | Gases (petroleum), gas concn. reabsorber distn.: Refinery gas: [A complex combination of hydrocarbons produced by distillation of products from combined gas streams in a gas concentration reabsorber. It consists predominantly of hydrogen, carbon monoxide, carbon dioxide, nitrogen, hydrogen sulfide and hydrocarbons having carbon numbers in the range of C1 through C3.] |  |  | N
295-397-2 | 92045-15-3 | Gases (petroleum), gas oil diethanolamine scrubber off: Refinery gas: [A complex combination produced by desulfurization of gas oils with diethanolamine. It consists predominantly of hydrogen sulfide, hydrogen and aliphatic hydrocarbons having carbon numbers in the range of C1 through C5.] |  |  | N
295-398-8 | 92045-16-4 | Gases (petroleum), gas oil hydrodesulfurization effluent: Refinery gas: [A complex combination obtained by separation of the liquid phase from the effluent from the hydrogenation reaction. It consists predominantly of hydrogen, hydrogen sulfide and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C3.] |  |  | N
295-399-3 | 92045-17-5 | Gases (petroleum), gas oil hydrodesulfurization purge: Refinery gas: [A complex combination of gases obtained from the reformer and from the purges from the hydrogenation reactor. It consists predominantly of hydrogen and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
270-777-0 | 68477-94-1 | Gases (petroleum), gas recovery plant depropanizer overheads: Petroleum gas: [A complex combination of hydrocarbons obtained by fractionation of miscellaneous hydrocarbon streams. It consists predominantly of hydrocarbons having carbon numbers in the range of C1 through C4, predominantly propane.] |  |  | N
270-778-6 | 68477-95-2 | Gases (petroleum), Girbotol unit feed: Petroleum gas: [A complex combination of hydrocarbons that is used as the feed into the Girbatol unit to remove hydrogen sulfide. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C4.] |  |  | N
272-876-4 | 68919-04-0 | Gases (petroleum), heavy distillate hydrotreater desulfurization stripper off: Refinery gas: [A complex combination stripped from the liquid product of the heavy distillate hydrotreater desulfurization process. It consists of hydrogen, hydrogen sulfide, and saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
271-001-3 | 68513-16-6 | Gases (petroleum), hydrocracking depropanizer off, hydrocarbon-rich: Petroleum gas: [A complex combination of hydrocarbon produced by the distillation of products from a hydrocracking process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C4. It may also contain small amounts of hydrogen and hydrogen sulfide.] |  |  | N
272-182-1 | 68783-06-2 | Gases (petroleum), hydrocracking low-pressure separator: Refinery gas: [A complex combination obtained by the liquid-vapor separation of the hydrocracking process reactor effluent. It consists predominantly of hydrogen and saturated hydrocarbons having carbon numbers predominantly in the range of C1 through C3.] |  |  | N
270-779-1 | 68477-96-3 | Gases (petroleum), hydrogen absorber off: Refinery gas: [A complex combination obtained by absorbing hydrogen from a hydrogen rich stream. It consists of hydrogen, carbon monoxide, nitrogen, and methane with small amounts of C2 hydrocarbons.] |  |  | N
295-400-7 | 92045-18-6 | Gases (petroleum), hydrogenator effluent flash drum off: Refinery gas: [A complex combination of gases obtained from flash of the effluents after the hydrogenation reaction. It consists predominantly of hydrogen and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
270-780-7 | 68477-97-4 | Gases (petroleum), hydrogen-rich: Refinery gas: [A complex combination separated as a gas from hydrocarbon gases by chilling. It consists primarily of hydrogen with various small amounts of carbon monoxide, nitrogen, methane, and C2 hydrocarbons.] |  |  | N
272-775-5 | 68911-58-0 | Gases (petroleum), hydrotreated sour kerosine depentanizer stabilizer off: Refinery gas: [The complex combination obtained from the depentanizer stabilization of hydrotreated kerosine. It consists primarily of hydrogen, methane, ethane, and propane with various small amounts of nitrogen, hydrogen sulfide, carbon monoxide and hydrocarbons having carbon numbers predominantly in the range of C4 through C5.] |  |  | N
272-776-0 | 68911-59-1 | Gases (petroleum), hydrotreated sour kerosine flash drum: Refinery gas: [A complex combination obtained from the flash drum of the unit treating sour kerosine with hydrogen in the presence of a catalyst. It consists primarily of hydrogen and methane with various small amounts of nitrogen, carbon monoxide, and hydrocarbons having carbon numbers predominantly in the range of C2 through C5.] |  |  | N
270-781-2 | 68477-98-5 | Gases (petroleum), hydrotreater blend oil recycle, hydrogen-nitrogen-rich: Refinery gas: [A complex combination obtained from recycled hydrotreated blend oil. It consists primarily of hydrogen and nitrogen with various small amounts of carbon monoxide, carbon dioxide and hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
270-782-8 | 68477-99-6 | Gases (petroleum), isomerized naphtha fractionator, C4-rich, hydrogen sulfide-free: Petroleum gas |  |  | N
272-878-5 | 68919-05-1 | Gases (petroleum), light straight run gasoline fractionation stabilizer off: Petroleum gas: [A complex combination of hydrocarbons obtained by the fractionation of light straight-run gasoline. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
271-002-9 | 68513-17-7 | Gases (petroleum), light straight-run naphtha stabilizer off: Petroleum gas: [A complex combination of hydrocarbons obtained by the stabilization of light straight-run naphtha. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C6.] |  |  | N
295-401-2 | 92045-19-7 | Gases (petroleum), naphtha steam cracking high-pressure residual: Refinery gas: [A complex combination obtained as a reaction mass of the non-condensable portions from the product of a naphtha steam cracking process as well as residual gases obtained during the preparation of subsequent products. It consists predominantly of hydrogen and paraffinic and olefinic hydrocarbons having carbon numbers predominantly in the range of C1 through C5 with which natural gas may also be mixed.] |  |  | N
272-879-0 | 68919-06-2 | Gases (petroleum), naphtha unifiner desulfurization stripper off: Petroleum gas: [A complex combination of hydrocarbons produced by a naphtha unifiner desulfurization process and stripped from the naphtha product. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
271-258-1 | 68527-15-1 | Gases (petroleum), oil refinery gas distn. off: Refinery gas: [A complex combination separated by distillation of a gas stream containing hydrogen, carbon monoxide, carbon dioxide and hydrocarbons having carbon numbers in the range of C1 through C6 or obtained by cracking ethane and propane. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C2, hydrogen, nitrogen, and carbon monoxide.] |  |  | N
272-343-6 | 68814-90-4 | Gases (petroleum), platformer products separator off: Refinery gas: [A complex combination obtained from the chemical reforming of naphthenes to aromatics. It consists of hydrogen and saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C4.] |  |  | N
272-880-6 | 68919-07-3 | Gases (petroleum), platformer stabilizer off, light ends fractionation: Refinery gas: [A complex combination obtained by the fractionation of the light ends of the platinum reactors of the platformer unit. It consists of hydrogen, methane, ethane and propane.] |  |  | N
272-881-1 | 68919-08-4 | Gases (petroleum), preflash tower off, crude distn.: Refinery gas: [A complex combination produced from the first tower used in the distillation of crude oil. It consists of nitrogen and saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
270-783-3 | 68478-00-2 | Gases (petroleum), recycle, hydrogen-rich: Refinery gas: [A complex combination obtained from recycled reactor gases. It consists primarily of hydrogen with various small amounts of carbon monoxide, carbon dioxide, nitrogen, hydrogen sulfide, and saturated aliphatic hydrocarbons having carbon numbers in the range of C1 through C5.] |  |  | N
272-183-7 | 68783-07-3 | Gases (petroleum), refinery blend: Petroleum gas: [A complex combination obtained from various processes. It consists of hydrogen, hydrogen sulfide and hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
272-338-9 | 68814-67-5 | Gases (petroleum), refinery: Refinery gas: [A complex combination obtained from various petroleum refining operations. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C3.] |  |  | N
271-003-4 | 68513-18-8 | Gases (petroleum), reformer effluent high-pressure flash drum off: Refinery gas: [A complex combination produced by the high-pressure flashing of the effluent from the reforming reactor. It consists primarily of hydrogen with various small amounts of methane, ethane, and propane.] |  |  | N
271-005-5 | 68513-19-9 | Gases (petroleum), reformer effluent low-pressure flash drum off: Refinery gas: [A complex combination produced by low-pressure flashing of the effluent from the reforming reactor. It consists primarily of hydrogen with various small amounts of methane, ethane, and propane.] |  |  | N
270-784-9 | 68478-01-3 | Gases (petroleum), reformer make-up, hydrogen-rich: Refinery gas: [A complex combination obtained from the reformers. It consists primarily of hydrogen with various small amounts of carbon monoxide and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
270-788-0 | 68478-04-6 | Gases (petroleum), reforming hydrotreater make-up, hydrogen-rich: Refinery gas: [A complex combination obtained from the reforming hydrotreating process. It consists primarily of hydrogen with various small amounts of carbon monoxide and aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
270-787-5 | 68478-03-5 | Gases (petroleum), reforming hydrotreater, hydrogen-methane-rich: Refinery gas: [A complex combination obtained from the reforming hydrotreating process. It consists primarily of hydrogen and methane with various small amounts of carbon monoxide, carbon dioxide, nitrogen and saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C5.] |  |  | N
270-785-4 | 68478-02-4 | Gases (petroleum), reforming hydrotreater: Refinery gas: [A complex combination obtained from the reforming hydrotreating process. It consists primarily of hydrogen, methane, and ethane with various small amounts of hydrogen sulfide and aliphatic hydrocarbons having carbon numbers predominantly in the range of C3 through C5.] |  |  | N
295-402-8 | 92045-20-0 | Gases (petroleum), residue visbaking off: Refinery gas: [A complex combination obtained from viscosity reduction of residues in a furnace. It consists predominantly of hydrogen sulfide and paraffinic and olefinic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
271-625-6 | 68602-84-6 | Gases (petroleum), secondary absorber off, fluidized catalytic cracker overheads fractionator: Refinery gas: [A complex combination produced by the fractionation of the overhead products from the catalytic cracking process in the fluidized catalytic cracker. It consists of hydrogen, nitrogen, and hydrocarbons having carbon numbers predominantly in the range of C1 through C3.] |  |  | N
273-269-7 | 68955-33-9 | Gases (petroleum), sponge absorber off, fluidized catalytic cracker and gas oil desulfurizer overhead fractionation: Refinery gas: [A complex combination obtained by the fractionation of products from the fluidized catalytic cracker and gas oil desulfurizer. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
295-404-9 | 92045-22-2 | Gases (petroleum), steam-cracker C3-rich: Petroleum gas: [A complex combination of hydrocarbons produced by the distillation of products from a steam cracking process. It consists predominantly of propylene with some propane and boils in the range of approximately – 70 °C to 0 °C (– 94 °F to 32 °F).] |  |  | N
273-270-2 | 68955-34-0 | Gases (petroleum), straight-run naphtha catalytic reformer stabilizer overhead: Petroleum gas: [A complex combination of hydrocarbons obtained by the catalytic reforming of straight-run naphtha and the fractionation of the total effluent. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C2 through C4.] |  |  | N
272-882-7 | 68919-09-5 | Gases (petroleum), straight-run naphtha catalytic reforming off: Petroleum gas: [A complex combination of hydrocarbons obtained by the catalytic reforming of straight-run naphtha and fractionation of the total effluent. It consists of methane, ethane, and propane.] |  |  | N
272-883-2 | 68919-10-8 | Gases (petroleum), straight-run stabilizer off: Petroleum gas: [A complex combination of hydrocarbons obtained from the fractionation of the liquid from the first tower used in the distillation of crude oil. It consists of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
272-884-8 | 68919-11-9 | Gases (petroleum), tar stripper off: Refinery gas: [A complex combination obtained by the fractionation of reduced crude oil. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
270-789-6 | 68478-05-7 | Gases (petroleum), thermal cracking distn.: Refinery gas: [A complex combination produced by distillation of products from a thermal cracking process. It consists of hydrogen, hydrogen sulfide, carbon monoxide, carbon dioxide and hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
272-885-3 | 68919-12-0 | Gases (petroleum), unifiner stripper off: Refinery gas: [A combination of hydrogen and methane obtained by fractionation of the products from the unifiner unit.] |  |  | N
273-265-5 | 68955-28-2 | Gases (petroleum, light steam-cracked, butadiene conc.: Petroleum gas: [A complex combination of hydrocarbons produced by the distillation of products from a thermal cracking process. It consists of hydrocarbons having a carbon number predominantly of C4.] |  |  | N
297-458-9 | 93572-29-3 | Gasoline, C<sub>5-11</sub>, high-octane stabilised reformed: Low boiling point cat-reformed naphtha: [A complex high octane combination of hydrocarbons obtained by the catalytic dehydrogenation of a predominantly naphthenic naphtha. It consists predominantly of aromatics and non-aromatics having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>11</sub> and boiling in the range of approximately 45°C to 185°C (113°F to 365°F).] |  |  | N
302-691-7 | 94114-55-3 | Gasoline, coal solvent extn., hydrocracked naphtha: [Motor fuel produced by the reforming of the refined naphtha fraction of the products of hydrocracking of coal extract or solution produced by the liquid solvent extraction or supercritical gas extraction processes and boiling in the range of approximately 30 °C to 180 °C (86 °F to 356 °F). Composed primarily of aromatic and naphthenic hydrocarbons, their alkyl derivatives and alkyl hydrocarbons having carbon numbers in the range of C<sub>4</sub> through C<sub>9</sub>.] |  |  | N
232-349-1 | 8006-61-9 | Gasoline, natural: Low boiling point naphtha: [A complex combination of hydrocarbons separated from natural gas by processes such as refrigeration or absorption. It consists predominantly of saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>8</sub> and boiling in the range of approximately minus 20°C to 120°C (-4°F to 248°F).] |  |  | N
271-726-5 | 68606-10-0 | Gasoline, pyrolysis, debutanizer bottoms: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained from the fractionation of depropanizer bottoms. It consists of hydrocarbons having carbon numbers predominantly greater than C<sub>5</sub>.] |  |  | N
302-639-3 | 94114-03-1 | Gasoline, pyrolysis, hydrogenated: Low boiling point naphtha-unspecified: [A distillation fraction from the hydrogenation of pyrolysis gasoline boiling in the range of approximately 20°C to 200°C (68°F to 392°F).] |  |  | N
271-727-0 | 68606-11-1 | Gasoline, straight-run, topping-plant: Low boiling point naphtha: [A complex combination of hydrocarbons produced from the topping plant by the distillation of crude oil. It boils in the range of approximately 36.1°C to 193.3°C (97°F to 380°F).] |  |  | N
271-025-4 | 68514-15-8 | Gasoline, vapor-recovery: Low boiling point naphtha: [A complex combination of hydrocarbons separated from the gases from vapor recovery systems by cooling. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>11</sub> and boiling in the range of approximately -20°C to 196°C(-4°F to 384°F).] |  |  | N
289-220-8 | 86290-81-5 | Gasoline: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons consisting primarily of paraffins, cycloparaffins, aromatic and olefinic hydrocarbons having carbon numbers predominantly greater than C<sub>3</sub> and boiling in the range of 30°C to 260°C (86°F to 500°F).] |  |  | N
278-636-5 | 77182-82-2 | glufosinate ammonium (ISO): ammonium 2-amino-4-(hydroxymethylphosphinyl)butyrate |  |  | N | Agricultural herbicide
| 12035-71-1 | heazlewoodite |  |  | N | Nickel sulphide mineral
204-273-9 | 118-74-1 | hexachlorobenzene |  |  | N | Banned pesticide
211-653-8 | 680-31-9 | hexamethylphosphoric triamide: hexamethylphosphoramide |  |  | N | Solvent and process chemical
206-114-9 | 302-01-2, 7803-57-8 | hydrazine |  |  | N | Process chemical and rocket fuel
405-030-1 |  | hydrazine bis(3-carboxy-4-hydroxybenzensulfonate): |  |  | N | None known (probably used as a lab chemical or process chemical)
414-850-9 |  | hydrazine-trinitromethane |  |  | N | Rocket fuel
204-563-5 | 122-66-7 | hydrazobenzene: 1,2-diphenylhydrazine |  |  | N | Process chemical
309-745-9 | 100801-63-6 | Hydrocarbon oils, arom., mixed with polyethylene and polypropylene, pyrolyzed, light oil fraction: Heat Treatment Products: [The oil obtained from the heat treatment of a polyethylene/polypropylene reaction mass with coal tar pitch or aromatic oils. It consists predominantly of benzene and its homologs boiling in a range of approximately 70°C to 120°C (158°F to 248°F).] |  |  | N
309-748-5 | 100801-65-8 | Hydrocarbon oils, arom., mixed with polyethylene, pyrolyzed, light oil fraction: Heat Treatment Products: [The oil obtained from the heat treatment of polyethylene with coal tar pitch or aromatic oils. It consists predominantly of benzene and its homologs boiling in a range of 70°C to 120°C (158°F to 248°F).] |  |  | N
309-749-0 | 100801-66-9 | Hydrocarbon oils, arom., mixed with polystyrene, pyrolyzed, light oil fraction: Heat Treatment Products: [The oil obtained from the heat treatment of polystyrene with coal tar pitch or aromatic oils. It consists predominantly of benzene and its homologs boiling in a range of approximately 70°C to 210°C (158°F to 410°F).] |  |  | N
307-753-7 | 97722-04-8 | hydrocarbons C<sub>26-55</sub>, arom-rich |  |  | N
270-690-8 | 68476-50-6 | Hydrocarbons, C<sub>≥5</sub>, C<sub>5-6</sub>-rich: Low boiling point naphtha - unspecified |  |  | N
307-757-9 | 97722-08-2 | Hydrocarbons, C<sub>11-17</sub>, solvent-extd. light naphthenic: Gasoil - unspecified: [A complex combination of hydrocarbons obtained by extraction of the aromatics from a light naphthenic distillate having a visciosity of 2.2 cSt at 40 °C (104 °F). It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>11</sub> through C<sub>17</sub> and boiling in the range of approximately 200 °C to 300 °C (392 °F to 572 °F).] |  |  | N
307-660-1 | 97675-86-0 | Hydrocarbons, C<sub>12-20</sub>, hydrotreated paraffinic, distn. lights: Gasoil - unspecified: [A complex combination of hydrocarbons obtained as first runnings from the vacuum distillation of effluents from the treatment of heavy paraffins with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>12</sub> through C<sub>20</sub> and boiling in the range of approximately 230 °C to 350 °C (446 °F to 662 °F). It produces a finished oil having a viscosity of 2cSt at 100 °C (212 °F).] |  |  | N
307-758-4 | 97722-09-3 | Hydrocarbons, C<sub>13-27</sub>, solvent-extd. light naphthenic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by extraction of the aromatics from a light naphthenic distillate having a viscosity of 9.5cSt at 40 °C (104 °F). It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>13</sub> through C<sub>27</sub> and boiling in the range of approximately 240 °C to 400 °C (464 °F to 752 °F.] |  |  | N
305-971-7 | 95371-04-3 | Hydrocarbons, C<sub>13-30</sub>, arom.-rich, solvent-extd. naphthenic distillate: Baseoil - unspecified |  |  | N
307-760-5 | 97722-10-6 | Hydrocarbons, C<sub>14-29</sub>, solvent-extd. light naphthenic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by extraction of the aromatics from a light naphthenic distillate having a viscosity of 16cSt at 40 °C (104 °F). It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>14</sub> through C<sub>29</sub> and boiling in the range of approximately 250 °C to 425 °C (482 °F to 797 °F).] |  |  | N
307-659-6 | 97675-85-9 | Hydrocarbons, C<sub>16-20</sub>, hydrotreated middle distillate, distn. lights: Gasoil - unspecified: [A complex combination of hydrocarbons obtained as first runnings from the vacuum distillation of effluents from the treatment of a middle distillate with hydrogen. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>16</sub> through C<sub>20</sub> and boiling in the range of approximately 290 °C to 350 °C (554 °F to 662 °F). It produces a finished oil having a viscosity of 2cSt at 100 °C (212 °F).] |  |  | N
305-972-2 | 95371-05-4 | Hydrocarbons, C<sub>16-32</sub>, arom. rich, solvent-extd. naphthenic distillate: Baseoil - unspecified |  |  | N
308-132-3 | 97862-82-3 | Hydrocarbons, C<sub>17-30</sub>, hydrotreated distillates, distn. lights: Baseoil - unspecified |  |  | N
307-661-7 | 97675-87-1 | Hydrocarbons, C<sub>17-30</sub>, hydrotreated solvent-deasphalted atm. distn. residue, distn. lights: Baseoil - unspecified: [A complex combination of hydrocarbons obtained as first runnings from the vacuum distillation of effluents from the treatment of a solvent deasphalted short residue with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>17</sub> through C<sub>30</sub> and boiling in the range of approximately 300 °C to 400 °C (572 °F to 752 °F). It produces a finished oil having a viscosity of 4cSt at approximately 100 °C (212 °F).] |  |  | N
307-755-8 | 97722-06-0 | Hydrocarbons, C<sub>17-40</sub>, hydrotreated solvent-deasphalted distn. residue, vacuum distn. lights: Baseoil - unspecified: [A complex combination of hydrocarbons obtained as first runnings from the vacuum distillation of effluents from the catalytic hydrotreatment of a solvent deasphalted short residue having a viscosity of 8cSt at approximately 100 °C (212 °F). It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>17</sub> through C<sub>40</sub> and boiling in the range of approximately 300 °C to 500 °C (592 °F to 932 °F).] |  |  | N
300-257-1 | 93924-61-9 | Hydrocarbons, C<sub>20-50</sub>, residual oil hydrogenation vacuum distillate: Baseoil - unspecified |  |  | N
292-617-9 | 90640-95-2 | Hydrocarbons, C<sub>20-50</sub>, solvent dewaxed heavy paraffinic, hydrotreated: Baseoil - unspecified: [A complex combination of hydrocarbons produced by treating dewaxed heavy paraffinic distillate with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub>.] |  |  | N
308-289-8 | 97926-70-0 | Hydrocarbons, C<sub>20-58</sub>, hydrotreated: Baseoil - unspecified |  |  | N
270-687-1 | 68476-47-1 | Hydrocarbons, C<sub>2-6</sub>, C<sub>6-8</sub> catalytic reformer: Low boiling point cat-reformed naphtha |  |  | N
308-131-8 | 97862-81-2 | Hydrocarbons, C<sub>27-42</sub>, dearomatized: Baseoil - unspecified |  |  | N
308-290-3 | 97926-71-1 | Hydrocarbons, C<sub>27-42</sub>, naphthenic: Baseoil - unspecified |  |  | N
308-287-7 | 97926-68-6 | Hydrocarbons, C<sub>27-45</sub>, dearomatized: Baseoil - unspecified |  |  | N
308-133-9 | 97862-83-4 | Hydrocarbons, C<sub>27-45</sub>, naphthenic vacuum distn.: Baseoil - unspecified |  |  | N
270-686-6 | 68476-46-0 | Hydrocarbons, C<sub>3-11</sub>, catalytic cracker distillates: Low boiling point cat-cracked naphtha: [A complex combination of hydrocarbons produced by the distillations of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>3</sub> through C<sub>11</sub> and boiling in a range approximately up to 204°C (400°F).] |  |  | N
310-012-0 | 102110-14-5 | Hydrocarbons, C<sub>3-6</sub>, C<sub>5</sub>-rich, steam-cracked naphtha: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by distillation of steam-cracked naphtha. It consists predominantly of hydrocarbons having carbon numbers in the range of C<sub>3</sub> through C<sub>6</sub>, predominantly C<sub>5</sub>.] |  |  | N
305-975-9 | 95371-08-7 | Hydrocarbons, C<sub>37-65</sub>, hydrotreated deasphalted vacuum distn. residues: Baseoil - unspecified |  |  | N
305-974-3 | 95371-07-6 | Hydrocarbons, C<sub>37-68</sub>, dewaxed deasphalted hydrotreated vacuum distn. residues: Baseoil - unspecified |  |  | N
295-445-2 | 92045-63-1 | Hydrocarbons, C<sub>4-11</sub>, naphtha-cracking, arom.-free: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained from prehydrogenated cracked naphtha after distillative separation of benzene- and toluene-containing hydrocarbon cuts and a higher boiling fraction. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>11</sub> and boiling in the range of approximately 30°C to 205°C (86°F to 401°F).] |  |  | N
295-443-1 | 92045-61-9 | Hydrocarbons, C<sub>4-12</sub>, naphtha-cracking, hydrotreated: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained by distillation from the product of a naphtha steam cracking process and subsequent catalytic selective hydrogenation of gum formers. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>12</sub> and boiling in the range of approximately 30°C to 230°C (86°F to 446°F).] |  |  | N
295-298-4 | 91995-38-9 | Hydrocarbons, C<sub>4-6</sub>, depentanizer lights, arom. hydrotreater: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained as first runnings from the depentanizer column before hydrotreatment of the aromatic charges. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>6</sub>, predominantly pentanes and pentenes, and boiling in the range of approximately 25°C to 40°C (77°F to 104°F).] |  |  | N
310-013-6 | 102110-15-6 | Hydrocarbons, C<sub>5</sub>-rich, dicyclopentadiene-contg.: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by distillation of the products from a steam-cracking process. It consists predominantly of hydrocarbons having carbon numbers of C<sub>5</sub> and dicyclopentadiene and boiling in the range of approximately 30°C to 170°C (86°F to 338°F).] |  |  | N
270-695-5 | 68476-55-1 | Hydrocarbons, C<sub>5</sub>-rich: Low boiling point naphtha - unspecified |  |  | N
297-466-2 | 93572-36-2 | Hydrocarbons, C<sub>5-11</sub>, nonaroms.-rich, reforming light fraction: Low boiling point cat-reformed naphtha: [A complex combination of hydrocarbons obtained by separation from the platformate-containing fraction. It consists predominantly of nonaromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>11</sub> and boiling in the range of approximately 35°C to 125°C (94°F to 257°F), benzene and toluene.] |  |  | N
309-871-4 | 101316-67-0 | Hydrocarbons, C<sub>6</sub>-rich, hydrotreated light naphtha distillates, solvent-refined: Low boiling point modified naphtha: [A complex combination of hydrocarbons obtained by distillation of hydrotreated naphtha followed by solvent extraction. It consists predominantly of saturated hydrocarbons and boiling in the range of approximately 65°C to 70°C (149°F to 158°F).] |  |  | N
297-852-0 | 93763-33-8 | Hydrocarbons, C<sub>6-11</sub>, hydrotreated, dearomatized: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained as solvents which have been subjected to hydrotreatment in order to convert aromatics to naphthenes by catalytic hydrogenation.] |  |  | N
295-446-8 | 92045-64-2 | Hydrocarbons, C<sub>6-7</sub>, naphtha-cracking, solvent-refined: Low boiling point modified naphtha: [A complex combination of hydrocarbons obtained by the sorption of benzene from a catalytically fully hydrogenated benzene-rich hydrocarbon cut that was distillatively obtained from prehydrogenated cracked naphtha. It consists predominantly of paraffinic and naphthenic hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>7</sub> and boiling in the range of approximately 70°C to 100°C (158°F to 212°F).] |  |  | N
309-870-9 | 101316-66-9 | Hydrocarbons, C<sub>6-8</sub>, hydrogenated sorption-dearomatized, toluene raffination: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained during the sorptions of toluene from a hydrocarbon fraction from cracked gasoline treated with hydrogen in the presence of a catalyst. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>8</sub> and boiling in the range of approximately 80°C to 135°C (176°F to 275°F).] |  |  | N
297-465-7 | 93572-35-1 | Hydrocarbons, C<sub>7-12</sub>, C<sub>≥9</sub>-arom.-rich, reforming heavy fraction: Low boiling point cat-reformed naphtha: [A complex combination of hydrocarbons obtained by separation from the platformate-containing fraction. It consists predominantly of nonaromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>12</sub> and boiling in the range of approximately 120°C to 210°C (248°F to 380°F) and C<sub>9</sub> and higher aromatic hydrocarbons.] |  |  | N
295-444-7 | 92045-62-0 | Hydrocarbons, C<sub>8-11</sub>, naphtha-cracking, toluene cut: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by distillation from prehydrogenated cracked naphtha. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>8</sub> through C<sub>11</sub> and boiling in the range of approximately 130°C to 205°C (266°F to 401°F).] |  |  | N
309-974-4 | 101794-97-2 | Hydrocarbons, C<sub>8-12</sub>, catalytic cracker distillates: Low boiling point cat-cracked naphtha: [A complex combination of hydrocarbons obtained by distillation of products from a catalytic cracking process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>8</sub> through C<sub>12</sub> and boiling in the range of approximately 140°C to 210°C (284°F to 410°F).] |  |  | N
309-987-5 | 101896-28-0 | Hydrocarbons, C<sub>8-12</sub>, catalytic cracking, chem. neutralized, sweetened: Low boiling point cat-cracked naphtha |  |  | N
295-794-0 | 92128-94-4 | Hydrocarbons, C<sub>8-12</sub>, catalytic-cracking, chem. neutralized: Low boiling point cat-cracked naphtha: [A complex combination of hydrocarbons produced by the distillation of a cut from the catalytic cracking process, having undergone an alkaline washing. It consists predominantly of hydrocarbons having carbon numbers in the range of C<sub>8</sub> through C<sub>12</sub> and boiling in the range of approximately 130°C to 210°C (266°F to 410°F).] |  |  | N
297-853-6 | 93763-34-9 | Hydrocarbons, C<sub>9-12</sub>, hydrotreated, dearomatized: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained as solvents which have been subjected to hydrotreatment in order to convert aromatics to naphthenes by catalytic hydrogenation.] |  |  | N
271-259-7 | 68527-16-2 | Hydrocarbons, C1-3: Petroleum gas: [A complex combination of hydrocarbons having carbon numbers predominantly in the range of C1 through C3 and boiling in the range of approximately minus 164 °C to minus 42 °C (– 263 °F to – 44 °F).] |  |  | N
271-261-8 | 68527-19-5 | Hydrocarbons, C1-4, debutanizer fraction: Petroleum gas |  |  | N
271-038-5 | 68514-36-3 | Hydrocarbons, C1-4, sweetened: Petroleum gas: [A complex combination of hydrocarbons obtained by subjecting hydrocarbon gases to a sweetening process to convert mercaptans or to remove acidic impurities. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C4 and boiling in the range of approximately – 164 °C to – 0.5 °C (– 263 °F to 31 °F).] |  |  | N
271-032-2 | 68514-31-8 | Hydrocarbons, C1-4: Petroleum gas: [A complex combination of hydrocarbons provided by thermal cracking and absorber operations and by distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C4 and boiling in the range of approximately minus 164 °C to minus 0.5 °C (– 263 °F to 31 °F).] |  |  | N
270-689-2 | 68476-49-3 | Hydrocarbons, C2-4, C3-rich: Petroleum gas |  |  | N
271-734-9 | 68606-25-7 | Hydrocarbons, C2-4: Petroleum gas |  |  | N
271-735-4 | 68606-26-8 | Hydrocarbons, C3: Petroleum gas |  |  | N
270-681-9 | 68476-40-4 | Hydrocarbons, C3-4: Petroleum gas |  |  | N
270-990-9 | 68512-91-4 | Hydrocarbons, C3-4-rich, petroleum distillate: Petroleum gas: [A complex combination of hydrocarbons produced by distillation and condensation of crude oil. It consists of hydrocarbons having carbon numbers in the range of C3 through C5, predominantly C3 through C4.] |  |  | N
306-004-1 | 95465-89-7 | Hydrocarbons, C4, 1,3-butadiene- and isobutene-free: Petroleum gas |  |  | N
295-405-4 | 92045-23-3 | Hydrocarbons, C4, steam-cracker distillate: Petroleum gas: [A complex combination of hydrocarbons produced by the distillation of the products of a steam cracking process. It consists predominantly of hydrocarbons having a carbon number of C4, predominantly 1-butene and 2-butene, containing also butane and isobutene and boiling in the range of approximately minus 12 °C to 5 °C (10.4 °F to 41 °F).] |  |  | N
289-339-5 | 87741-01-3 | Hydrocarbons, C4: Petroleum gas |  |  | N
270-682-4 | 68476-42-6 | Hydrocarbons, C4-5: Petroleum gas |  |  | N
297-857-8 | 93763-38-3 | Hydrocarbons, hydrocracked paraffinic distn. residues, solvent-dewaxed: Baseoil - unspecified |  |  | N
295-436-3 | 92045-55-1 | Hydrocarbons, hydrotreated light naphtha distillates, solvent-refined: Low boiling point modified naphtha: [A combination of hydrocarbons obtained from the distillation of hydrotreated naphtha followed by a solvent extraction and distillation process. It consists predominantly of saturated hydrocarbons boiling in the range of approximately 94°C to 99°C (201°F to 210°F).] |  |  | N
206-019-2 | 288-32-4 | imidazole |  |  | N | Corrosion inhibitor (in water) and process chemical used to make imidazole derivatives
244-959-5 | 22398-80-7 | indium phosphide |  |  | 2 | Semiconductor used in laser diodes and high frequency electronics - no patient contact
200-857-2 | 75-28-5 | isobutane (containing ≥ 0,1 % butadiene (203-450-8)) |  |  | N | Gas
208-819-7 | 542-56-3 | isobutyl nitrite |  |  | N | Medicinal and recreational drug
201-143-3 | 78-79-5 | isoprene (stabilised): 2-methyl-1,3-butadiene |  |  | N | Monomer used to make rubber. Occurs naturally
| 143-50-0 | Kepone (Chlordecone) |  |  | N | Banned (POPs Regulation) insecticide
265-667-4 | 65277-42-1 | ketoconazole: 1-[4-[4-[[(2<i>SR</i>,4<i>RS</i>)-2-(2,4-dichlorophenyl)-2-(imidazol-1-ylmethyl)-1,3-dioxolan-4-yl]methoxy]phenyl]piperazin-1-yl]ethanone |  |  | N | Antifungal drug used primarily to treat fungal infections.
239-290-0 | 15245-44-0 | lead 2,4,6-trinitro-<i>m</i>-phenylene dioxide: lead 2,4,6-trinitroresorcinoxide: lead styphnate |  |  | N | Detonator
215-630-3 | 1335-32-6 | lead acetate, basic |  |  | 2 | MRI/CT contrast agent, but should be no patient exposure, unless leaks occur.
231-846-0 | 7758-97-6 | lead chromate |  |  | 1 | Pigment (Not RoHS compliant)
235-759-9 | 12656-85-8 | lead chromate molybdate sulfate red: C.I. Pigment Red 104: [This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77605.] |  |  | 1 | Pigment (Not RoHS compliant)
206-104-4 | 301-04-2, 6080-56-4 | lead di(acetate) |  |  | 2 | MRI/CT contrast agent, but should be no patient exposure, unless leaks occur.
236-542-1 | 13424-46-9 | lead diazide: lead azide |  |  | N | Detonator
247-278-1 | 25808-74-6 | lead hexafluorosilicate |  |  | N | Lead refining
232-064-2 | 7784-40-9 | lead hydrogen arsenate |  |  | N | Insecticide (now banned) and an SVHC
215-693-7 | 1344-37-2 | lead sulfochromate yellow: C.I. Pigment Yellow 34: [This substance is identified in the Colour Index by Colour Index Constitution Number, C.I. 77603.] |  |  | 1 | Pigment (Not RoHS compliant)
401-750-5 | 17570-76-2 | lead(II) methanesulphonate |  |  | N | Electroplating chemical
266-012-5 | 65996-78-3 | Light oil (coal), coke-oven: Crude benzole: [The volatile organic liquid extracted from the gas evolved in the high temperature (greater than 700°C (1292°F)) destructive distillation of coal. Composed primarily of benzene, toluene, and xylenes. May contain other minor hydrocarbon constituents.] |  |  | N
292-635-7 | 90641-11-5 | Light oil (coal), semi-coking process: Fresh oil: [The volatile organic liquid condensed from the gas evolved in the low-temperature (less than 700°C (1292°F)) destructive distillation of coal. Composed primarily of C<sub>6-10</sub> hydrocarbons.] |  |  | N
232-453-7 | 8032-32-4 | Ligroine: Low boiling point naphtha: [A complex combination of hydrocarbons obtained by the fractional distillation of petroleum. This fraction boils in a range of approximately 20°C to 135°C (58°F to 275°F).] |  |  | N
206-356-5 | 330-55-2 | linuron (ISO): 3-(3,4-dichlorophenyl)-1-methoxy-1-methylurea |  |  | N | Agricultural herbicide
620-400-4 | 12031-65-1 | lithium nickel dioxide |  |  | 2 | Lithium ion battery anode
249-644-6 | 29457-72-5 | lithium perfluorooctane sulfonate: lithium heptadecafluorooctanesulfonate |  |  | N | Insecticide and laboratory chemical
278-011-7 | 74869-21-9 | Lubricating greases: Grease: [A complex combination of hydrocarbons having carbon numbers predominantly in the range of C<sub>12</sub> through C<sub>50.</sub> May contain organic salts of alkali metals, alkaline earth metals, and/or aluminium compounds.] |  |  | 3 | May be used in medical devices although unlikely to be patient contact
297-474-6 | 93572-43-1 | Lubricating oils (petroleum), base oils, paraffinic: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by refining of crude oil. It consists predominantly of aromatics, naphthenics and paraffinics and produces a finished oil with a viscosity of 120 SUS at 100 °F (23cSt at 40 °C).] |  |  | 3 | Lubricants
309-874-0 | 101316-69-2 | Lubricating oils (petroleum), C <sub>>25</sub>, solvent-extd., deasphalted, dewaxed, hydrogenated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by solvent extraction and hydrogenation of vacuum distillation residues. It consists predominantly of hydrocarbons having carbon numbers predominantly greater than C<sub>25</sub> and produces a finished oil with a viscosity in the order of 32cSt to 37cSt at 100 °C (212 °F).] |  |  | 3 | Lubricants
276-737-9 | 72623-86-0 | Lubricating oils (petroleum), C<sub>15-30</sub>, hydrotreated neutral oil-based: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treating light vacuum gas oil and heavy vacuum gas oil with hydrogen in the presence of a catalyst in a two stage process with dewaxing being carried out between the two stages. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil having a viscosity of approximately 15cSt at 40 °C. It contains a relatively large proportion of saturated hydrocabons.] |  |  | 3 | Lubricants
309-875-6 | 101316-70-5 | Lubricating oils (petroleum), C<sub>17-32</sub>, solvent-extd., dewaxed, hydrogenated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by solvent extraction and hydrogenation of atmospheric distillation residues. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>17</sub> through C<sub>32</sub> and produced a finished oil with a viscosity in the order of 17cSt to 23cSt at 40 °C (104 °F.] |  |  | 3 | Lubricants
295-423-2 | 92045-42-6 | Lubricating oils (petroleum), C<sub>17-35</sub>, solvent-extd., dewaxed, hydrotreated: Baseoil - unspecified |  |  | 3 | Lubricants
307-034-8 | 97488-95-4 | Lubricating oils (petroleum), C<sub>18-27</sub>, hydrocracked solvent-dewaxed: Baseoil - unspecified |  |  | 3 | Lubricants
305-594-8 | 94733-15-0 | Lubricating oils (petroleum), C<sub>18-40</sub>, solvent-dewaxed hydrocracked distillate-based: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by solvent deparaffination of the distillation residue from hydrocracked petroleum. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>18</sub> through C<sub>40</sub> and boiling in the range of approximately 370 °C to 550 °C (698 °F to 1022 °F).] |  |  | 3 | Lubricants
305-595-3 | 94733-16-1 | Lubricating oils (petroleum), C<sub>18-40</sub>, solvent-dewaxed hydrogenated raffinate-based: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by solvent deparaffination of the hydrogenated raffinate obtained by solvent extraction of a hydrotreated petroleum distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>18</sub> through C<sub>40</sub> and boiling in the range of approximately 370 °C to 550 °C (698 °F to 1022 °F).] |  |  | 3 | Lubricants
309-876-1 | 101316-71-6 | Lubricating oils (petroleum), C<sub>20-35</sub>, solvent-extd., dewaxed, hydrogenated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by solvent extraction and hydrogenation of atmospheric distillation residues. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>35</sub> and produces a finished oil with a viscosity in the order of 37cSt to 44cSt at 40 °C (104 °F).] |  |  | 3 | Lubricants
276-736-3 | 72623-85-9 | Lubricating oils (petroleum), C<sub>20-50</sub>, hydrotreated neutral oil-based, high-viscosity: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treating light vacuum gas oil, heavy vacuum gas oil, and solvent deasphalted residual oil with hydrogen in the presence of a catalyst in a two stage process with dewaxing being carried out between the two stages. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil having a viscosity of approximately 112cSt at 40 °C. It contains a relatively large proportion of saturated hydrocarbons.] |  |  | 3 | Lubricants
276-738-4 | 72623-87-1 | Lubricating oils (petroleum), C<sub>20-50</sub>, hydrotreated neutral oil-based: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treating light vacuum gas oil, heavy vacuum gas oil and solvent deasphalted residual oil with hydrogen in the presence of a catalyst in a two stage process with dewaxing being carried out between the two stages. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of approximately 32cSt at 40 °C. It contains a relatively large proportion of saturated hydrocarbons.] |  |  | 3 | Lubricants
309-877-7 | 101316-72-7 | Lubricating oils (petroleum), C<sub>24-50</sub>, solvent-extd., dewaxed, hydrogenated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by solvent extraction and hydrogenation of atmospheric distillation residues. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>24</sub> through C<sub>50</sub> and produces a finished oil with a viscosity in the order of 16cSt to 75cSt at 40 °C (104 °F).] |  |  | 3 | Lubricants
295-424-8 | 92045-43-7 | Lubricating oils (petroleum), hydrocracked nonarom. solvent-deparaffined: Baseoil - unspecified |  |  | 3 | Lubricants
278-012-2 | 74869-22-0 | Lubricating oils: Baseoil - unspecified: [A complex combination of hydrocarbons obtained from solvent extraction and dewaxing processes. It consists predominantly of saturated hydrocarbons having carbon numbers in the range C<sub>15</sub> through C<sub>50</sub>.] |  |  | 3 | Lubricants
| 12427-38-2 | Maneb |  |  | N | Fungicide
231-106-7 | 7439-97-6 | mercury |  |  | 2 | Measurement instruments such as sphygmomanometers, although most uses are now banned
| 137-42-8 | Metam Natrium |  |  | N | Fungicide
210-894-6 | 625-45-6 | methoxyacetic acid |  |  | 1 | Solvent for resins and paints. Also used as a plasticiser, in colouring agents and for textile treatment
403-230-3 | 77402-05-2 | methyl acrylamidoglycolate (containing ≥ 0,1 % acrylamide) |  |  | N | Used to make polymers
401-890-7 | 77402-03-0 | methyl acrylamidomethoxyacetate (containing ≥ 0,1 % acrylamid) |  |  | N | Used to make polymers
209-765-7 | 592-62-1 | methyl-<i>O</i>NN-azoxymethyl acetate: methyl azoxy methyl acetate |  |  | N | Pharmaceutical manufacture
200-471-4 | 60-34-4 | methylhydrazine |  |  | N | Rocket fuel
| 1314-04-1 | millerite |  |  | N | Nickel sulphide mineral
| 2385-85-5 | Mirex |  |  | N | Insecticide
268-585-7 | 68130-36-9 | molybdenum nickel hydroxide oxide phosphate |  |  | N | Process catalyst
| 12673-58-4 | molybdenum nickel oxide |  |  | N | Process catalyst
238-034-5 | 14177-55-0 | molybdenum nickel tetraoxide |  |  | N | Process catalyst
227-062-3 | 5625-90-1 | N,N′-methylenedimorpholine: N,N′-methylenebismorpholine: [formaldehyde released from N,N′-methylenebismorpholine]: [MBM] |  |  | N | Biocide
292-636-2 | 90641-12-6 | Naphtha (coal), distn. residues: Light Oil Redistillate, high boiling: [The residue remaining from the distillation of recovered naphtha. Composed primarily of naphthalene and condensation products of indene and styrene.] |  |  | N
302-690-1 | 94114-54-2 | Naphtha (coal), solvent extn., hydrocracked: [Fraction of the distillate obtained by hydrocracking of coal extract or solution produced by the liquid solvent extraction or supercritical gas extraction processes and boiling in the range of approximately 30°C to 180°C (86°F to 356°F). Composed primarily of aromatic, hydrogenated aromatic and naphthenic compounds, their alkyl derivatives and alkanes with carbon numbers predominantly in the range of C<sub>4</sub> to C<sub>9</sub>. Nitrogen, sulfur and oxygen-containing aromatic and hydrogenated aromatic compounds are also present.] |  |  | N
265-115-2 | 64742-15-0 | Naphtha (petroleum), acid-treated: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained as a raffinate from a sulfuric acid treating process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>12</sub> and boiling in the range of approximately 90°C to 230°C (194°F to 446°F).] |  |  | N
271-635-0 | 68603-08-7 | Naphtha (petroleum), arom.-contg.: Low boiling point naphtha - unspecified |  |  | N
295-430-0 | 92045-49-3 | Naphtha (petroleum), C<sub>4-12</sub>, butane-alkylate, isooctane-rich: Low boiling point modified naphtha: [A complex combination of hydrocarbons obtained by alkylation of butanes. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>12</sub>, rich in isooctane, and boiling in the range of approximately 35°C to 210°C (95°F to 410°F).] |  |  | N
272-185-8 | 68783-09-5 | Naphtha (petroleum), catalytic cracked light distd.: Low boiling point cat-cracked naphtha: [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>1</sub> through C<sub>5</sub>.] |  |  | N
265-170-2 | 64742-66-1 | Naphtha (petroleum), catalytic dewaxed: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained from the catalytic dewaxing of a petroleum fraction. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>12</sub> and boiling in the range of approximately 35°C to 230°C (95°F to 446°F).] |  |  | N
285-510-3 | 85116-59-2 | Naphtha (petroleum), catalytic reformed light, arom.-free fraction: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons remaining after removal of aromatic compounds from catalytic reformed light naphtha in a selective absorption process. It consists predominantly of paraffinic and cyclic compounds having carbon numbers predominantly in the range of C<sub>5</sub> to C<sub>8</sub> and boiling in the range of approximately 66°C to 121°C (151°F to 250°F).] |  |  | N
273-271-8 | 68955-35-1 | Naphtha (petroleum), catalytic reformed: Low boiling point cat-reformed naphtha: [A complex combination of hydrocarbons produced by the distillation of products from a catalytic reforming process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>12</sub> and boiling in the range of approximately 30°C to 220°C (90°F to 430°F). It contains a relatively large proportion of aromatic and branched chain hydrocarbons. This stream may contain 10 vol. % or more benzene.] |  |  | N
265-122-0 | 64742-22-9 | Naphtha (petroleum), chemically neutralized heavy: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>12</sub> and boiling in the range of approximately 65°C to 230°C (149°F to 446°F).] |  |  | N
265-123-6 | 64742-23-0 | Naphtha (petroleum), chemically neutralized light: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons produced by a treating process to remove acidic materials. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>11</sub> and boiling in the range of approximately -20°C to 190°C (-4°F to 374°F).] |  |  | N
271-262-3 | 68527-21-9 | Naphtha (petroleum), clay-treated full-range straight-run: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons resulting from treatment of full-range straight-run naphtha with natural or modified clay, usually in a percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>11</sub> and boiling in the range of approximately -20°C to 220°C (-4°F to 429°F).] |  |  | N
271-263-9 | 68527-22-0 | Naphtha (petroleum), clay-treated light straight-run: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons resulting from treatment of light straight-run naphtha with a natural or modified clay, usually in a percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>10</sub> and boiling in the range of approximately 93°C to 180°C (200°F to 356°F).] |  |  | N
271-267-0 | 68527-27-5 | Naphtha (petroleum), full-range alkylate, butane-contg.: Low boiling point modified naphta: [A complex combination of hydrocarbons produced by the distillation of the reaction products of isobutane with monoolefinic hydrocarbons usually ranging in carbon numbers from C<sub>3</sub> through C<sub>5</sub>. It consists of predominantly branched chain saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>12</sub> with some butanes and boiling in the range of approximately 35°C to 200°C (95°F to 428°F).] |  |  | N
265-066-7 | 64741-64-6 | Naphtha (petroleum), full-range alkylate: Low boiling point modified naphtha: [A complex combination of hydrocarbons produced by distillation of the reaction products of isobutane with monoolefinic hydrocarbons usually ranging in carbon numbers from C<sub>3</sub> through C<sub>5</sub> It consists of predominantly branched chain saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>12</sub> and boiling in the range of approximately 90°C to 220°C (194°F to 428°F).] |  |  | N
270-991-4 | 68513-02-0 | Naphtha (petroleum), full-range coker: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons produced by the distillation of products from a fluid coker. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>15</sub> and boiling in the range of approximately 43°C to 250°C (110°F-500°F).] |  |  | N
272-895-8 | 68919-37-9 | Naphtha (petroleum), full-range reformed: Low boiling point cat-reformed naphtha: [A complex combination of hydrocarbons produced by the distillation of the products from a catalytic reforming process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>12</sub> and boiling in the range of approximately 35°C to 230°C (95°F to 446°F).] |  |  | N
265-042-6 | 64741-42-0 | Naphtha (petroleum), full-range straight-run: Low boiling point naphtha: [A complex combination of hydrocarbons produced by distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>11</sub> and boiling in the range of approximately -20°C to 220°C (-4°F to 428°F).] |  |  | N
265-067-2 | 64741-65-7 | Naphtha (petroleum), heavy alkylate: Low boiling point modified naphtha: [A complex combination of hydrocarbons produced by distillation of the reaction products of isobutane with monoolefinic hydrocarbons usually ranging in carbon numbers from C<sub>3</sub> to C<sub>5</sub>. It consists of predominantly branched chain saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>9</sub> through C<sub>12</sub> and boiling in the range of approximately 150°C to 220°C (302°F to 428°F).] |  |  | N
295-431-6 | 92045-50-6 | Naphtha (petroleum), heavy catalytic cracked, sweetened: Low boiling point cat-cracked naphtha: [A complex combination of hydrocarbons obtained by subjecting a catalytic cracked petroleum distillate to a sweetening process to convert mercaptans or to remove acidic impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>12</sub> and boiling in the range of approximately 60°C to 200°C (140°F to 392°F).] |  |  | N
265-055-7 | 64741-54-4 | Naphtha (petroleum), heavy catalytic cracked: Low boiling point cat-cracked naphtha: [A complex combination of hydrocarbons produced by a distillation of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>12</sub> and boiling in the range of approximately 65°C to 230°C (148°F to 446°F). It contains a relatively large proportion of unsaturated hydrocarbons.] |  |  | N
265-070-9 | 64741-68-0 | Naphtha (petroleum), heavy catalytic reformed: Low boiling point cat-reformed naphtha: [A complex combination of hydrocarbons produced from the distillation of products from a catalytic reforming process. It consists of predominantly aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>12</sub> and boiling in the range of approximately 90°C to 230°C (194°F to 446°F).] |  |  | N
265-079-8 | 64741-78-2 | Naphtha (petroleum), heavy hydrocracked: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons from distillation of the products from a hydrocracking process. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>12</sub>, and boiling in the range of approximately 65°C to 230°C (148°F to 446°F).] |  |  | N
295-432-1 | 92045-51-7 | Naphtha (petroleum), heavy steam-cracked, hydrogenated: Low boiling point hydrogen treated naphtha |  |  | N
309-945-6 | 101631-20-3 | Naphtha (petroleum), heavy straight run, arom.-contg.: Low boiling point naphtha: [A complex combination of hydrocarbons obtained from a distillation process of crude petroleum. It consists predominantly of hydrocarbons having carbon numbers in the range of C<sub>8</sub> through C<sub>12</sub> and boiling in the range of approximately 130°C to 210°C (266°F to 410°F).] |  |  | N
265-041-0 | 64741-41-9 | Naphtha (petroleum), heavy straight-run: Low boiling point naphtha: [A complex combination of hydrocarbons produced by distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>12</sub> and boiling in the range of approximately 65°C to 230°C (149°F to 446°F).] |  |  | N
265-085-0 | 64741-83-9 | Naphtha (petroleum), heavy thermal cracked: Low boiling point thermally cracked naphtha: [A complex combination of hydrocarbons from distillation of the products from a thermal cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>12</sub> and boiling in the range of approximately 65°C to 220°C (148°F to 428°F).] |  |  | N
309-879-8 | 101316-76-1 | Naphtha (petroleum), hydrodesulfurised full-range coker: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by fractionation from hydrodesulfurised coker distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> to C<sub>11</sub> and boiling in the range of approximately 23°C to 196°C (73°F to 385°F).] |  |  | N
295-433-7 | 92045-52-8 | Naphtha (petroleum), hydrodesulfurized full-range: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained from a catalytic hydrodesulfurization process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>11</sub> and boiling in the range of approximately 30°C to 250°C (86°F to 482°F).] |  |  | N
295-434-2 | 92045-53-9 | Naphtha (petroleum), hydrodesulfurized light, dearomatized: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by distillation of hydrodesulfurized and dearomatized light petroleum fractions. It consists predominantly of C<sub>7</sub> paraffins and cycloparaffins boiling in a range of approximately 90°C to 100°C (194°F to 212°F).] |  |  | N
265-178-6 | 64742-73-0 | Naphtha (petroleum), hydrodesulfurized light: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained from a catalytic hydrodesulfurization process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>11</sub> and boiling in the range of approximately -20°C to 190°C (-4°F to 374°F).] |  |  | N
285-511-9 | 85116-60-5 | Naphtha (petroleum), hydrodesulfurized thermal cracked light: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained by fractionation of hydrodesulfurized thermal cracker distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> to C<sub>11</sub> and boiling in the range of approximately 23°C to 195°C (73°F to 383°F).] |  |  | N
265-185-4 | 64742-82-1 | naphtha (petroleum), hydrodesulphurized heavy: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained from a catalytic hydrodesulfurization process. It consists of hydrocarbons having carbon numbers predominantly in the range of C7 through C12 and boiling in the range of approximately 90 °C to 230 °C (194 °F to 446 °F).] |  |  | N
265-150-3 | 64742-48-9 | Naphtha (petroleum), hydrotreated heavy: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>13</sub> and boiling in the range of approximately 65°C to 230°C (149°F to 446°F).] |  |  | N | Solvent for paint, but will not be present in medical devices
295-438-4 | 92045-57-3 | Naphtha (petroleum), hydrotreated light steam-cracked: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained by treating a petroleum fraction, derived from a pyrolysis process, with hydrogen in the presence of a catalyst. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>11</sub> and boiling in the range of approximately 35°C to 190°C (95°F to 374°F).] |  |  | N
285-512-4 | 85116-61-6 | Naphtha (petroleum), hydrotreated light, cycloalkane-contg.: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained from the distillation of a petroleum fraction. It consists predominantly of alkanes and cycloalkanes boiling in the range of approximately -20°C to 190°C (-4°F to 374°F).] |  |  | N
265-151-9 | 64742-49-0 | Naphtha (petroleum), hydrotreated light: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>11</sub> and boiling in the range of approximately minus 20°C to 190°C (-4°F to 374°F).] |  |  | N | Solvent used for thinning oil-based paint, enamel & varnish. Will not occur in medical devices
295-440-5 | 92045-58-4 | Naphtha (petroleum), isomerization, C<sub>6</sub>-fraction: Low boiling point modified naphtha: [A complex combination of hydrocarbons obtained by distillation of a gasoline which has been catalytically isomerized. It consists predominantly of hexane isomers boiling in the range of approximately 60°C to 66°C (140°F to 151°F).] |  |  | N
265-073-5 | 64741-70-4 | Naphtha (petroleum), isomerization: Low boiling point modified naphtha: [A complex combination of hydrocarbons obtained from catalytic isomerization of straight chain paraffinic C<sub>4</sub> through C<sub>6</sub> hydrocarbons. It consists predominantly of saturated hydrocarbons such as isobutane, isopentane, 2,2-dimethylbutane, 2-methylpentane, and 3-methylpentane.] |  |  | N
265-068-8 | 64741-66-8 | Naphtha (petroleum), light alkylate: Low boiling point modified naphtha: [A complex combination of hydrocarbons produced by distillation of the reaction products of isobutane with monoolefinic hydrocarbons usually ranging in carbon numbers from C<sub>3</sub> through C<sub>5</sub>. It consists of predominantly branched chain saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>10</sub> and boiling in the range of approximately 90°C to 160°C (194°F to 320°F).] |  |  | N
295-441-0 | 92045-59-5 | Naphtha (petroleum), light catalytic cracked sweetened: Low boiling point cat-cracked naphtha: [A complex combination of hydrocarbons obtained by subjecting naphtha from a catalytic cracking process to a sweetening process to convert mercaptans or to remove acidic impurities. It consists predominantly of hydrocarbons boiling in a range of approximately 35°C to 210°C (95°F to 410°F).] |  |  | N
265-056-2 | 64741-55-5 | Naphtha (petroleum), light catalytic cracked: Low boiling point cat-cracked naphtha: [A complex combination of hydrocarbons produced by the distillation of products from a catalytic cracking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>11</sub> and boiling in the range of approximately -20°C to 190°C (-4°F to 374°F). It contains a relatively large proportion of unsaturated hydrocarbons.] |  |  | N
270-993-5 | 68513-03-1 | Naphtha (petroleum), light catalytic reformed, arom.-free: Low boiling point cat-reformed naphtha: [A complex combination of hydrocarbons obtained from distillation of products from a catalytic reforming process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>8</sub> and boiling in the range of approximately 35°C to 120°C (95°F to 248°F). It contains a relatively large proportion of branched chain hydrocarbons with the aromatic components removed.] |  |  | N
265-065-1 | 64741-63-5 | Naphtha (petroleum), light catalytic reformed: Low boiling point cat-reformed naphtha: [A complex combination of hydrocarbons produced from the distillation of products from a catalytic reforming process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>11</sub> and boiling in the range of approximately 35°C to 190°C (95°F to 374°F). It contains a relatively large proportion of aromatic and branched chain hydrocarbons. This stream may contain 10 vol. % or more benzene.] |  |  | N
296-028-8 | 92201-97-3 | Naphtha (petroleum), light heat-soaked, steam-cracked: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by the fractionation of steam cracked naphtha after recovery from a heat soaking process. It consists predominantly of hydrocarbons having a carbon number predominantly in the range of C<sub>4</sub> through C<sub>6</sub> and boiling in the range of approximately 0°C to 80°C (32°F to 176°F).] |  |  | N
265-071-4 | 64741-69-1 | Naphtha (petroleum), light hydrocracked: Low boiling naphtha - unspecified: [A complex combination of hydrocarbons from distillation of the products from a hydrocracking process. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>10</sub>, and boiling in the range of approximately -20°C to 180°C (-4°F to 356°F).] |  |  | N
271-264-4 | 68527-23-1 | Naphtha (petroleum), light steam-cracked arom.: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons produced by distillation of products from a steam-cracking process. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>9</sub> and boiling in the range of approximately 110°C to 165°C (230°F to 329°F).] |  |  | N
308-713-1 | 98219-46-6 | Naphtha (petroleum), light steam-cracked, debenzenized, thermally treated: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by the treatment and distillation of debenzenized light steam-cracked petroleum naphtha. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>12</sub> and boiling in the range of approximately 95°C to 200°C (203°F to 392°F).] |  |  | N
271-266-5 | 68527-26-4 | Naphtha (petroleum), light steam-cracked, debenzenized: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons produced by distillation of products from a steam-cracking process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>12</sub> and boiling in the range of approximately 80°C to 218°C (176°F to 424°F).] |  |  | N
296-942-7 | 93165-55-0 | Naphtha (petroleum), light steam-cracked, hydrogenated: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons produced from the separation and subsequent hydrogenation of the products of a steam-cracking process to produce ethylene. It consists predominantly of saturated and unsaturated paraffins, cyclic paraffins and cyclic aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>10</sub> and boiling in the range of approximately 50°C to 200°C (122°F to 392°F). The proportion of benzene hydrocarbons may vary up to 30 wt. % and the stream may also contain small amounts of sulfur and oxygenated compounds.] |  |  | N
308-714-7 | 98219-47-7 | Naphtha (petroleum), light steam-cracked, thermally treated: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by the treatment and distillation of light steam-cracked petroleum naphtha. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>6</sub> and boiling in the range of approximately 35°C to 80°C (95°F to 176°F).] |  |  | N
265-187-5 | 64742-83-2 | Naphtha (petroleum), light steam-cracked: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by the distillation of the products from a steam cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>11</sub> and boiling in the range of approximately minus 20°C to 190°C (-4°F to 374°F). This stream is likely to contain 10 vol. % or more benzene.] |  |  | N
265-046-8 | 64741-46-4 | Naphtha (petroleum), light straight-run: Low boiling point naphtha: [A complex combination of hydrocarbons produced by distillation of crude oil. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>10</sub> and boiling in the range of approximately -20°C to 180°C (-4°F to 356°F).] |  |  | N
295-447-3 | 92045-65-3 | Naphtha (petroleum), light thermal cracked, sweetened: Low boiling point thermally cracked naphtha: [A complex combination of hydrocarbons obtained by subjecting a petroleum distillate from the high temperature thermal cracking of heavy oil fractions to a sweetening process to convert mercaptans. It consists predominantly of aromatics, olefins and saturated hydrocarbons boiling in the range of approximately 20°C to 100°C (68°F to 212°F).] |  |  | N
265-075-6 | 64741-74-8 | Naphtha (petroleum), light thermal cracked: Low boiling point thermally cracked naphtha: [A complex combination of hydrocarbons from distillation of products from a thermal cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>8</sub> and boiling in the range of approximately -10 °C to 130 °C (14 °F to 266 °F).] |  |  | N
295-442-6 | 92045-60-8 | Naphtha (petroleum), light, C<sub>5</sub>-rich, sweetened: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by subjecting a petroleum naphtha to a sweetening process to convert mercaptans or to remove acidic impurities. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>5</sub>, predominantly C<sub>5</sub>, and boiling in the range of approximately minus 10°C to 35°C (14°F to 95°F).] |  |  | N
272-206-0 | 68783-66-4 | Naphtha (petroleum), light, sweetened: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by subjecting a petroleum distillate to a sweetening process to convert mercaptans or to remove acidic impurities. It consists predominantly of saturated and unsaturated hydrocarbons having carbon numbers predominantly in the range of C<sub>3</sub> through C<sub>6</sub> and boiling in the range of approximately -20°C to 100°C (-4°F to 212°F).] |  |  | N
265-095-5 | 64741-92-0 | Naphtha (petroleum), solvent-refined heavy: Low boiling point modified naphtha: [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>12</sub> and boiling in the range of approximately 90°C to 230°C (194°F to 446°F).] |  |  | N
307-035-3 | 97488-96-5 | Naphtha (petroleum), solvent-refined hydrodesulfurized heavy: Gasoil - unspecified |  |  | N
265-086-6 | 64741-84-0 | Naphtha (petroleum), solvent-refined light: Low boiling point modified naphtha: [A complex combination of hydrocarbons obtained as the raffinate from a solvent extraction process. It consists predominantly of aliphatic hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>11</sub> and boiling in the range of approximately 35°C to 190°C (95°F to 374°F).] |  |  | N
271-138-9 | 68516-20-1 | Naphtha (petroleum), steam-cracked middle arom.: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons produced by the distillation of products from a steam-cracking process. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>7</sub> through C<sub>12</sub> and boiling in the range of approximately 130°C to 220°C (266°F to 428°F).] |  |  | N
309-976-5 | 101795-01-1 | Naphtha (petroleum), sweetened light: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by subjecting a petroleum naphtha to a sweetening process to convert mercaptans or to remove acidic impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>8</sub> and boiling in the range of approximately 20°C to 130°C (68°F to 266°F).] |  |  | N
265-089-2 | 64741-87-3 | Naphtha (petroleum), sweetened: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by subjecting a petroleum naphtha to a sweetening process to convert mercaptans or to remove acidic impurities. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>12</sub> and boiling in the range of approximately -10°C to 230°C (14°F to 446°F).] |  |  | N
272-186-3 | 68783-12-0 | Naphtha (petroleum), unsweetened: Low boiling point naphtha: [A complex combination of hydrocarbons produced from the distillation of naphtha streams from various refinery processes. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>12</sub> and boiling in the range of approximately 0°C to 230°C (25°F to 446°F).] |  |  | N
232-443-2 | 8030-30-6 | Naphtha: Low boiling point naphtha: [Refined, partly refined, or unrefined petroleum products produced by the distillation of natural gas. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>6</sub> and boiling in the range of approximately 100°C to 200°C (212°F to 392°F).] |  |  | N | Component of fuels
265-172-3 | 64742-68-3 | Naphthenic oils (petroleum), catalytic dewaxed heavy: Baseoil - unspecified: [A complex combination of hydrocarbons obtained from a catalytic dewaxing process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-173-9 | 64742-69-4 | Naphthenic oils (petroleum), catalytic dewaxed light: Baseoil - unspecified: [A complex combination of hydrocarbons obtained from a catalytic dewaxing process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-179-1 | 64742-75-2 | Naphthenic oils (petroleum), complex dewaxed heavy: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by removing straight chain paraffin hydrocarbons as a solid by treatment with an agent such as urea. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil having a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-180-7 | 64742-76-3 | Naphthenic oils (petroleum), complex dewaxed light: Baseoil - unspecified: [A complex combination of hydrocarbons obtained from a catalytic dewaxing process. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil having a viscosity less than 100 SUS at 100 °F (19cSt at 40 °C). It contains relatively few normal paraffins.] |  |  | N
265-048-9 | 64741-48-6 | Natural gas (petroleum), raw liq. mix: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons separated as a liquid from natural gas in a gas recycling plant by processes such as refrigeration or absorption. It consists mainly of saturated aliphatic hydrocarbons having carbon numbers in the range of C<sub>2</sub> through C<sub>8</sub>.] |  |  | N
265-047-3 | 64741-47-5 | Natural gas condensates (petroleum): Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons separated as a liquid from natural gas in a surface separator by retrograde condensation. It consists mainly of hydrocarbons having carbon numbers predominantly in the range of C<sub>2</sub> to C<sub>20</sub>. It is a liquid at atmospheric temperature and pressure.] |  |  | N
272-896-3 | 68919-39-1 | Natural gas condensates: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons separated and/or condensed from natural gas during transportation and collected at the wellhead and/or from the production, gathering, transmission, and distribution pipelines in deeps, scrubbers, etc. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>2</sub> through C<sub>8</sub>.] |  |  | N
257-447-1 | 51818-56-5 | neodecanoic acid, nickel salt |  |  | N | Catalyst used to make rubber but likely to be <0.1%
220-250-6 | 2687-91-4 | N-ethyl-2-pyrrolidone: 1-ethylpyrrolidin-2-one |  |  | 2 | Solvent in photoresist stripper formulations. Ethylpyrrolidone can be used in electrolyte formulations for electrolytic capacitors (according to Eastman who are a manufacturer), ink-jet printing ink solvent, liquid crystal cells, solvent and in pharmaceuticals.
240-841-2 | 16812-54-7 | nickel (II) sulfide |  |  | N | Chemical intermediate
258-051-1 | 52625-25-9 | nickel 3,5-bis(<i>tert</i>-butyl)-4-hydroxybenzoate (1:2) |  |  | 1 | Rubber adhesives
239-086-1 | 14998-37-9 | nickel acetate |  |  | 1 | Light stabilisers in PP and PE - impart a green colour. Dye mordant for textiles, and as a sealer for anodized aluminium.
248-169-1 | 27016-75-7 | nickel arsenide |  |  | N | Chemical intermediate
271-853-6 | 68610-24-2 | nickel barium titanium primrose priderite: C.I. Pigment Yellow 157: C.I. 77900 |  |  | 1 | Pigment
224-699-9 | 4454-16-4 | nickel bis(2-ethylhexanoate) |  |  | 2 | Catalyst used to make rubber but likely to be <0.1%
223-463-2 | 3906-55-6 | nickel bis(4-cyclohexylbutyrate) |  |  | N | Catalyst
254-642-3 | 39819-65-3 | nickel bis(benzenesulfonate) |  |  | N | Catalyst
242-522-3 | 18718-11-1 | nickel bis(dihydrogen phosphate) |  |  | N | Catalyst
284-349-6 | 84852-37-9 | nickel bis(isononanoate) |  |  | N | Catalyst
238-511-8 | 14507-36-9 | nickel bis(phosphinate) |  |  | N | Catalyst
237-396-1 | 13770-89-3 | nickel bis(sulfamidate): nickel sulfamate |  |  | N | Process chemical
238-753-4 | 14708-14-6 | nickel bis(tetrafluoroborate) |  |  | N | Process chemical
235-723-2 | 12619-90-8 | nickel boride |  |  | 1 | Used as a process chemical but also used as a hard coating (to replace chrome)
234-493-0 | 12007-00-0 | nickel boride (NiB) |  |  | 1 | Used as a process chemical but also used as a hard coating (to replace chrome)
| 65229-23-4 | nickel boron phosphide |  |  | N | Process chemical
222-068-2 | 3333-67-3 | nickel carbonate: basic nickel carbonate: carbonic acid, nickel (2+) salt |  |  | 1 | Process chemical, industrial catalyst, pigment, mainly used in ceramics
238-766-5 | 14721-18-7 | nickel chromate |  |  | N | Process chemical
206-761-7 | 373-02-4, 6018-89-9 | nickel di(acetate) |  |  | 1 | Anodised aluminium sealing. Up to 1 - 2% in thin anodised coating
235-103-1 | 12068-61-0 | nickel diarsenide |  |  | N | Process chemical
209-046-8 | 553-71-9 | nickel dibenzoate |  |  | N | Catalyst
238-596-1 | 14550-87-9 | nickel dibromate |  |  | N | Process chemical
236-665-0 | 13462-88-9 | nickel dibromide |  |  | N | Process chemical
267-897-0 | 67952-43-6 | nickel dichlorate |  |  | N | Process chemical
231-743-0 | 7718-54-9, 7791-20-0 | nickel dichloride |  |  | N | Process chemical
239-646-5 | 15586-38-6 | nickel dichromate |  |  | N | Process chemical
209-160-8 | 557-19-7 | nickel dicyanide |  |  | N | Process chemical
233-071-3 | 10028-18-9 | nickel difluoride |  |  | N | Process chemical
222-101-0 | 3349-06-2 | nickel diformate |  |  | N | Process chemical
235-008-5 | 12054-48-7 | nickel dihydroxide |  |  | N | NiCd batteries
236-666-6 | 13462-90-3 | nickel diiodide |  |  | N | Process chemical
| 16039-61-5 | nickel dilactate |  |  | N | Process chemical
236-068-5 | 13138-45-9, 13478-00-7 | nickel dinitrate |  |  | N | Process chemical, industrial catalyst
234-823-3 | 12035-36-8 | nickel dioxide |  |  | 1 | Pigment used mainly in glass, ceramics, porcelain enamels and in catalysts. Used to make ferrites (but not present as NiO).
237-124-1 | 13637-71-3 | nickel diperchlorate: perchloric acid, nickel(II) salt |  |  | N | Process chemical
237-563-9 | 13842-46-1 | nickel dipotassium bis(sulfate) |  |  | N | Process chemical
235-379-3 | 12201-89-7 | nickel disilicide |  |  | 2 | Semiconductor memory and other devices
237-205-1 | 13689-92-4 | nickel dithiocyanate |  |  | N | Process chemical
257-970-5 | 52502-12-2 | nickel divanadium hexaoxide |  |  | 2 | Electrochromic material. Also used as photocell electrode layer
247-430-7 | 26043-11-8 | nickel hexafluorosilicate |  |  | N | Process chemical
238-278-2 | 14332-34-4 | nickel hydrogen phosphate |  |  | N | Process chemical
234-348-1 | 11113-74-9 | nickel hydroxide |  |  | N | Process chemical and in NiCd & NiMH batteries
248-585-3 | 27637-46-3 | nickel isooctanoate |  |  | 2 | Catalyst used to make rubber but likely to be <0.1%
273-749-6 | 69012-50-6 | nickel matte |  |  | N | Refinery by-product
215-215-7 | 1313-99-1 | nickel monoxide |  |  | 1 | Pigment used mainly in glass, ceramics, porcelain enamels and in catalysts. Used to make ferrites (but not present as NiO).
208-933-7 | 547-67-1 | nickel oxalate |  |  | N | Process chemical
234-323-5 | 11099-02-8 | nickel oxide |  |  | 1 | Pigment used mainly in glass, ceramics, porcelain enamels and in catalysts. Used to make ferrites (but not present as NiO).
252-840-4 | 36026-88-7 | nickel phosphinate |  |  | N | Process chemical
| 11132-10-8 | nickel potassium fluoride |  |  | N | Process chemical
239-125-2 | 15060-62-5 | nickel selenate |  |  | N | Process chemical
215-216-2 | 1314-05-2 | nickel selenide |  |  | N | Process chemical
250-788-7 | 31748-25-1 | nickel silicate (3:4) |  |  | N | Process chemical
232-104-9 | 7786-81-4, 10101-97-0, 10101-98-1 | nickel sulfate |  |  | N | Process chemical, industrial catalyst
234-349-7 | 11113-75-0, 12035-72-2 | nickel sulfide |  |  | N | Process chemical
235-260-6 | 12142-88-0 | nickel telluride |  |  | 1 | Used in electronic industries for special resistors and capacitors
239-974-9 | 15852-21-8 | nickel tellurium tetraoxide |  |  | N | Process chemical
239-967-0 | 15851-52-2 | nickel tellurium trioxide |  |  | N | Process chemical
234-824-9 | 12035-38-0 | nickel tin trioxide: nickel stannate |  |  | 2 | Additive in ceramic capacitors, may be used for electroplating
235-752-0 | 12653-76-8 | nickel titanium oxide |  |  | 1 | Pigment, gas sensors, semiconductor, high density data storage, solid fuel cells and as a catalyst
234-825-4 | 12035-39-1 | nickel titanium trioxide |  |  | 1 | Pigment, gas sensors, semiconductor, high density data storage, solid fuel cells and as a catalyst
239-876-6 | 15780-33-3 | nickel triuranium decaoxide |  |  | N | Catalyst
238-032-4 | 14177-51-6 | nickel tungsten tetraoxide |  |  | 1 | Thin film coatings
274-755-1 | 70692-93-2 | nickel zirkonium trioxide |  |  | 1 | Pigment/ceramic
242-533-3 | 18721-51-2 | nickel(II) hydrogen citrate |  |  | N | Catalyst
287-468-1 | 85508-43-6 | nickel(II) isodecanoate |  |  | 2 | Catalyst used to make rubber but likely to be <0.1%
249-555-2 | 29317-63-3 | nickel(II) isooctanoate |  |  | 2 | Catalyst used to make rubber but likely to be <0.1%
287-469-7 | 85508-44-7 | nickel(II) neodecanoate |  |  | 2 | Catalyst used to make rubber but likely to be <0.1%
300-094-6 | 93920-10-6 | nickel(II) neononanoate |  |  | 2 | Catalyst used to make rubber but likely to be <0.1%
300-093-0 | 93920-09-3 | nickel(II) neoundecanoate |  |  | 2 | Catalyst used to make rubber but likely to be <0.1%
225-656-7 | 4995-91-9 | nickel(II) octanoate |  |  | 2 | Catalyst used to make rubber but likely to be <0.1%
237-138-8 | 13654-40-5 | nickel(II) palmitate |  |  | N | Catalyst
222-102-6 | 3349-08-4 | nickel(II) propionate |  |  | N | Process chemical
233-263-7 | 10101-96-9 | nickel(II) selenite |  |  | N | Process chemical
244-578-4 | 21784-78-1 | nickel(II) silicate |  |  | N | Process chemical
218-744-1 | 2223-95-2 | nickel(II) stearate: nickel(II) octadecanoate |  |  | N | Catalyst used to make rubber but likely to be <0.1%
231-827-7 | 7757-95-1 | nickel(II) sulfite |  |  | N | Process chemical
240-235-8 | 16083-14-0 | nickel(II) trifluoroacetate |  |  | N | Process chemical
238-076-4 | 14216-75-2 | nitric acid, nickel salt |  |  | N | Process chemical
202-716-0 | 98-95-3 | nitrobenzene |  |  | N | Process chemical
217-406-0 | 1836-75-5 | nitrofen (ISO): 2,4-dichlorophenyl 4-nitrophenyl ether |  |  | N | Herbicide
210-698-0 | 621-64-7 | nitrosodipropylamine |  |  | N | Lab chemical. Contraminant of some weedkillers
206-400-3 | 335-76-2 | nonadecafluorodecanoic acid |  |  | 1 | SVHC: Main uses is for waterproofing, may have been used to make plasticisers (very unlikely), lubricants, wetting agents, surfactants and corrosion inhibitors
933-378-9 | 776297-69-9 | n-pentyl-isopentylphthalate |  |  | 1 | Plasticiser
271-112-7 | 68515-84-4 | olivine, nickel green |  |  | 1 | Pigment
237-560-2 | 13840-56-7 | orthoboric acid, sodium salt |  |  | 1 | Sodium borate, borax, preservative and flame retardant.
243-867-2 | 20543-06-0 | oxalic acid, nickel salt |  |  | N | Process chemical
417-210-7 | 70987-78-9 | oxiranemethanol, 4-methylbenzene-sulfonate, (<i>S</i>)- |  |  | N | Process chemical
| 50-29-3 | p,p’-DDT |  |  | N | Pesticide
265-174-4 | 64742-70-7 | Paraffin oils (petroleum), catalytic dewaxed heavy: Baseoil - unspecified: [A complex combination of hydrocarbons obtained from a catalytic dewaxing process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>20</sub> through C<sub>50</sub> and produces a finished oil with a viscosity of at least 100 SUS at 100 °F (19cSt at 40 °C).] |  |  | N
265-176-5 | 64742-71-8 | Paraffin oils (petroleum), catalytic dewaxed light: Baseoil - unspecified: [A complex combination of hydrocarbons obtained from a catalytic dewxing process. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C<sub>15</sub> through C<sub>30</sub> and produces a finished oil with a viscosity of less than 100 SUS at 100 °F (19cSt at 40 °C).] |  |  | N
295-810-6 | 92129-09-4 | Paraffin oils (petroleum), solvent-refined dewaxed heavy: Baseoil - unspecified: [A complex combination of hydrocarbons obtained from sulfur-containing paraffinic crude oil. It consists predominantly of a solvent refined deparaffinated lubricating oil with a viscosity of 65cSt at 50 °C.] |  |  | N
308-297-1 | 97926-77-7 | Paraffin waxes (coal), brown-coal high-temp tar, clay-treated: Coal Tar Extract: [A complex combination of hydrocarbons obtained by the treatment of lignite carbonization tar with bentonite for removal of trace constituents and impurities. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C<sub>12</sub>.] |  |  | N
308-298-7 | 97926-78-8 | Paraffin waxes (coal), brown-coal high-temp tar, silicic acid-treated: Coal Tar Extract: [A complex combination of hydrocarbons obtained by the treatment of lignite carbonization tar with silicic acid for removal of trace constituents and impurities. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C<sub>12</sub>.] |  |  | N
308-296-6 | 97926-76-6 | Paraffin waxes (coal), brown-coal high-temp. tar, carbon-treated: Coal Tar Extract: [A complet combination of hydrocarbons obtained by the treatment of lignite carbonization tar with activated carbon for removal of trace constituents and impurities. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C<sub>12</sub>.] |  |  | N
295-455-7 | 92045-72-2 | Paraffin waxes (coal), brown-coal-high-temp. tar, hydrotreated: Coal Tar Extract: [A complex combination of hydrocarbons obtained from lignite carbonization tar by solvent crystallisation (solvent deoiling), by sweating or an adducting process treated with hydrogen in the presence of a catalyst. It consists predominantly of straight and branched chain saturated hydrocarbons having carbon numbers predominantly greater than C<sub>12</sub>.] |  |  | N
295-454-1 | 92045-71-1 | Paraffin waxes (coal), brown-coal-high-temp. tar: Coal Tar Extract: [A complex combination of hydrocarbons obtained from lignite carbonization tar by solvent crystallisation (solvent deoiling), by sweating or an adducting process. It consists predominantly of straight and branched chain saturated hydrocarbons having carbon numbers predominantly greater than C<sub>12</sub>.] |  |  | N
| 1336-36-3 | PCB |  |  | N | Banned since 1980s, was used in transformers, capacitors, cables
239-172-9 | 13517-20-9, 15120-21-5 | perboric acid (H3BO2(O2)), monosodium salt trihydrate: [containing < 0,1 % (w/w) of particles with an aerodynamic diameter of below 50 µm] |  |  | N | Laundry detergent additive
231-556-4 | 7632-04-4, 10332-33-9, 10486-00-7 | perboric acid (HBO(O2)), sodium salt, tetrahydrate sodium peroxoborate hexahydrate: [containing < 0,1 % (w/w) of particles with an aerodynamic diameter of below 50 µm] |  |  | N | Laundry detergent additive
234-390-0 | 10332-33-9, 11138-47-9, 12040-72-1, 37244-98-7 | perboric acid, sodium salt, monohydrate: perboric acid, sodium salt: [containing < 0,1 % (w/w) of particles with an aerodynamic diameter of below 50 µm]: [containing < 0,1 % (w/w) of particles with an aerodynamic diameter of below 50 µm] |  |  | N | Laundry detergent additive
206-801-3 | 375-95-1 | perfluorononan-1-oic acid |  |  | 1 | SVHC. Also called PFNA, used as a processing aid for manufacture of fluoropolymers such as PVDF. Also used as an additive in lubricating oils, in cleaning agents, for waterproofing and textile coatings and in liquid crystal displays. Occurs in water and stain resistant textiles. Will not exceed 0.1% in a material, but could as an ultra-thin waterproof coating
| 4149-60-4 | perfluorononan-1-oic acid ammonium salts |  |  | 1 | SVHC. Used as a processing aid for manufacture of fluoropolymers such as PVDF. Also used as an additive in lubricating oils, in cleaning agents, for waterproofing and textile coatings and in liquid crystal displays. Occurs in water and stain resistant textiles. Will not exceed 0.1% in a material, but could as an ultra-thin waterproof coating
| 21049-39-8 | perfluorononan-1-oic acid sodium salts |  |  | 1 | SVHC. Used as a processing aid for manufacture of fluoropolymers such as PVDF. Also used as an additive in lubricating oils, in cleaning agents, for waterproofing and textile coatings and in liquid crystal displays. Occurs in water and stain resistant textiles. Will not exceed 0.1% in a material, but could as an ultra-thin waterproof coating
217-179-8 | 1763-23-1 | perfluorooctane sulfonic acid: heptadecafluorooctane-1-sulfonic acid |  |  | 1 | PFOS - banned
206-397-9 | 335-67-1 | perfluorooctanoic acid |  |  | 1 | PFOA - water and stain resistant fabrics - applied as very thin coatings
285-098-5 | 85029-74-9 | Petrolatum (petroleum), alumina-treated: Petrolatum: [A complex combination of hydrocarbons obtained when petrolatum is treated with Al<sub>2</sub>O<sub>3</sub> to remove polar components and impurities. It consists predominantly of saturated, crystalline, and liquid hydrocarbons having carbon numbers predominantly greater than C<sub>25</sub>.] |  |  | N
308-149-6 | 97862-97-0 | Petrolatum (petroleum), carbon-treated: Petrolatum: [A complex combination of hydrocarbons obtained by the treatment of petroleum petrolatum with activated carbon for the removal of trace polar constituents and impurities. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub>.] |  |  | N
309-706-6 | 100684-33-1 | Petrolatum (petroleum), clay-treated: Petrolatum: [A complex combination of hydrocarbons obtained by treatment of petrolatum with bleaching earth for the removal of traces of polar constituents and impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of greater than C<sub>25</sub>.] |  |  | N
295-459-9 | 92045-77-7 | Petrolatum (petroleum), hydrotreated: Petrolatum: [A complex combination of hydrocarbons obtained as a semi-solid from dewaxed paraffinic residual oil treated with hydrogen in the presence of a catalyst. It consists predominantly of saturated microcrystalline and liquid hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub>.] |  |  | N
265-206-7 | 64743-01-7 | Petrolatum (petroleum), oxidized: Petrolatum: [A complex combination of organic compounds, predominantly high molecular weight carboxylic acids, obtained by the air oxidation of petrolatum.] |  |  | N
308-150-1 | 97862-98-1 | Petrolatum (petroleum), silicic acid-treated: Petrolatum: [A complex combination of hydrocarbons obtained by the treatment of petroleum petrolatum with silicic acid for the removal of trace polar constituents and impurities. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub>.] |  |  | N
232-373-2 | 8009-03-8 | Petrolatum: Petrolatum: [A complex combination of hydrocarbons obtained as a semi-solid from dewaxing paraffinic residual oil. It consists predominantly of saturated crystalline and liquid hydrocarbons having carbon numbers predominantly greater than C<sub>25.</sub>] |  |  | 1 | Also called petroleum jelly, mainly has medical uses, but is used as a lubricant. Petrolatums with different CAS numbers not likely to be used
295-463-0 | 92045-80-2 | Petroleum gases, liquefied, sweetened, C4 fraction: Petroleum gas: [A complex combination of hydrocarbons obtained by subjecting a liquified petroleum gas mix to a sweetening process to oxidize mercaptans or to remove acidic impurities. It consists predominantly of C4 saturated and unsaturated hydrocarbons.] |  |  | N
270-705-8 | 68476-86-8 | Petroleum gases, liquefied, sweetened: Petroleum gas: [A complex combination of hydrocarbons obtained by subjecting liquefied petroleum gas mix to a sweetening process to convert mercaptans or to remove acidic impurities. It consists of hydrocarbons having carbon numbers predominantly in the range of C3 through C7 and boiling in the range of approximately – 40 °C to 80 °C (– 40 °F to 176 °F).] |  |  | N
270-704-2 | 68476-85-7 | Petroleum gases, liquefied: Petroleum gas: [A complex combination of hydrocarbons produced by the distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly in the range of C3 through C7 and boiling in the range of approximately – 40 °C to 80 °C (– 40 °F to 176 °F).] |  |  | N
271-058-4 | 68514-79-4 | Petroleum products, hydrofiner-powerformer reformates: Low boiling point cat-reformed naphtha: [The complex combination of hydrocarbons obtained in a hydrofiner-powerformer process and boiling in a range of approximately 27°C to 210°C (80°F to 410°F).] |  |  | N
271-750-6 | 68607-11-4 | Petroleum products, refinery gases: Refinery gas: [A complex combination which consists primarily of hydrogen with various small amounts of methane, ethane, and propane.] |  |  | N
232-298-5 | 8002-05-9 | Petroleum: Crude oil: [A complex combination of hydrocarbons, It consists predominantly of aliphatic, alicyclic and aromatic hydrocarbons. It may also contain small amounts of nitrogen, oxygen and sulfur compounds. This category encompasses light, medium, and heavy petroleums, as well as the oils extended from tar sands. Hydrocarbonaceous materials requiring major chemical changes for their recovery or conversion to petroleum refinery feedstocks such as crude shale oils; upgraded shale oils and liquid coal fuels are not included in this definition.] |  |  | N
310-169-5 | 122070-78-4 | Phenanthrene, distn. residues: Heavy Anthracene Oil Redistillate: [Residue from the distillation of crude phenanthrene boiling in the approximate range of 340 °C to 420 °C (644 °F to 788 °F). It consists predominantly of phenanthrene, anthracene and carbazole.] |  |  | N
616-100-8 | 74499-35-7 | phenol, (tetrapropenyl) derivatives |  |  | N | Process chemical
640-104-9 | 210555-94-5 | phenol, 4-dodecyl-, branched: phenol, 2-dodecyl-, branched: phenol, 3-dodecyl-, branched |  |  | N | Process chemical
310-154-3 | 121158-58-5 | phenol, dodecyl-, branched |  |  | N | Process chemical
201-004-7 | 77-09-8 | phenolphthalein |  |  | N | pH indicator and laxative (SVHC)
284-881-9 | 84988-93-2 | Phenols, ammonia liquor ext.: Alkaline Extract: [The combination of phenols extracted, using isobutyl acetate, from the ammonia liquor condensed from the gas evolved in low-temperature (less than 700°C (1292°F)) destructive distillation of coal. It consists predominantly of a reaction mass of monohydric and dihydric phenols.] |  |  | N | Process chemical
293-435-2 | 91079-47-9 | Phenols, C<sub>9-11</sub>: Distillate Phenols |  |  | N | Process chemical
204-557-2 | 122-60-1 | phenyl glycidyl ether: 2,3-epoxypropyl phenyl ether: 1,2-epoxy-3-phenoxypropane |  |  | 1 | Process chemical for epoxy resins but also reported used as a stabiliser for various polymers (vinyl, epoxy, rubber), likely to be uncommon
202-873-5 | 100-63-0 | phenylhydrazine |  |  | N | Process chemical
248-259-0 | 27140-08-5 | phenylhydrazine hydrochloride |  |  | N | Process chemical
200-444-7 | 59-88-1 | phenylhydrazinium chloride |  |  | N | Process chemical
257-622-2 | 52033-74-6 | phenylhydrazinium sulphate (2:1) |  |  | N | Process chemical
| 17169-61-8 | phosphoric acid, calcium nickel salt |  |  | N | Industrial process catalyst
310-162-7 | 121575-60-8 | Pitch, coal tar, high-temp., heat-treated: Pitch: [The heat treated residue from the distillation of high temperature coal tar. A black solid with an approximate softening point from 80 °C to 180 °C (176 °F to 356 °F). Composed primarily of a complex mixture of three or more membered condensed ring aromatic hydrocarbons.] |  |  | N
302-650-3 | 94114-13-3 | Pitch, coal tar, high-temp., secondary: Pitch Redistillate: [The residue obtained during the distillation of high boiling fractions from bituminous coal high temperature tar and/or pitch coke oil, with a softening point of 140 °C to 170 °C (284 °F to 392 °F) according to DIN 52025. Composed primarily of tri- and polynuclear aromatic compounds which also contain heteroatoms.] |  |  | N
266-028-2 | 65996-93-2 | pitch, coal tar, high-temp.: [The residue from the distillation of high temperature coal tar. A black solid with an approximate softening point from 30 °C to 180 °C (86 °F to 356 °F). Composed primarily of a complex mixture of three or more membered condensed ring aromatic hydrocarbons.] |  |  | N
292-653-5 | 90669-58-2 | Pitch, coal tar, low-temp., heat-treated: Pitch Residue, oxidised: Pitch Residue, heat-treated: [A complex black solid obtained by the heat treatment of low temperature coal tar pitch. It has a softening point within the approximate range of 50 °C to 140 °C (122 °F to 284 °F). Composed primarily of a complex mixture of aromatic compounds.] |  |  | N
292-654-0 | 90669-59-3 | Pitch, coal tar, low-temp., oxidized: Pitch Residue, oxidised: [The product obtained by air-blowing, at elevated temperature, low-temperature coal tar pitch. It has a softening-point within the approximate range of 70 °C to 180 °C (158 °F to 356 °F). Composed primarily of a complex mixture of hydrocarbons.] |  |  | N
292-651-4 | 90669-57-1 | Pitch, coal tar, low-temp: Pitch Residue: [A complex black solid or semi-solid obtained from the distillation of a low temperature coal tar. It has a softening point within the approximate range of 40 °C to 180 °C (104 °F to 356 °F). Composed primarily of a complex mixture of hydrocarbons.] |  |  | N
269-109-0 | 68187-57-5 | Pitch, coal tar-petroleum: Pitch Residues: [The residue from the distillation of a mixture of coal tar and aromatic petroleum streams. A solid with a softening point from 40 °C to 180 °C (140 °F to 356 °F). Composed primarily of a complex combination of three or more membered condensed ring aromatic hydrocarbons.] |  |  | N
263-072-4 | 61789-60-4 | Pitch: Pitch |  |  | N
418-260-2 | 183196-57-8 | potassium 1-methyl-3-morpholinocarbonyl-4-[3-(1-methyl-3-morpholinocarbonyl-5-oxo-2-pyrazolin-4-ylidene)-1-propenyl]pyrazole-5-olate: [containing < 0.5 % <i>N,N</i>-dimethylformamide (EC no 200-679-5)] |  |  | N | Uncertain, probably used only as a process chemical
231-829-8 | 7758-01-2 | potassium bromate |  |  | N | Process chemical
232-140-5 | 7789-00-6 | potassium chromate |  |  | N | Process chemical
231-906-6 | 7778-50-9 | potassium dichromate |  |  | 3 | Process chemical and used for sealing anodised aluminium (to prevent dye leakage). Would not comply with RoHS and requires REACH authorisation if used in the EU.
220-527-1 | 2795-39-3 | potassium perfluorooctanesulfonate: potassium heptadecafluorooctane-1-sulfonate |  |  | 1 | A PFOS salt, used as a fluorosurfactant in stain and water-proof coatings for fabrics and textiles. Also added to waxes, polishes, paints, varnishes, and cleaning products
200-879-2 | 75-56-9 | propylene oxide: 1,2-epoxypropane: methyloxirane |  |  | N | Process chemical
269-929-9 | 68391-11-7 | Pyridine, alkyl derivs.: Crude Tar Bases: [The complex combination of polyalkylated pyridines derived from coal tar distillation or as high-boiling distillates approximately above 150°C (302°F) from the reaction of ammonia with acetaldehyde, formaldehyde or paraformaldehyde.] |  |  | N
205-711-1 | 148-24-3 | quinolin-8-ol: 8-hydroxyquinoline |  |  | N | Process chemical, biocide (antiseptic) and used to make OLEDs
202-051-6 | 91-22-5 | quinoline |  |  | N | Process chemical
270-088-5 | 68410-71-9 | Raffinates (petroleum), catalytic reformer ethylene glycol-water countercurrent exts.: Low boiling point modified naphtha: [A complex combination of hydrocarbons obtained as the raffinate from the UDEX extraction process on the catalytic reformer stream. It consists of saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>9</sub>.] |  |  | N
270-349-3 | 68425-35-4 | Raffinates (petroleum), reformer, Lurgi unit-sepd.: Low boiling point modified naphtha: [The complex combination of hydrocarbons obtained as a raffinate from a Lurgi separation unit. It consists predominantly of non-aromatic hydrocarbons with various small amounts of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>8</sub>.] |  |  | N
307-769-4 | 97722-19-5 | Raffinates (petroleum), steam-cracked C4 fraction cuprous ammonium acetate extn., C3-5 and C3-5 unsatd., butadiene-free: Petroleum gas |  |  | N
412-790-8 |  | reaction mass of: <i>N</i>-[3-hydroxy-2-(2-methylacryloylaminomethoxy)propoxymethyl]-2-methylacrylamide: <i>N</i>-[2,3-bis-(2-methylacryloylaminomethoxy)propoxymethyl]-2-methylacrylamide: methacrylamide: 2-methyl-<i>N</i>-(2-methylacryloylaminomethoxymethyl)-acrylamide: <i>N</i>-(2,3-dihydroxypropoxymethyl)-2-methylacrylamide |  |  | 4 | Uncertain but one reference indicates that it is used in cosmetics, so not likely to be applicable toi medical devices
421-550-1 |  | reaction mass of: 1,3,5-tris(3-aminomethylphenyl)-1,3,5-(1<i>H</i>,3<i>H</i>,5<i>H</i>)-triazine-2,4,6-trione: reaction mass of oligomers of 3,5-bis(3-aminomethylphenyl)-1-poly[3,5-bis(3-aminomethylphenyl)-2,4,6-trioxo-1,3,5-(1<i>H</i>,3<i>H</i>,5<i>H</i>)-triazin-1-yl]-1,3,5-(1<i>H</i>,3<i>H</i>,5<i>H</i>)-triazine-2,4,6-trione |  |  | 4 | Uncertain, registration dossier for 1,3,5-tris(oxiranylmethyl)-1,3,5-triazine-2,4,6(1H,3H,5H)-trione indicates that this is a process chemical only, used to make powder coatings, solder resist inks and special electronic tapes. Reacts in use so would not occur in products.
403-250-2 |  | reaction mass of: 4-[[bis-(4-fluorophenyl)methylsilyl]methyl]-4<i>H</i>-1,2,4-triazole: 1-[[bis-(4-fluorophenyl)methylsilyl]methyl]-1<i>H</i>-1,2,4-triazole |  |  | 4 | Uncertain, but may be used as a fungicide
435-960-3 |  | reaction mass of: dimethyl (2-(hydroxymethylcarbamoyl)ethyl)phosphonate: diethyl (2-(hydroxymethylcarbamoyl)ethyl)phosphonate: methyl ethyl (2-(hydroxymethylcarbamoyl)ethyl)phosphonate |  |  | 4 | Unknown, but indication that this may be used to give special effects to textiles
402-660-9 |  | reaction mass of: disodium 4-(3-ethoxycarbonyl-4-(5-(3-ethoxycarbonyl-5-hydroxy-1-(4-sulfonatophenyl)pyrazol-4-yl)penta-2,4-dienylidene)-4,5-dihydro-5-oxopyrazol-1-yl)benzenesulfonate: trisodium 4-(3-ethoxycarbonyl-4-(5-(3-ethoxycarbonyl-5-oxido-1-(4-sulfonatophenyl)pyrazol-4-yl)penta-2,4-dienylidene)-4,5-dihydro-5-oxopyrazol-1-yl)benzenesulfonate |  |  | 4 | Unknown. Registered for REACH by one legal entity (not identified) as a new substance, but uses not specified. Banned in cosmetics, which indicates an old use.
309-710-8 | 100684-37-5 | Residual oils (petroleum), carbon-treated solvent-dewaxed: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by the treatment of solvent-dewaxed petroleum residual oils with activated charcoal for the removal of trace polar constituents and impurities.] |  |  | N
294-843-3 | 91770-57-9 | Residual oils (petroleum), catalytic dewaxed: Baseoil - unspecified |  |  | N
309-711-3 | 100684-38-6 | Residual oils (petroleum), clay-treated solvent-dewaxed: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treatment of solvent-dewaxed petroleum residual oils with bleaching earth for the removal of trace polar constituents and impurities.] |  |  | N
265-143-5 | 64742-41-2 | Residual oils (petroleum), clay-treated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treatment of a residual oil with a natural or modified clay in either a contacting or percolation process to remove the trace amounts of polar compounds and impurities present. It consists of hydro-carbons having carbon numbers predominantly higher than C<sub>25</sub> and boiling above approximately 400 °C (752 °F).] |  |  | N
270-795-9 | 68478-16-0 | Residual oils (petroleum), deisobutanizer tower: Low boiling point naphtha - unspecified: [A complex residuum from the atmospheric distillation of the butane-butylene stream. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>6</sub>.] |  |  | N
270-795-9 | 68478-16-0 | Residual oils (petroleum), deisobutanizer tower: Low boiling point naphtha - unspecified: [A complex residuum from the atmospheric distillation of the butane-butylene stream. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C<sub>4</sub> through C<sub>6</sub>.] |  |  | N
295-499-7 | 92061-86-4 | Residual oils (petroleum), hydrocracked acid-treated solvent-dewaxed: Baseoil - unspecified: [A complex combination of hydrocarbons produced by solvent removal of paraffins from the residue of the distillation of acid-treated, hydrocracked heavy paraffins and boiling approximately above 380 °C (716 °F).] |  |  | N
292-656-1 | 90669-74-2 | Residual oils (petroleum), hydrotreated solvent dewaxed: Baseoil - unspecified |  |  | N
265-160-8 | 64742-57-0 | Residual oils (petroleum), hydrotreated: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists of hydrocarbons having carbon numbers predominantly greater than C<sub>25</sub> and boiling above approximately 400 °C (752 °F).] |  |  | N
265-096-0 | 64741-95-3 | Residual oils (petroleum), solvent deasphalted: Baseoil - unspecified: [A complex combination of hydrocarbons obtained as the solvent soluble fraction from C<sub>3</sub>-C<sub>4</sub> solvent deasphalting of a residuum. It consists of hydrocarbons having carbon numbers predominantly higher than C<sub>25</sub> and boiling above approximately 400 °C (752 °F).] |  |  | N
265-166-0 | 64742-62-7 | Residual oils (petroleum), solvent-dewaxed: Baseoil - unspecified: [A complex combination of hydrocarbons obtained by removal of long, branched chain hydrocarbons from a residual oil by solvent crystallization. It consists of hydrocarbons having carbon numbers predominantly greater than C<sub>25</sub> and boiling above approximately 400 °C (752 °F).] |  |  | N
298-754-0 | 93821-66-0 | Residual oils (petroleum): Heavy Fuel oil: [A complex combination of hydrocarbons, sulfur compounds and metal-containing organic compounds obtained as the residue from refinery fractionation cracking processes. It produces a finished oil with a viscosity above 2cSt. at 100 °C.] |  |  | N
265-101-6 | 64742-01-4 | Residual oils (petroleum,) solvent-refined: Baseoil - unspecified: [A complex combination by hydrocarbons obtained as the solvent insoluble fraction from solvent refining of a residuum using a polar organic solvent such as phenol or furfural. It consists of hydrocarbons having carbon numbers predominantly higher than C<sub>25</sub> and boiling above approximately 400 °C (752 °F).] |  |  | N
295-505-8 | 92061-92-2 | Residues (coal tar), anthracene oil distn.: Anthracene Oil Fraction: [The residue from the fraction distillation of crude anthracene boiling in the approximate range of 340°C to 400°C (644°F to 752°F). It consists predominantly of tri- and polynuclear aromatic and heterocyclic hydrocarbons.] |  |  | N
295-506-3 | 92061-93-3 | Residues (coal tar), creosote oil distn.: Wash Oil Redistillate: [The residue from the fractional distillation of wash oil boiling in the approximate range of 270°C to 330°C (518°F to 626°F). It consists predominantly of dinuclear aromatic and heterocyclic hydrocarbons.] |  |  | N
295-507-9 | 92061-94-4 | Residues (coal tar), pitch distn.: Pitch Redistillate: [Residue from the fractional distillation of pitch distillate boiling in the range of approximately 400 °C to 470 °C (752 °F to 846 °F). Composed primarily of polynuclear aromatic hydrocarbons, and heterocyclic compounds.] |  |  | N
302-681-2 | 94114-46-2 | Residues (coal), liq. solvent extn.: [A cohesive powder composed of coal mineral matter and undissolved coal remaining after extraction of coal by a liquid solvent.] |  |  | N
271-010-2 | 68513-66-6 | Residues (petroleum), alkylation splitter, C4-rich: Petroleum gas: [A complex residuum from the distillation of streams various refinery operations. It consists of hydrocarbons having carbon numbers in the range of C4 through C5, predominantly butane and boiling in the range of approximately – 11.7 °C to 27.8 °C (11 °F to 82 °F).] |  |  | N
265-045-2 | 64741-45-3 | Residues (petroleum), atm. tower: Heavy Fuel oil: [A complex residuum from the atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub> and boiling above approximately 350 °C (662 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
269-777-3 | 68333-22-2 | Residues (petroleum), atmospheric: Heavy Fuel oil: [A complex residuum from atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly greater than C<sub>11</sub> and boiling above approximately 200 °C (392 °F). This stream is likely to contain 5 wt. % or more of 4-to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
270-791-7 | 68478-12-6 | Residues (petroleum), butane splitter bottoms: Low boiling point naphtha - unspecified: [A complex residuum from the distillation of butane stream. It consists of aliphatic hydrocarbons having carbon numbers predominantly in the range of C4 through C6.] |  |  | N
270-794-3 | 68478-15-9 | Residues (petroleum), C<sub>6-8</sub> catalytic reformer: Low boiling point cat-reformed naphtha: [A complex residuum from the catalytic reforming of C<sub>6-8</sub> feed. It consists of hydrocarbons having carbon numbers predominantly in the range of C<sub>2</sub> through C<sub>6</sub>.] |  |  | N
295-511-0 | 92061-97-7 | Residues (petroleum), catalytic cracking: Heavy Fuel oil: [A complex combination of hydrocarbons produced as the residual fraction from the distillation of the products from a catalytic cracking process. It consists predominantly of hydrocarbons having carbon numbers predominantly greater than C<sub>11</sub> and boiling above approximately 200 °C (392 °F).] |  |  | N
270-792-2 | 68478-13-7 | Residues (petroleum), catalytic reformer fractionator residue distn.: Heavy Fuel oil: [A complex residuum from the distillation of catalytic reformer fractionator residue. It boils approximately above 399 °C (750 °F).] |  |  | N
265-069-3 | 64741-67-9 | Residues (petroleum), catalytic reformer fractionator: Heavy Fuel oil: [A complex combination of hydrocarbons produced as the residual fraction from distillation of the product from a catalytic reforming process. It consists of predominantly aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>10</sub> through C<sub>25</sub> and boiling in the range of approximately 160 °C to 400 °C (320 °F to 725 °F). This stream is likely to contain 5 wt. % or more of 4- or 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
272-187-9 | 68783-13-1 | Residues (petroleum), coker scrubber, Condensed-ring-arom.-contg.: Heavy Fuel oil: [A very complex combination of hydrocarbons produced as the residual fraction from the distillation of vaccum residuum and the products from a thermal cracking process. It consists predominantly of hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub> and boiling above approximately 350 °C (662 °F). This stream is likely to contain 5 wt.% or more of 4- to 6-membered condensed rind aromatic hydrocarbons.] |  |  | N
270-983-0 | 68512-61-8 | Residues (petroleum), heavy coker and light vacuum: Heavy Fuel oil: [A complex combination of hydrocarbons produced as the residual fraction from the distillation of heavy coker gas oil and light vacuum gas oil. It consists predominantly of hydrocarbons having carbon numbers predominantly greater than C<sub>13</sub> and boiling above approximately 230 °C (446 °F).] |  |  | N
270-796-4 | 68478-17-1 | Residues (petroleum), heavy coker gas oil and vacuum gas oil: Heavy Fuel oil: [A complex combination of hydrocarbons produced as the residual fraction from the distillation of heavy coker gas oil and vacuum gas oil. It predominantly consists of hydrocarbons having carbon numbers predominantly greater than C<sub>13</sub> and boiling above approximately 230 °C (446 °F).] |  |  | N
265-076-1 | 64741-75-9 | Residues (petroleum), hydrocracked: Heavy Fuel oil: [A complex combination of hydrocarbons produced as the residual fraction from distillation of the products of a hydrocracking process. It consists of hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub> and boiling above approximately 350 °C (662 °F).] |  |  | N
265-181-2 | 64742-78-5 | Residues (petroleum), hydrodesulfurized atmospheric tower: Heavy Fuel oil: [A complex combination of hydrocarbons obtained by treating an atmospheric tower residuum with hydrogen in the presence of a catalyst under conditions primarily to remove organic sulfur compounds. It consists of hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub> and boiling above approximately 350 °C (662 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
295-514-7 | 92062-00-5 | Residues (petroleum), hydrogenated steam-cracked naphtha: Cracked gasoil: [A complex combination of hydrocarbons obtained as a residual fraction from the distillation of hydrotreated steam-cracked naphtha. It consists predominantly of hydrocarbons boiling in the range of approximately 200 °C to 350 °C (32 °F to 662 °F).] |  |  | N
270-984-6 | 68512-62-9 | Residues (petroleum), light vacuum: Heavy Fuel oil: [A complex residuum from the vacuum distillation of the residuum from the atmospheric distillation of crude oil. It consists of hydrocarbons having carbon numbers predominantly greater than C<sub>13</sub> and boiling above approximately 230 °C (446 °F).] |  |  | N
297-905-8 | 93763-85-0 | Residues (petroleum), steam-cracked heat-soaked naphtha: Cracked gasoil: [A complex combination of hydrocarbons obtained as residue from the distillation of steam cracked heat soaked naphtha and boiling in the range of approximately 150 °C to 350 °C (302 °F to 662 °F).] |  |  | N
310-057-6 | 102110-55-4 | Residues (petroleum), steam-cracked light, arom.: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained by the distillation of the products of steam cracking or similar processes after taking off the very light products resulting in a residue starting with hydrocarbons having carbon numbers greater than C<sub>5</sub>. It consists predominantly of aromatic hydrocarbons having carbon numbers greater than C<sub>5</sub> and boiling above approximately 40°C (104°F).] |  |  | N
271-013-9 | 68513-69-9 | Residues (petroleum), steam-cracked light: Heavy Fuel oil: [A complex residuum from the distillation of the products from a steam-cracking process. It consists predominantly of aromatic and unsaturated hydrocarbons having carbon numbers greater than C<sub>7</sub> and boiling in the range of approximately 101 °C to 555 °C (214 °F to 1030 °F).] |  |  | N
295-517-3 | 92062-04-9 | Residues (petroleum), steam-cracked naphtha distn.: Cracked gasoil: [A complex combination of hydrocarbons obtained as a column bottom from the separation of effluents from steam cracking naphtha at a high temperature. It boils in the range of approximately 147 °C to 300 °C (297 °F to 572 °F) and produces a finished oil having a viscosity of 18cSt at 50 °C.] |  |  | N
292-657-7 | 90669-75-3 | Residues (petroleum), steam-cracked, distillates: Heavy Fuel oil: [A complex combination of hydrocarbons obtained during the production of refined petroleum tar by the distillation of steam cracked tar. It consists predominantly of aromatic and other hydrocarbons and organic sulfur compounds.] |  |  | N
273-272-3 | 68955-36-2 | Residues (petroleum), steam-cracked, resinous: Heavy Fuel oil: [A complex residuum from the distillation of steam-cracked petroleum residues.] |  |  | N
265-193-8 | 64742-90-1 | Residues (petroleum), steam-cracked: Heavy Fuel oil: [A complex combination of hydrocarbons obtained as the residual fraction from the distillation of the products of a steam cracking process (including steam cracking to produce ethylene). It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly greater than C<sub>14</sub> and boiling above approximately 260 °C (500 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
265-081-9 | 64741-80-6 | Residues (petroleum), thermal cracked: Heavy Fuel oil: [A complex combination of hydrocarbons produced as the residual fraction from distillation of the product from a thermal cracking process. It consists predominantly of unsaturated hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub> and boiling above approximately 350 °C (662 °F). This stream is likely to contain 5 wt. % or more of 4- to 6-membered condensed ring aromatic hydrocarbons.] |  |  | N
271-763-7 | 68607-30-7 | Residues (petroleum), topping plant, low-sulfur: Heavy Fuel oil: [A low-sulfur complex combination of hydrocarbons produced as the residual fraction from the topping plant distillation of crude oil. It is the residuum after the straight-run gasoline cut, kerosene cut and gas oil cut have been removed.] |  |  | N
292-658-2 | 90669-76-4 | Residues (petroleum), vacuum, light: Heavy Fuel oil: [A complex residuum from the vacuum distillation of the residuum from atmospheric distillation of crude oil. It consists predominantly of hydrocarbons having carbon numbers predominantly greater than C<sub>24</sub> and boiling above approximately 390 °C (734 °F).] |  |  | N
308-733-0 | 98219-64-8 | Residues, steam cracked, thermally treated: Heavy Fuel oil: [A complex combination of hydrocarbons obtained by the treatment and distillation of raw steam-cracked naphtha. It consists predominantly of unsaturated hydrocarbons boiling in the range above approximately 180 °C (356 °F).] |  |  | N
202-345-4 | 94-59-7 | safrole: 5-allyl-1,3-benzodioxole |  |  | N | Banned food flavouring
209-030-0 | 553-00-4 | salts of 2-naphthylamine |  |  | N | Banned intermediate
210-313-6 | 612-52-2 | salts of 2-naphthylamine |  |  | N | Banned intermediate
210-322-5 | 612-82-8 | salts of 4,4'-bi-<i>o</i>-toluidine: salts of 3,3'-dimethylbenzidine: salts of <i>o</i>-tolidine |  |  | N | Process chemical
265-294-7 | 64969-36-4 | salts of 4,4'-bi-<i>o</i>-toluidine: salts of 3,3'-dimethylbenzidine: salts of <i>o</i>-tolidine |  |  | N | Process chemical
277-985-0 | 74753-18-7 | salts of 4,4'-bi-<i>o</i>-toluidine: salts of 3,3'-dimethylbenzidine: salts of <i>o</i>-tolidine |  |  | N | Process chemical
244-236-4 | 21136-70-9 | salts of benzidine [ |  |  | N | Process chemical
252-984-8 | 36341-27-2 | salts of benzidine [ |  |  | N | Process chemical
208-519-6 | 531-85-1 | salts of benzidine [ |  |  | N | Process chemical
208-520-1 | 531-86-2 | salts of benzidine [ |  |  | N | Process chemical
| 68130-19-8 | silicic acid, lead nickel salt: |  |  | N | Process chemical
253-461-7 | 37321-15-6 | silicic acid, nickel salt |  |  | 1 | Green pigment
292-659-8 | 90669-77-5 | Slack wax (petroleum), acid-treated: Slack wax: [A complex combination of hydrocarbons obtained as a raffinate by treatment of a petroleum slack wax fraction with sulfuric acid treating process. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub>.] |  |  | N
309-723-9 | 100684-49-9 | Slack wax (petroleum), carbon-treated: Slack wax: [A complex combination of hydrocarbons obtained by treatment of petroleum slack wax with activated charcoal for the removal of trace polar constituents and impurities.] |  |  | N
292-660-3 | 90669-78-6 | Slack wax (petroleum), clay-treated: Slack wax: [A complex combination of hydrocarbons obtained by treatment of a petroleum slack wax fraction with natural or modified clay in either a contacting or percolation process. It consists predominantly of saturated straight and branched hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub>.] |  |  | N
295-523-6 | 92062-09-4 | Slack wax (petroleum), hydrotreated: Slack wax: [A complex combination of hydrocarbons obtained by treating slack wax with hydrogen in the presence of a catalyst. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub>.] |  |  | N
308-155-9 | 97863-04-2 | Slack wax (petroleum), low-melting, carbon-treated: Slack wax: [A complex combination of hydrocarbons obtained by the treatment of low-melting slack wax with activated carbon for the removal of trace polar constituents and impurities. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C<sub>12</sub>.] |  |  | N
308-156-4 | 97863-05-3 | Slack wax (petroleum), low-melting, clay-treated: Slack wax: [A complex combination of hydrocarbons obtained by the treatment of low-melting petroleum slack wax with bentonite for removal of trace polar constituents and impurities. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C<sub>12</sub>.] |  |  | N
295-525-7 | 92062-11-8 | Slack wax (petroleum), low-melting, hydrotreated: Slack wax: [A complex combination of hydrocarbons obtained by treatment of low-melting petroleum slack wax with hydrogen in the presence of a catalyst. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C12.] |  |  | N
308-158-5 | 97863-06-4 | Slack wax (petroleum), low-melting, silicic acid-treated: Slack wax: [A complex combination of hydrocarbons obtained by the treatment of low-melting petroleum slack wax with silicic acid for the removal of trace polar constituents and impurities. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C<sub>12</sub>.] |  |  | N
295-524-1 | 92062-10-7 | Slack wax (petroleum), low-melting: Slack wax: [A complex combination of hydrocarbons obtained from a petroleum fraction by solvent deparaffination. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C<sub>12</sub>.] |  |  | N
265-165-5 | 64742-61-6 | Slack wax (petroleum): Slack wax: [A complex combination of hydrocarbons obtained from a petroleum fraction by solvent crystallization (solvent dewaxing) or as a distillation fraction from a very waxy crude. It consists predominantly of saturated straight and branched chain hydrocarbons having carbon numbers predominantly greater than C<sub>20</sub>.] |  |  | N
305-433-1 | 94551-87-8 | slimes and sludges, copper electrolyte refining, decopperised |  |  | N
295-859-3 | 92129-57-2 | slimes and sludges, copper electrolytic refining, decopperised, nickel sulfate |  |  | N | Process waste ujsed as a metals feedstock
231-889-5 | 7775-11-3 | sodium chromate |  |  | N | Process chemical
234-190-3 | 7789-12-0, 10588-01-9 | sodium dichromate |  |  | 3 | Process chemical and aluminium anodised coating sealant (to prevent dye leakage)
| 3830-45-3 | sodium nonadecafluorodecanoate |  |  | 1 | Main uses is for waterproofing, may have been used to make plasticisers (very unlikely), lubricants, wetting agents, surfactants and corrosion inhibitors. Unlikely to exceed 0.1%
239-172-9 | 13517-20-9, 15120-21-5 | sodium perborate: [containing = 0,1 % (w/w) of particles with an aerodynamic diameter of below 50 µm] |  |  | N | Laundry detergent additive
239-172-9 | 13517-20-9, 15120-21-5 | sodium perborate: [containing = 0,1 % (w/w) of particles with an aerodynamic diameter of below 50 µm] |  |  | N | Laundry detergent additive
231-556-4 | 7632-04-4, 10332-33-9, 10486-00-7 | sodium peroxometaborate: sodium peroxoborate: [containing < 0,1 % (w/w) of particles with an aerodynamic diameter of below 50 µm] |  |  | N | Laundry detergent additive
287-500-4 | 85536-19-2 | Solvent naphtha (coal), coumarone-styrene contg.: Light Oil Redistillate, intermediate boiling |  |  | N
287-498-5 | 85536-17-0 | Solvent naphtha (coal), light: Light Oil Redistillate, low boiling |  |  | N
287-502-5 | 85536-20-5 | Solvent naphtha (coal), xylene-styrene cut: Light Oil Redistillate, intermediate boiling |  |  | N
266-013-0 | 65996-79-4 | Solvent naphtha (coal): Light Oil Extract Residues, high boiling: [The distillate from either high temperature coal tar, coke oven light oil, or coal tar oil alkaline extract residue having an approximate distillation range of 130°C to 210°C (266°F to 410°F). Composed primarily of indene and other polycyclic ring systems containing a single aromatic ring. May contain phenolic compounds and aromatic nitrogen bases.] |  |  | N
295-529-9 | 92062-15-2 | Solvent naphtha (petroleum), hydrotreated light naphthenic: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists predominantly of cycloparaffinic hydrocarbons having carbon numbers predominantly in the range of C<sub>6</sub> through C<sub>7</sub> and boiling in the range of approximately 73°C to 85°C (163°F to 185°F).] |  |  | N | Solvent for adhesives, paints, inks, should not be present in medical devices
265-192-2 | 64742-89-8 | Solvent naphtha (petroleum), light aliph.: Low boiling point naphtha: [A complex combination of hydrocarbons obtained from the distillation of crude oil or natural gasoline. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C<sub>5</sub> through C<sub>10</sub> and boiling in the range of approximately 35°C to 160°C (95°F to 320°F).] |  |  | N | Solvent for adhesives, paints, inks, should not be present in medical devices
270-988-8 | 68512-78-7 | Solvent naphtha (petroleum), light arom., hydrotreated: Low boiling point hydrogen treated naphtha: [A complex combination of hydrocarbons obtained by treating a petroleum fraction with hydrogen in the presence of a catalyst. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>8</sub> through C<sub>10</sub> and boiling in the range of approximately 135°C to 210°C (275°F to 410°F).] |  |  | N | Solvent for adhesives, paints, inks, should not be present in medical devices
265-199-0 | 64742-95-6 | Solvent naphtha (petroleum), light arom.: Low boiling point naphtha - unspecified: [A complex combination of hydrocarbons obtained from distillation of aromatic streams. It consists predominantly of aromatic hydrocarbons having carbon numbers predominantly in the range of C<sub>8</sub> through C<sub>10</sub> and boiling in the range of approximately 135°C to 210°C (275°F to 410°F).] |  |  | N | Solvent for adhesives, paints, inks, should not be present in medical devices
232-489-3 | 8052-41-3 | stoddard solvent: Low boiling point naphtha - unspecified: [A colourless, refined petroleum distillate that is free from rancid or objectionable odours and that boils in a range of approximately 148,8 °C to 204,4 °C (300 °F to 400 °F).] |  |  | N
232-142-6 | 7789-06-2 | strontium chromate |  |  | 1 | SVHC: Paint primers
| 100-42-5 | Styrene |  |  | N | Process chemical & monomer
202-476-7 | 96-09-3 | styrene oxide: (epoxyethyl)benzene: phenyloxirane |  |  | N | Process chemical
202-388-9 | 95-06-7 | sulfallate (ISO): 2-chloroallyl <i>N</i>,<i>N</i>-dimethyldithiocarbamate |  |  | N | Herbicide
270-802-5 | 68478-21-7 | Tail gas (petroleum), catalytic cracked clarified oil and thermal cracked vacuum residue fractionation reflux drum: Petroleum gas: [A complex combination of hydrocarbons obtained from fractionation of catalytic cracked clarified oil and thermal cracked vacuum residue. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
269-617-2 | 68307-98-2 | Tail gas (petroleum), catalytic cracked distillate and catalytic cracked naphtha fractionation absorber: Petroleum gas: [The complex combination of hydrocarbons from the distillation of the products from catalytic cracked distillates and catalytic cracked naphtha. It consists predominantly of hydrocarbons having carbon numbers in the range of C1 through C4.] |  |  | N
273-170-9 | 68952-77-2 | Tail gas (petroleum), catalytic cracked distillate and naphtha stabilizer: Petroleum gas: [A complex combination of hydrocarbons obtained by the fractionation of catalytic cracked naphtha and distillate. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
270-803-0 | 68478-22-8 | Tail gas (petroleum), catalytic cracked naphtha stabilization absorber: Petroleum gas: [A complex combination of hydrocarbons obtained from the stabilization of catalytic cracked naphtha. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
270-805-1 | 68478-25-1 | Tail gas (petroleum), catalytic cracker refractionation absorber: Refinery gas: [A complex combination of hydrocarbons obtained from refractionation of products from a catalytic cracking process. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C3.] |  |  | N
270-804-6 | 68478-24-0 | Tail gas (petroleum), catalytic cracker, catalytic reformer and hydrodesulfurizer combined fractionater: Petroleum gas: [A complex combination of hydrocarbons obtained from the fractionation of products from catalytic cracking, catalytic reforming and hydrodesulfurizing processes treated to remove acidic impurities. It consists predominantly of hydrocarbons having cabon numbers predominantly in the range of C1 through C5.] |  |  | N
273-173-5 | 68952-79-4 | Tail gas (petroleum), catalytic hydrodesulfurized naphtha separator: Refinery gas: [A complex combination of hydrocarbons obtained from the hydrodesulfurization of naphtha. It consists of hydrogen, methane, ethane, and propane.] |  |  | N
269-618-8 | 68307-99-3 | Tail gas (petroleum), catalytic polymn. naphtha fractionation stabilizer: Petroleum gas: [A complex combination of hydrocarbons from the fractionation stabilization products from polymerization of naphtha. It consists predominantly of hydrocarbons having carbon numbers in the range of C1 through C4.] |  |  | N
269-619-3 | 68308-00-9 | Tail gas (petroleum), catalytic reformed naphtha fractionation stabilizer, hydrogen sulfide-free: Petroleum gas: [A complex combination of hydrocarbons obtained from fractionation stabilization of catalytic reformed naphtha and from which hydrogen sulfide has been removed by amine treatment. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
270-806-7 | 68478-26-2 | Tail gas (petroleum), catalytic reformed naphtha fractionation stabilizer: Petroleum gas: [A complex combination of hydrocarbons obtained from the fractionation stabilization of catalytic reformed naphtha. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
270-807-2 | 68478-27-3 | Tail gas (petroleum), catalytic reformed naphtha separator: Refinery gas: [A complex combination of hydrocarbons obtained from the catalytic reforming of straight run naphtha. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
270-808-8 | 68478-28-4 | Tail gas (petroleum), catalytic reformed naphtha stabilizer: Refinery gas: [A complex combination of hydrocarbons obtained from the stabilization of catalytic reformed naphtha. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
270-809-3 | 68478-29-5 | Tail gas (petroleum), cracked distillate hydrotreater separator: Refinery gas: [A complex combination of hydrocarbons obtained by treating cracked distillates with hydrogen in the presence of a catalyst. It consists of hydrogen and saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
269-620-9 | 68308-01-0 | Tail gas (petroleum), cracked distillate hydrotreater stripper: Petroleum gas: [A complex combination of hydrocarbons obtained by treating thermal cracked distillates with hydrogen in the presence of a catalyst. It consists predominantly of saturated hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
269-623-5 | 68308-03-2 | Tail gas (petroleum), gas oil catalytic cracking absorber: Petroleum gas: [A complex combination of hydrocarbons obtained from the distillation of products from the catalytic cracking of gas oil. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
269-625-6 | 68308-05-4 | Tail gas (petroleum), gas recovery plant deethanizer: Petroleum gas: [A complex combination of hydrocarbons from the distillation of products from miscellaneous hydrocarbon streams. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
269-624-0 | 68308-04-3 | Tail gas (petroleum), gas recovery plant: Petroleum gas: [A complex combination of hydrocarbons from the distillation of products from miscellaneous hydrocarbon streams. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
269-626-1 | 68308-06-5 | Tail gas (petroleum), hydrodesulfurized distillate and hydrodesulfurized naphtha fractionator, acid-free: Petroleum gas: [A complex combination of hydrocarbons obtained from fractionation of hydrodesulfurized naphtha and distillate hydrocarbon streams and treated to remove acidic impurities. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
270-810-9 | 68478-30-8 | Tail gas (petroleum), hydrodesulfurized straight-run naphtha separator: Refinery gas: [A complex combination of hydrocarbons obtained from hydrodesulfurization of straight-run naphtha. It consists of hydrogen and saturated aliphatic hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
269-627-7 | 68308-07-6 | Tail gas (petroleum), hydrodesulfurized vacuum gas oil stripper, hydrogen sulfide-free: Petroleum gas: [A complex combination of hydrocarbons obtained from stripping stabilization of catalytic hydrodesulfurized vacuum gas oil and from which hydrogen sulfide has been removed by amine treatment. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
269-628-2 | 68308-08-7 | Tail gas (petroleum), isomerized naphtha fractionation stabilizer: Petroleum gas: [A complex combination of hydrocarbons obtained from the fractionation stabilization products from isomerized naphtha. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
269-629-8 | 68308-09-8 | Tail gas (petroleum), light straight-run naphtha stabilizer, hydrogen sulfide-free: Petroleum gas: [A complex combination of hydrocarbons obtained from fractionation stabilization of light straight run naphtha and from which hydrogen sulfide has been removed by amine treatment. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
269-631-9 | 68308-11-2 | Tail gas (petroleum), propane-propylene alkylation feed prep deethanizer: Petroleum gas: [A complex combination of hydrocarbons obtained from the distillation of the reaction products of propane with propylene. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
270-813-5 | 68478-32-0 | Tail gas (petroleum), saturate gas plant mixed stream, C4-rich: Petroleum gas: [A complex combination of hydrocarbons obtained from the fractionation stabilization of straight-run naphtha, distillation tail gas and catalytic reformed naphtha stabilizer tail gas. It consists of hydrocarbons having carbon numbers in the range of C3 through C6, predominantly butane and isobutane.] |  |  | N
270-814-0 | 68478-33-1 | Tail gas (petroleum), saturate gas recovery plant, C1-2-rich: Petroleum gas: [A complex combination of hydrocarbons obtained from fractionation of distillate tail gas, straight-run naphtha, catalytic reformed naphtha stabilizer tail gas. It consists predominantly of hydrocarbons having carbon numbers in the range of C1through C5, predominantly methane and ethane.] |  |  | N
269-630-3 | 68308-10-1 | Tail gas (petroleum), straight-run distillate hydrodesulfurizer, hydrogen sulfide-free: Petroleum gas: [A complex combination of hydrocarbons obtained from catalytic hydrodesulfurization of straight run distillates and from which hydrogen sulfide has been removed by amine treatment. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C4.] |  |  | N
273-174-0 | 68952-80-7 | Tail gas (petroleum), straight-run naphtha hydrodesulfurizer: Refinery gas: [A complex combination obtained from the hydrodesulfurization of straight-run naphtha. It consists of hydrogen and hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
273-176-1 | 68952-82-9 | Tail gas (petroleum), thermal cracked hydrocarbon fractionation stabilizer, petroleum coking: Petroleum gas: [A complex combination of hydrocarbons obtained from the fractionation stabilization of thermal cracked hydrocarbons from petroleum coking process. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
273-175-6 | 68952-81-8 | Tail gas (petroleum), thermal-cracked distillate, gas oil and naphtha absorber: petroleum gas: [A complex combination of hydrocarbons obtained from the separation of thermal-cracked distillates, naphtha and gas oil. It consists pedrominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
269-632-4 | 68308-12-3 | Tail gas (petroleum), vacuum gas oil hydrodesulfurizer, hydrogen sulfide-free: Petroleum gas: [A complex combination of hydrocarbons obtained from catalytic hydrodesulfurization of vacuum gas oil and from which hydrogen sulfide has been removed by amine treatment. It consists predominantly of hydrocarbons having carbon numbers predominantly in the range of C1 through C6.] |  |  | N
270-815-6 | 68478-34-2 | Tail gas (petroleum), vacuum residues thermal cracker: Petroleum gas: [A complex combination of hydrocarbons obtained from the thermal cracking of vacuum residues. It consists of hydrocarbons having carbon numbers predominantly in the range of C1 through C5.] |  |  | N
284-896-0 | 84989-07-1 | Tar acids, 3,5-xylenol fraction: Distillate Phenols: [The fraction of tar acids, rich in 3,5-dimethylphenol, recovered by distillation of low-temperature coal tar acids.] |  |  | N
295-536-7 | 92062-22-1 | Tar acids, brown-coal gasification: Crude Phenols: [A complex combination of organic compounds obtained from brown coal gasification. Composed primarily of C<sub>6-10</sub> hydroxy aromatic phenols and their homologs.] |  |  | N
302-662-9 | 94114-29-1 | Tar acids, brown-coal, C<sub>2</sub>-alkylphenol fraction: Distillate Phenols: [The distillate from the acidification of alkaline washed lignite tar distillate boiling in the range of approximately 200°C to 230°C (392°F to 446°F). Composed primarily of <i>m</i>- and <i>p</i>-ethylphenol as well as cresols and xylenols.] |  |  | N
309-888-7 | 101316-86-3 | Tar acids, brown-coal, crude: Crude Phenols: [An acidified alkaline extract of brown coal tar distillate. Composed primarily of phenol and phenol homologs.] |  |  | N
266-019-3 | 65996-85-2 | Tar acids, coal, crude: Crude Phenols: [The reaction product obtained by neutralizing coal tar oil alkaline extract with an acidic solution, such as aqueous sulfuric acid, or gaseous carbon dioxide, to obtain the free acids. Composed primarily of tar acids such as phenol, cresols, and xylenols.] |  |  | N
271-418-0 | 68555-24-8 | Tar acids, cresylic, residues: Distillate Phenols: [The residue from crude coal tar acids after removal of phenol, cresols, xylenols and any higher boiling phenols. A black solid with a melting point approximately 80°C (176°F). Composed primarily of polyalkylphenols, resin gums, and inorganic salts.] |  |  | N
272-361-4 | 68815-21-4 | Tar acids, cresylic, sodium salts, caustic solns.: Alkaline Extract |  |  | N
295-540-9 | 92062-26-5 | Tar acids, cresylic: Distillate Phenols: [A complex combination of organic compounds obtained from brown coal and boiling in the range of approximately 200°C to 230°C (392°F to 446°F). It contains chiefly phenols and pyridine bases.] |  |  | N
306-251-5 | 96690-55-0 | Tar acids, distn. residues: Distillate Phenols: [A residue from the distillation of crude phenol from coal. It consists predominantly of phenols having carbon numbers in the range of C<sub>8</sub> through C<sub>10</sub> with a softening point of 60°C to 80°C (140°F to 176°F).] |  |  | N
284-891-3 | 84989-03-7 | Tar acids, ethylphenol fraction: Distillate Phenols: [The fraction of tar acids, rich in 3- and 4-ethylphenol, recovered by distillation of low-temperature coal tar crude tar acids.] |  |  | N
284-892-9 | 84989-04-8 | Tar acids, methylphenol fraction: Distillate Phenols: [The fraction of tar acid rich in 3- and 4-methylphenol, recovered by distillation of low-temperature coal tar crude tar acids.] |  |  | N
284-893-4 | 84989-05-9 | Tar acids, polyalkylphenol fraction: Distillate Phenols: [The fraction of tar acids, recovered by distillation of low-temperature coal tar crude tar acids, having an approximate boiling range of 225°C to 320°C (437°F to 608°F). Composed primarily of polyalkylphenols.] |  |  | N
270-713-1 | 68477-23-6 | Tar acids, residues, distillates, first-cut: Distillate Phenols: [The residue from the distillation in the range of 235°C to 355°C (481°F to 697°F) of light carbolic oil.] |  |  | N
284-895-5 | 84989-06-0 | Tar acids, xylenol fraction: Distillate Phenols: [The fraction of tar acids, rich in 2,4- and 2,5-dimethylphenol, recovered by distillation of low-temperature coal tar crude tar acids.] |  |  | N
295-541-4 | 92062-27-6 | Tar bases, coal, aniline fraction: Distillate Bases: [The distillation fraction boiling in the range of approximately 180 °C to 200 °C (356 °F to 392 °F) from the crude bases obtained by dephenolating and debasing the carbolated oil from the distillation of coal tar. It contains chiefly aniline, collidines, lutidines and toluidines.] |  |  | N
295-543-5 | 92062-28-7 | Tar bases, coal, collidine fraction: Distillate Bases: [The distillation fraction boiling in the range of approximately 181 °C to 186 °C (356 °F to 367 °F) from the crude bases obtained from the neutralized, acid-extracted base-containing tar fractions obtained by the distillation of bituminous coal tar. It contains chiefly aniline and collidines.] |  |  | N
266-018-8 | 65996-84-1 | Tar bases, coal, crude: Crude Tar Bases: [The reaction product obtained by neutralizing coal tar base extract oil with an alkaline solution, such as aqueous sodium hydroxide, to obtain the free bases. Composed primarily of such organic bases as acridine, phenanthridine, pyridine, quinoline and their alkyl derivatives.] |  |  | N
295-544-0 | 92062-29-8 | Tar bases, coal, distn. residues: Distillate Bases: [The distillation residue remaining after the distillation of the neutralized, acid-extracted base-containing tar fractions obtained by the distillation of coal tars. It contains chiefly aniline, collidines, quinoline and quinoline derivatives and toluidines.] |  |  | N
293-766-2 | 91082-52-9 | Tar bases, coal, lutidine fraction: Distillate Bases |  |  | N
295-548-2 | 92062-33-4 | Tar bases, coal, picoline fraction: Distillate Bases: [Pyridine bases boiling in the range of approximately 125°C to 160°C (257°F 320°F) obtained by distillation of neutralized acid extract of the base-containing tar fraction obtained by the distillation of bituminous coal tars. Composed chiefly of lutidines and picolines.] |  |  | N
274-560-1 | 70321-67-4 | Tar bases, coal, quinoline derivs. fraction: Distillate Bases |  |  | N
293-767-8 | 91082-53-0 | Tar bases, coal, toluidine fraction: Distillate Bases |  |  | N
271-020-7 | 68513-87-1 | Tar bases, quinoline derivs.: Distillate Bases |  |  | N
309-885-0 | 101316-83-0 | Tar brown-coal: [An oil distilled from brown-coal tar. Composed primarily of aliphatic, naphthenic and one- to three-ring aromatic hydrocarbons, their alkyl derivates, heteroaromatics and one- and two-ring phenols boiling in the range of approximately 150 °C to 360 °C (302 °F to 680 °F).] |  |  | N
302-674-4 | 94114-40-6 | Tar oils, brown-coal: Light Oil: [The distillate from lignite tar boiling in the range of approximately 80°C to 250°C (176°F to 482°F). Composed primarily of aliphatic and aromatic hydrocarbons and monobasic phenols.] |  |  | N
309-889-2 | 101316-87-4 | Tar oils, coal, low-temp.: Tar Oil, high boiling: [A distillate from low-temperature coal tar. Composed primarily of hydrocarbons, phenolic compounds and aromatic nitrogen bases boiling in the range of approximately 160°C to 340°C (320°F to 644°F).] |  |  | N
266-016-7 | 65996-82-9 | Tar oils, coal: Carbolic Oil: [The distillate from high temperature coal tar having an approximate distillation range of 130°C to 250°C (266°F to 410°F). Composed primarily of naphthalene, alkylnaphthalenes, phenolic compounds, and aromatic nitrogen bases.] |  |  | N
309-886-6 | 101316-84-1 | Tar, brown-coal, low-temp.: [A tar obtained from low temperature carbonization and low temperature gasification of brown coal. Composed primarily of aliphatic, naphthenic and cyclic aromatic hydrocarbons, heteroaromatic hydrocarbons and cyclic phenols.] |  |  | N
295-535-1 | 92062-20-9 | Tar, coal, high-temp., distn. and storage residues: Coal Tar Solids Residue: [Coke- and ash-containing solid residues that separate on distillation and thermal treatment of bituminous coal high temperature tar in distillation installations and storage vessels. Consists predominantly of carbon and contains a small quantity of hetero compounds as well as ash components.] |  |  | N
273-615-7 | 68990-61-4 | Tar, coal, high-temp., high-solids: Coal Tar Solids Residue: [The condensation product obtained by cooling, to approximately ambient temperature, the gas evolved in the high temperature (greater than 700 °C (1292 °F)) destructive distillation of coal. Composed primarily of a complex mixture of condensed ring aromatic hydrocarbons with a high solid content of coal-type materials.] |  |  | N
309-726-5 | 100684-51-3 | Tar, coal, high-temp., residues: Coal Tar Solids Residue: [Solids formed during the coking of bituminous coal to produce crude bituminous coal high temperature tar. Composed primarily of coke and coal particles, highly aromatized compounds and mineral substances.] |  |  | N
266-024-0 | 65996-89-6 | Tar, coal, high-temp.: Coal tar: [The condensation product obtained by cooling, to approximately ambient temperature, the gas evolved in the high temperature (greater than 700 °C (1292 °F)) destructive distillation of coal. A black viscous liquid denser than water. Composed primarily of a complex mixture of condensed ring aromatic hydrocarbons. May contain minor amounts of phenolic compounds and aromatic nitrogen bases.] |  |  | N
309-887-1 | 101316-85-2 | Tar, coal, low-temp., distn. residues: Tar Oil, intermediate boiling: [Residues from fractional distillation of low temperature coal tar to remove oils that boil in a range up to approximately 300 °C (572 °F). Composed primarily of aromatic compounds.] |  |  | N
266-025-6 | 65996-90-9 | Tar, coal, low-temp.: Coal oil: [The condensation product obtained by cooling, to approximately ambient temperature, the gas evolved in low temperature (less than 700 °C (1292 °F)) destructive distillation of coal. A black viscous liquid denser than water. Composed primarily of condensed ring aromatic hydrocarbons, phenolic compounds, aromatic nitrogen bases, and their alkyl derivatives.] |  |  | N
293-764-1 | 91082-50-7 | Tar, coal, storage residues: Coal Tar Solids Residue: [The deposit removed from crude coal tar storages. Composed primarily of coal tar and carbonaceous particulate matter.] |  |  | N
232-361-7 | 8007-45-2 | Tar, coal: Coal tar: [The by-product from the destructive distillation of coal. Almost black semisolid. A complex combination of aromatic hydro-carbons, phenolic compounds, nitrogen bases and thiophene.] |  |  | N
235-541-3 | 12267-73-1 | tetraboron disodium heptaoxide, hydrate |  | Repr. 1b | 1 | SVHC: Also known as borax, has many uses including in household cleaning products, in steel welding fluxes, for silver soldering, as an anti-fungal agent (e.g. in wood), as a fire retardant for cellulose, paper and wood, in welding fluxes, in brake fluids, metal working fluids, in antifreeze as a corrosion inhibitor and to make other boron compounds. It is also used in starch adhesives used in corrugated cardboard packaging
236-669-2 | 13463-39-3 | tetracarbonylnickel: nickel tetracarbonyl |  |  | N | Process chemical
| 3563-45-9 | Tetrachloro-DDT |  |  | N | Banned insecticide
202-625-6 | 97-99-4 | tetrahydro-2-furylmethanol: tetrahydrofurfuryl alcohol |  |  | N | Solvent and process chemical
407-330-8 | 61571-06-0 | tetrahydrothiopyran-3-carboxaldehyde |  |  | N | Process chemical
220-012-1 | 2602-46-2 | tetrasodium 3,3'-[[1,1'-biphenyl]-4,4'-diylbis(azo)]bis[5-amino-4-hydroxynaphthalene-2,7-disulphonate]: C.I. Direct Blue 6 |  |  | 1 | Pigment
601-147-9 | 111988-49-9 | thiacloprid (ISO): (Z)-3-(6-chloro-3-pyridylmethyl)-1,3-thiazolidin-2-ylidenecyanamide: {(2Z)-3-[(6-chloropyridin-3-yl)methyl]-1,3-thiazolidin-2-ylidene}cyanamide |  |  | N | Agricultural insecticide
200-541-4 | 62-55-5 | thioacetamide |  |  | N | Previously used as an organic solvent in the leather, textile, and paper industries, as an accelerator in the vulcanization of buna rubber (synthetic polybutadiene), and as a stabilizer of motor fuel. However, there is no evidence that it is currently used for any of these purposes. Now used as a precess and lab chemmical
| 137-26-8 | Thiram |  |  | N | Agricultural fungicide
265-697-8 | 65321-67-7 | toluene-2,4-diammonium sulphate: 4-methyl-m-phenylenediamine sulfate |  |  | N | Process chemical
| 8001-35-2 | Toxaphene (Camphechlor) |  |  | N | Insecticide
259-537-6 | 55219-65-3 | triadimenol (ISO): (1RS,2RS;1RS,2SR)-1-(4-chlorophenoxy)-3,3-dimethyl-1-(1H-1,2,4-triazol-1-yl)butan-2-ol: α-tert-butyl-β-(4-chlorophenoxy)-1H-1,2,4-triazole-1-ethanol |  |  | N | Agricultural fungicide
201-167-4 | 79-01-6 | trichloroethylene: trichloroethene |  |  | N | Solvent
246-347-3 | 24602-86-6 | tridemorph (ISO): 2,6-dimethyl-4-tridecylmorpholine |  |  | N | Agricultural fungicide
427-700-2 | 15606-95-8 | triethyl arsenate |  |  | N | Process chemical
614-708-8 | 68694-11-1 | triflumizole (ISO): (1E)-N-[4-chloro-2-(trifluoromethyl)phenyl]-1-(1H-imidazol-1-yl)-2-propoxyethanimine |  |  | N | Agricultural fungicide
235-688-3 | 12519-85-6 | trihydrogen hydroxybis[orthosilicato(4-)]trinickelate(3-) |  |  | N | Industrial process catalyst
231-205-5 | 7446-27-7 | trilead bis(orthophosphate) |  |  | 1 | PVC (and rarely other polymer) stabiliser, used to make special glasses (but not present as trilead bis(orthophosphate)), some medicinal uses.
236-771-7 | 13477-70-8 | trinickel bis(arsenate): nickel(II) arsenate |  |  | N | Process chemical
| 74646-29-0 | trinickel bis(arsenite) |  |  | N | Process chemical
233-844-5 | 10381-36-9 | trinickel bis(orthophosphate) |  |  | N | Process chemical
234-495-1 | 12007-02-2 | trinickel boride |  |  | 1 | Used as a process chemical but also used as a hard coating (to replace chrome)
234-829-6 | 12035-72-2 | trinickel disulfide: nickel subsulfide |  |  | N | Catalyst
| 12137-12-1 | trinickel tetrasulfide |  |  | N | Catalyst
204-118-5 | 115-96-8 | tris(2-chloroethyl)phosphate |  |  | 1 | SVHC: Flame retardant and additive plasticizer used in various polymers including rigid polyurethane and flexible polyurethane foam, textile back coatings, PVC (uncommon), cellulose esters and polyester resin
413-590-3 | 164058-22-4 | trisodium [4'-(8-acetylamino-3,6-disulfonato-2-naphthylazo)-4''-(6-benzoylamino-3-sulfonato-2-naphthylazo)-biphenyl-1,3',3'',1'''-tetraolato-<i>O</i>,<i>O</i>',<i>O</i>'',<i>O</i>''']copper(II) |  |  | 1 | Believed to be used as a pigment (colour unknown)
246-677-8 | 25155-23-1 | trixylyl phosphate |  |  | 1 | SVHC: Fire retardant for many types of plastics, fire retardant in hydraulic fluids (e.g. for industrial generators).
200-123-1 | 51-79-6 | urethane (INN): ethyl carbamate |  |  | N | Intermediate. Previously used as a neoplastic agent
256-599-6 | 50471-44-8 | vinclozolin (ISO): <i>N</i>-3,5-dichlorophenyl-5-methyl-5-vinyl-1,3-oxazolidine-2,4-dione |  |  | N | Agricultural fungicide
200-831-0 | 75-01-4 | vinyl chloride: chloroethylene |  |  | N | Monomer
201-377-6 | 81-81-2 | warfarin (ISO): 4-hydroxy-3-(3-oxo-1-phenylbutyl)-2H-chromen-2-one |  |  | N | Rat poison and medicinal drug
295-549-8 | 92062-34-5 | Waste solids, coal-tar pitch coking: Coal Tar Solids Residue: [The combination of wastes formed by the coking of bituminous coal tar pitch. It consists predominantly of carbon.] |  |  | N | Waste
| 12122-67-7 | Zineb |  |  | N | Agricultural fungicide
226-009-1 | 5216-25-1 | α, α,α,4-tetrachlorotoluene: <i>p</i>-chlorobenzotrichloride |  |  | N | Process chemical
202-634-5 | 98-07-7 | α,α,α-trichlorotoluene: benzotrichloride |  |  | N | Process chemical
202-853-6 | 100-44-7 | α-chlorotoluene: benzyl chloride |  |  | N | Process chemical
|  | lead compounds with the exception of those specified elsewhere in this Annex |  |  | 1 | Various uses, but most have no patient contact. Includes lead zinconium niobate (PZT) used in ultrasound transducers and piezo devices, lead titanate (used in some electronic components), lead dioxide (used corrosion resistant paint and in lead-acid batteries), as PVC stabilisers (see REACH SVHC list of lead compounds), leaded glass (e.g. optical glass), white pigments (hydoxide and carbonate are banned by REACH).
|  | lead alkyls |  |  | N | Fuel additive and process chemical
|  | zinc chromates including zinc potassium chromate |  |  | 1 | Primer coatings
|  | methacrylamide |  |  | N | Monomer
|  | reaction products of paraformaldehyde and 2-hydroxypropylamine (ratio 3:2) | [MBO] |  | N | Biocide
|  | [formaldehyde released from 3,3′-methylenebis[5-methyloxazolidine] |  |  | N | Biocide
|  | [formaldehyde released from oxazolidin] |  |  | N | Biocide
|  | reaction products of paraformaldehyde with 2-hydroxypropylamine (ratio 1:1) | [HPT] |  | N | Biocide
|  | [formaldehyde released from α,α,α-trimethyl-1,3,5-triazine-1,3,5(2H,4H,6H)-triethanol] |  |  | N | Biocide
|  | Benzidine based azo dyes | 4,4'-diarylazobiphenyl dyes, with the exception of those specified elsewhere in this Annex |  | 1 | Pigments mainly used with textiles
|  | [formaldehyde released from N,N′-methylenebismorpholine] | [MBM] |  | N | Biocide
|  | Chromium (VI) compounds, with the exception of barium chromate and of compounds specified elsewhere in this Annex |  |  | 1 | Pigments, corrosion inhibitors
|  | e-glass fibres, [Calcium-aluminium-silicate fibres with random orientation with the following representative composition (% given by weight): SiO2 50,0-56,0 %, Al2O3 13,0-16,0 %, B2O3 5,8-10,0 %, Na2O < 0,6 %, K2O < 0,4 %, CaO 15,0-24,0 %, MgO < 5,5 %, Fe2O3 < 0,5 %, F2 < 1,0 %. Process: typically produced by flame attenuation and rotary process. (Additional individual elements may be present at low levels; the process list does not preclude innovation).] |  |  | 1 | Refractory insulation
|  | salts of hydrazine |  |  | N | Process chemicals
|  | reaction products of paraformaldehyde and 2-hydroxypropylamine (ratio 3:2): [formaldehyde released from 3,3′-methylenebis[5-methyloxazolidine]: [formaldehyde released from oxazolidin]: [MBO]: reaction products of paraformaldehyde with 2-hydroxypropylamine (ratio 1:1): [formaldehyde released from α,α,α-trimethyl-1,3,5-triazine-1,3,5(2H,4H,6H)-triethanol]: [HPT] |  |  | N | Biocide